## Documents
Slug | Name | Short Code | Sections | Requirements | Created
architecture-decisions | Architecture Decisions | ARC | 1 | 0 | 2026-03-15T00:01:47.006Z
interface-requirements | Interface Requirements | IFC | 1 | 18 | 2026-03-15T00:01:46.446Z
stakeholder-requirements | Stakeholder Requirements | STK | 1 | 6 | 2026-03-15T00:01:44.814Z
subsystem-requirements | Subsystem Requirements | SUB | 1 | 46 | 2026-03-15T00:01:45.813Z
system-requirements | System Requirements | SYS | 1 | 12 | 2026-03-15T00:01:45.320Z
verification-plan | Verification Plan | VER | 1 | 19 | 2026-03-15T00:01:47.569Z

## Requirements
Ref | Text | Document | Section | Pattern | Verification | QA Score | Compliance | Rationale | Tags
IFC-INTERFACEDEFINITIONS-001 | The interface between the ECG Acquisition Module and the Signal Conditioning Unit SHALL convey digitised ECG samples at 500 Hz per channel via a synchronous serial data bus with a maximum latency of 2 ms. | interface-requirements | Interface Definitions | [5 empty] | interface, vital-signs-acquisition, ecg, session-170
IFC-INTERFACEDEFINITIONS-002 | The interface between the Pulse Oximetry Module and the Signal Conditioning Unit SHALL transmit red and infrared photoplethysmogram samples at a minimum rate of 100 Hz with time-stamped synchronisation to the system clock. | interface-requirements | Interface Definitions | [5 empty] | interface, vital-signs-acquisition, spo2, session-170
IFC-INTERFACEDEFINITIONS-003 | The interface between the NIBP Measurement Module and the Signal Conditioning Unit SHALL report cuff pressure readings at 200 Hz during inflation and deflation, and SHALL signal measurement completion with systolic, diastolic, and MAP values. | interface-requirements | Interface Definitions | [5 empty] | interface, vital-signs-acquisition, nibp, session-170
IFC-INTERFACEDEFINITIONS-004 | The interface between the Signal Conditioning Unit and the Network and Communication Subsystem SHALL output calibrated vital sign parameter packets at 1 Hz containing heart rate, SpO2, NIBP, and temperature values conforming to IEEE 11073 encoding. | interface-requirements | Interface Definitions | [5 empty] | interface, vital-signs-acquisition, network, session-170
IFC-INTERFACEDEFINITIONS-005 | The interface between the Vital Signs Acquisition Subsystem and the Alarm Detection Engine SHALL deliver calibrated vital sign samples as time-stamped parameter-value pairs at the acquisition sampling rate, with parameter identity, unit of measure, and signal quality indicator. | interface-requirements | Interface Definitions | [5 empty] | interface, alarm-management, session-171
IFC-INTERFACEDEFINITIONS-006 | The interface between the Alarm Detection Engine and the Alarm Prioritization Module SHALL convey alarm condition events containing the violated parameter identity, threshold value, measured value, direction of violation, and detection timestamp. | interface-requirements | Interface Definitions | [5 empty] | interface, alarm-management, session-171
IFC-INTERFACEDEFINITIONS-007 | The interface between the Alarm Notification Controller and the Central Monitoring Station SHALL transmit alarm notification messages containing patient identifier, bed location, alarm priority, parameter identity, and current value using a reliable transport with delivery confirmation. | interface-requirements | Interface Definitions | [5 empty] | interface, alarm-management, session-171
IFC-INTERFACEDEFINITIONS-008 | The interface between the Alarm Notification Controller and the nurse call system SHALL conform to the facility nurse call protocol and transmit crisis and warning alarms with bed location and alarm type within 3 seconds of alarm activation. | interface-requirements | Interface Definitions | [5 empty] | interface, alarm-management, session-171
IFC-INTERFACEDEFINITIONS-009 | The interface between the Alarm Management Subsystem and the Alert Dashboard Controller SHALL deliver alarm events via an asynchronous message queue with guaranteed delivery, conveying alarm identifier, priority level, parameter name, threshold value, measured value, and patient identifier within 500 milliseconds of alarm activation. | interface-requirements | Interface Definitions | [5 empty] | interface, central-monitoring-station, session-172
IFC-INTERFACEDEFINITIONS-010 | The interface between the Vital Signs Acquisition Subsystem and the Patient Tile Manager SHALL stream digitised vital sign samples via a publish-subscribe transport, with each message containing the patient identifier, parameter type, sample value, and acquisition timestamp, at a rate matching the source sampling frequency. | interface-requirements | Interface Definitions | [5 empty] | interface, central-monitoring-station, session-172
IFC-INTERFACEDEFINITIONS-011 | The interface between the Alert Dashboard Controller and the Alarm Management Subsystem SHALL convey clinician alarm acknowledgement and silence commands, including the operator identity, action type, and alarm identifier, with a round-trip confirmation latency of no more than 1 second. | interface-requirements | Interface Definitions | [5 empty] | interface, central-monitoring-station, session-172
IFC-INTERFACEDEFINITIONS-012 | The interface between the HL7/FHIR Interface Engine and the Clinical Data Repository SHALL transfer patient demographic, medication, and allergy data using an internal normalized patient context schema, with each update carrying a timestamp and source EHR system identifier. | interface-requirements | Interface Definitions | [5 empty] | interface, clinical-data-integration, session-173
IFC-INTERFACEDEFINITIONS-013 | The interface between the Clinical Data Repository and the Alarm Prioritization Module SHALL provide patient clinical context (active medications, known conditions) within 50 milliseconds upon query to support context-aware alarm priority adjustment. | interface-requirements | Interface Definitions | [5 empty] | interface, clinical-data-integration, session-173
IFC-INTERFACEDEFINITIONS-014 | The interface between the Data Export Controller and the Network and Communication Subsystem SHALL support TLS 1.2 or higher encrypted transport for all outbound clinical data exports, with mutual certificate authentication to the receiving EMR endpoint. | interface-requirements | Interface Definitions | [5 empty] | interface, clinical-data-integration, session-173
IFC-INTERFACEDEFINITIONS-015 | The interface between the Medical-Grade Network Switch and Wireless Access Controller SHALL use IEEE 802.1Q tagged VLAN trunks with dedicated VLAN IDs for clinical monitoring traffic and management traffic. | interface-requirements | Interface Definitions | [5 empty] | interface, network, session-174
IFC-INTERFACEDEFINITIONS-016 | The interface between the Medical-Grade Network Switch and Clinical Network Firewall SHALL carry all inter-VLAN traffic through a trunk port supporting IEEE 802.1Q at a minimum bandwidth of 10 Gbps. | interface-requirements | Interface Definitions | [5 empty] | interface, network, session-174
IFC-INTERFACEDEFINITIONS-017 | The interface between the Message Broker and Data Distribution Service and subscribing subsystems SHALL use a publish-subscribe protocol with topic hierarchies structured as patient/parameter/event-type, with message payloads conforming to a common JSON schema. | interface-requirements | Interface Definitions | [5 empty] | interface, network, session-174
IFC-INTERFACEDEFINITIONS-018 | The interface between the Time Synchronization Service and bedside monitoring devices SHALL use IEEE 1588 PTPv2 messages over multicast UDP, with the grandmaster announcing at a minimum rate of 1 message per second. | interface-requirements | Interface Definitions | [5 empty] | interface, network, session-174
STK-STAKEHOLDERNEEDS-001 | The Hospital Patient Monitoring System SHALL continuously monitor patient vital signs including heart rate, blood pressure, oxygen saturation, respiration rate, and body temperature without interruption during the patient's stay. | stakeholder-requirements | Stakeholder Needs | [5 empty] | stakeholder, session-169
STK-STAKEHOLDERNEEDS-002 | The Hospital Patient Monitoring System SHALL alert clinical staff within clinically acceptable timeframes when patient vital signs deviate from prescribed thresholds. | stakeholder-requirements | Stakeholder Needs | [5 empty] | stakeholder, session-169
STK-STAKEHOLDERNEEDS-003 | The Hospital Patient Monitoring System SHALL integrate with the hospital Electronic Health Record system to provide a unified patient data view and avoid duplicate data entry. | stakeholder-requirements | Stakeholder Needs | [5 empty] | stakeholder, session-169
STK-STAKEHOLDERNEEDS-004 | The Hospital Patient Monitoring System SHALL comply with applicable medical device regulations including IEC 60601, IEC 62304, and FDA 21 CFR Part 820 quality system requirements. | stakeholder-requirements | Stakeholder Needs | [5 empty] | stakeholder, session-169
STK-STAKEHOLDERNEEDS-005 | The Hospital Patient Monitoring System SHALL reduce clinically non-actionable alarms to mitigate alarm fatigue among nursing staff while maintaining sensitivity to genuine patient deterioration events. | stakeholder-requirements | Stakeholder Needs | [5 empty] | stakeholder, session-169
STK-STAKEHOLDERNEEDS-006 | The Hospital Patient Monitoring System SHALL protect patient health information in accordance with HIPAA privacy and security rules and hospital data governance policies. | stakeholder-requirements | Stakeholder Needs | [5 empty] | stakeholder, session-169
SUB-SUBSYSTEMREQUIREMENTS-001 | The ECG Acquisition Module SHALL acquire continuous electrocardiogram signals using a minimum 3-lead electrode configuration with a common-mode rejection ratio of at least 90 dB. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, vital-signs-acquisition, ecg, session-170
SUB-SUBSYSTEMREQUIREMENTS-002 | The ECG Acquisition Module SHALL digitise the analogue ECG signal with a bandwidth of 0.05 Hz to 150 Hz and a minimum sampling rate of 500 samples per second per channel. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, vital-signs-acquisition, ecg, session-170
SUB-SUBSYSTEMREQUIREMENTS-003 | The Pulse Oximetry Module SHALL measure arterial oxygen saturation (SpO2) with an accuracy of +/- 2% over the range of 70% to 100% SpO2. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, vital-signs-acquisition, spo2, session-170
SUB-SUBSYSTEMREQUIREMENTS-004 | The Pulse Oximetry Module SHALL derive pulse rate from the photoplethysmogram signal with an accuracy of +/- 3 beats per minute over the range 30 to 250 bpm. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, vital-signs-acquisition, spo2, session-170
SUB-SUBSYSTEMREQUIREMENTS-005 | The NIBP Measurement Module SHALL perform oscillometric blood pressure determination with an accuracy of +/- 5 mmHg for systolic and diastolic pressures over the range 40 to 260 mmHg. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, vital-signs-acquisition, nibp, session-170
SUB-SUBSYSTEMREQUIREMENTS-006 | The NIBP Measurement Module SHALL limit cuff pressure to a maximum of 300 mmHg and SHALL automatically deflate the cuff within 180 seconds of inflation initiation to prevent tissue ischaemia. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, vital-signs-acquisition, nibp, safety, session-170
SUB-SUBSYSTEMREQUIREMENTS-007 | The Signal Conditioning Unit SHALL reject 50 Hz and 60 Hz mains interference from all sensor input channels by at least 40 dB without attenuating the physiological signal bandwidth. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, vital-signs-acquisition, signal-conditioning, session-170
SUB-SUBSYSTEMREQUIREMENTS-008 | While operating on battery backup, the Vital Signs Acquisition Subsystem SHALL continue acquiring all vital sign parameters without data loss for a minimum of 30 minutes. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, vital-signs-acquisition, reliability, session-170
SUB-SUBSYSTEMREQUIREMENTS-009 | The Alarm Detection Engine SHALL evaluate each incoming vital sign sample against the configured upper and lower alarm thresholds for that parameter within 500 milliseconds of receipt. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-010 | The Alarm Detection Engine SHALL support compound alarm rules that evaluate Boolean combinations of up to four vital sign parameters to detect multi-parameter clinical deterioration patterns. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-011 | The Alarm Prioritization Module SHALL classify each alarm condition into one of three priority levels — crisis, warning, or advisory — in accordance with IEC 60601-1-8 Clause 6.3. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-012 | When a patient record indicates a known pre-existing condition for a monitored parameter, the Alarm Prioritization Module SHALL adjust the alarm priority according to the clinician-configured exception rules for that condition. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-013 | The Alarm Fatigue Mitigation Unit SHALL apply a configurable alarm delay of 0 to 60 seconds per parameter before promoting a threshold crossing to an active alarm, suppressing transient artifact-induced violations. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-014 | The Alarm Fatigue Mitigation Unit SHALL never suppress or delay crisis-priority alarms regardless of delay configuration or suppression rules. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-015 | When an alarm passes through fatigue mitigation filtering, the Alarm Notification Controller SHALL activate both an audible annunciator and a visual indicator at the bedside within 2 seconds of receiving the filtered alarm event. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-016 | The Alarm Notification Controller SHALL transmit crisis and warning alarms to the nurse call system and designated mobile devices within 5 seconds of alarm activation. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-017 | When a crisis alarm remains unacknowledged for a configurable escalation period (default 120 seconds), the Alarm Notification Controller SHALL escalate the notification to the charge nurse and unit supervisor. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-018 | The Alarm History and Audit Logger SHALL record every alarm event — including detection, prioritization, suppression, notification, acknowledgement, and silencing — with a UTC timestamp, patient identifier, and operator identity in a tamper-evident audit trail. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, alarm-management, session-171
SUB-SUBSYSTEMREQUIREMENTS-019 | The Display Rendering Engine SHALL render continuous ECG waveforms at a minimum refresh rate of 25 frames per second with a sweep speed selectable between 12.5 mm/s and 50 mm/s. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, display, session-172
SUB-SUBSYSTEMREQUIREMENTS-020 | The Display Rendering Engine SHALL update numeric vital sign readouts within 1 second of receiving new parameter values from the Patient Tile Manager. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, display, session-172
SUB-SUBSYSTEMREQUIREMENTS-021 | The Patient Tile Manager SHALL support simultaneous display of at least 32 patient tiles, each showing the active vital sign parameters, alarm status, and patient identifier for one monitored bed. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, tile-manager, session-172
SUB-SUBSYSTEMREQUIREMENTS-022 | When a patient alarm state changes, the Patient Tile Manager SHALL reorder the tile display within 2 seconds so that patients with crisis-priority alarms appear in the most prominent screen positions. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, tile-manager, session-172
SUB-SUBSYSTEMREQUIREMENTS-023 | The Trend Analysis Engine SHALL maintain rolling trend buffers of at least 72 hours duration for each monitored vital sign parameter per patient. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, trend, session-172
SUB-SUBSYSTEMREQUIREMENTS-024 | The Trend Analysis Engine SHALL detect gradual vital sign deterioration by computing a weighted early warning score at 5-minute intervals and SHALL flag patients whose score increases by 2 or more points within any 60-minute window. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, trend, session-172
SUB-SUBSYSTEMREQUIREMENTS-025 | The Alert Dashboard Controller SHALL provide a one-touch alarm acknowledgement control that records the acknowledging clinician identity, timestamp, and alarm identifier in the audit trail. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, alert-dashboard, session-172
SUB-SUBSYSTEMREQUIREMENTS-026 | When a clinician activates the alarm silence function, the Alert Dashboard Controller SHALL suppress the audible annunciator for a maximum of 120 seconds while maintaining the visual alarm indicator. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, alert-dashboard, session-172
SUB-SUBSYSTEMREQUIREMENTS-027 | The Multi-bed Overview Processor SHALL generate a ward-level acuity summary screen that colour-codes each bed by the highest active alarm priority and updates within 3 seconds of any alarm state change. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, multi-bed, session-172
SUB-SUBSYSTEMREQUIREMENTS-028 | The Multi-bed Overview Processor SHALL compute and display per-bed alarm frequency statistics over configurable time windows of 1, 4, 8, and 24 hours to support alarm burden assessment by charge nurses. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, central-monitoring-station, multi-bed, session-172
SUB-SUBSYSTEMREQUIREMENTS-029 | The HL7/FHIR Interface Engine SHALL support bidirectional communication using HL7 v2.x messaging (ADT, ORU, ORM message types) and FHIR R4 RESTful API operations for patient, observation, and device resources. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, clinical-data-integration, session-173
SUB-SUBSYSTEMREQUIREMENTS-030 | The HL7/FHIR Interface Engine SHALL process inbound HL7 messages and FHIR resource requests within 500 milliseconds of receipt under normal operating load. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, clinical-data-integration, session-173
SUB-SUBSYSTEMREQUIREMENTS-031 | When an ADT Admit event is received for a bed with an active bedside monitor, the Patient ADT Manager SHALL automatically associate the patient record with the monitor within 2 seconds of event receipt. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, clinical-data-integration, session-173
SUB-SUBSYSTEMREQUIREMENTS-032 | When an ADT Discharge event is received, the Patient ADT Manager SHALL disassociate the patient record from all bound monitors and archive the monitoring session data within 5 seconds. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, clinical-data-integration, session-173
SUB-SUBSYSTEMREQUIREMENTS-033 | The Clinical Data Repository SHALL refresh cached patient demographics and medication data from the EHR at a configurable interval not exceeding 15 minutes and upon receipt of any ADT update event for the patient. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, clinical-data-integration, session-173
SUB-SUBSYSTEMREQUIREMENTS-034 | The Clinical Data Repository SHALL respond to patient context queries from the Alarm Management Subsystem and Central Monitoring Station within 50 milliseconds. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, clinical-data-integration, session-173
SUB-SUBSYSTEMREQUIREMENTS-035 | The Data Export Controller SHALL export continuous vital sign observation data to the hospital EMR in HL7 ORU message format or FHIR Observation resources at a configurable interval of 1 to 60 minutes. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, clinical-data-integration, session-173
SUB-SUBSYSTEMREQUIREMENTS-036 | The Terminology and Coding Service SHALL map all exported vital sign measurements to their corresponding LOINC observation codes with a mapping completeness of at least 99% for standard vital signs (heart rate, SpO2, NIBP, temperature, respiratory rate). | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, clinical-data-integration, session-173
SUB-SUBSYSTEMREQUIREMENTS-037 | The Medical-Grade Network Switch SHALL forward vital sign data frames between any two ports with a maximum latency of 500 microseconds under full traffic load. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-038 | The Medical-Grade Network Switch SHALL provide dual redundant uplink paths with automatic failover completing within 3 seconds upon detection of link loss. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-039 | The Wireless Access Controller SHALL complete roaming handoff for mobile patient monitors between access points within 100 milliseconds without loss of vital sign data packets. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-040 | The Wireless Access Controller SHALL authenticate all wireless client devices using WPA3-Enterprise with 802.1X certificate-based authentication before granting network access. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-041 | The Time Synchronization Service SHALL maintain clock alignment across all connected monitoring devices to within 1 millisecond of the network master clock using IEEE 1588 Precision Time Protocol. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-042 | The Clinical Network Firewall SHALL enforce TLS 1.2 or higher on all data flows carrying patient health information between the monitoring VLAN and the hospital enterprise network. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-043 | The Clinical Network Firewall SHALL isolate the clinical monitoring network into a dedicated VLAN with access control lists permitting only authorized protocols and endpoints per IEC 80001-1. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-044 | The Message Broker and Data Distribution Service SHALL deliver alarm event messages to all subscribed consumers within 50 milliseconds of publication with guaranteed at-least-once delivery semantics. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-045 | The Message Broker and Data Distribution Service SHALL support topic-based routing with per-patient and per-parameter granularity, enabling subsystems to subscribe to specific vital sign streams without receiving unrelated traffic. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SUB-SUBSYSTEMREQUIREMENTS-046 | When the primary PTP grandmaster clock becomes unavailable, the Time Synchronization Service SHALL promote a secondary clock source and re-synchronize all devices within 10 seconds while maintaining sub-5 millisecond accuracy during the transition. | subsystem-requirements | Subsystem Requirements | [5 empty] | subsystem, network, session-174
SYS-SYSTEM-LEVELREQUIREMENTS-001 | The Hospital Patient Monitoring System SHALL acquire and display continuous ECG waveforms with a minimum of 3-lead configuration at a sampling rate of no less than 250 Hz. | system-requirements | System-Level Requirements | [5 empty] | system, vital-signs, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-002 | The Hospital Patient Monitoring System SHALL measure and display pulse oximetry (SpO2) with an accuracy of +/- 2% in the range of 70-100% saturation. | system-requirements | System-Level Requirements | [5 empty] | system, vital-signs, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-003 | The Hospital Patient Monitoring System SHALL perform non-invasive blood pressure measurements at configurable intervals ranging from 1 minute to 480 minutes and on manual demand. | system-requirements | System-Level Requirements | [5 empty] | system, vital-signs, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-004 | When a monitored vital sign exceeds a configured alarm threshold, the Hospital Patient Monitoring System SHALL generate an audible and visual alarm at the bedside within 10 seconds of threshold violation. | system-requirements | System-Level Requirements | [5 empty] | system, alarm, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-005 | The Hospital Patient Monitoring System SHALL display real-time vital signs and alarm status for all monitored patients on the central monitoring station simultaneously, supporting a minimum of 32 patients per station. | system-requirements | System-Level Requirements | [5 empty] | system, display, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-006 | The Hospital Patient Monitoring System SHALL export patient vital sign data and alarm events to the Electronic Health Record using HL7 FHIR R4 observation resources. | system-requirements | System-Level Requirements | [5 empty] | system, integration, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-007 | The Hospital Patient Monitoring System SHALL automatically associate bedside monitors with patient identity upon receiving ADT admit messages and disassociate upon discharge. | system-requirements | System-Level Requirements | [5 empty] | system, integration, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-008 | The Hospital Patient Monitoring System SHALL support configurable alarm delay periods of 0 to 60 seconds per parameter to suppress transient threshold crossings caused by patient movement or artifact. | system-requirements | System-Level Requirements | [5 empty] | system, alarm, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-009 | The Hospital Patient Monitoring System SHALL encrypt all patient data in transit using TLS 1.2 or higher and at rest using AES-256 encryption. | system-requirements | System-Level Requirements | [5 empty] | system, security, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-010 | The Hospital Patient Monitoring System SHALL maintain continuous monitoring operation during a single network path failure through redundant communication links with automatic failover completing within 5 seconds. | system-requirements | System-Level Requirements | [5 empty] | system, network, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-011 | The Hospital Patient Monitoring System SHALL store a minimum of 72 hours of trending data per patient including all monitored vital sign parameters at one-minute resolution. | system-requirements | System-Level Requirements | [5 empty] | system, storage, session-169
SYS-SYSTEM-LEVELREQUIREMENTS-012 | The Hospital Patient Monitoring System SHALL meet the safety and essential performance requirements of IEC 60601-1-8 for medical electrical equipment alarm systems. | system-requirements | System-Level Requirements | [5 empty] | system, regulatory, session-169
VER-VERIFICATIONMETHODS-001 | The threshold evaluation latency of the Alarm Detection Engine SHALL be verified by test, injecting vital sign samples at maximum acquisition rate and measuring the time from sample receipt to alarm condition event generation, confirming it does not exceed 500 milliseconds across 10000 consecutive samples. | verification-plan | Verification Methods | [5 empty] | verification, alarm-management, session-171
VER-VERIFICATIONMETHODS-002 | The IEC 60601-1-8 priority classification of the Alarm Prioritization Module SHALL be verified by inspection of the alarm priority mapping table and by test using a predefined set of alarm conditions covering all three priority levels, confirming correct classification for each. | verification-plan | Verification Methods | [5 empty] | verification, alarm-management, session-171
VER-VERIFICATIONMETHODS-003 | The crisis alarm bypass behaviour of the Alarm Fatigue Mitigation Unit SHALL be verified by test, configuring maximum delay and suppression settings and then injecting a crisis-priority alarm, confirming the alarm passes to the Notification Controller with zero delay. | verification-plan | Verification Methods | [5 empty] | verification, alarm-management, session-171
VER-VERIFICATIONMETHODS-004 | The tamper-evident audit trail of the Alarm History and Audit Logger SHALL be verified by test, generating a sequence of alarm events across all lifecycle states and confirming that every event is recorded with correct timestamp, patient identifier, and operator identity, and that any post-hoc modification is detectable. | verification-plan | Verification Methods | [5 empty] | verification, alarm-management, session-171
VER-VERIFICATIONMETHODS-005 | The automatic patient-monitor binding upon ADT Admit of the Patient ADT Manager SHALL be verified by integration test using a simulated HL7 ADT^A01 message sequence, measuring the time from message receipt to confirmed binding in the Clinical Data Repository, with pass criteria of binding completion within 2 seconds for 100 consecutive admit events. | verification-plan | Verification Methods | [5 empty] | verification, clinical-data-integration, session-173
VER-VERIFICATIONMETHODS-006 | The LOINC mapping completeness of the Terminology and Coding Service SHALL be verified by inspection of the terminology mapping table against the full set of vital sign measurement types produced by the Vital Signs Acquisition Subsystem, with pass criteria of a valid LOINC code assigned to at least 99% of standard measurement types. | verification-plan | Verification Methods | [5 empty] | verification, clinical-data-integration, session-173
VER-VERIFICATIONMETHODS-007 | The TLS encryption and mutual certificate authentication of the Data Export Controller interface SHALL be verified by penetration test capturing outbound traffic during clinical data export, confirming TLS 1.2 or higher negotiation and mutual certificate exchange, with pass criteria of zero plaintext observation data detected in captured traffic. | verification-plan | Verification Methods | [5 empty] | verification, clinical-data-integration, session-173
VER-VERIFICATIONMETHODS-008 | Verification of network switch latency SHALL be performed by injecting timestamped test frames at line rate across all active ports and measuring one-way delay using hardware timestamping, confirming 99th percentile latency remains below 500 microseconds. | verification-plan | Verification Methods | [5 empty] | verification, network, session-174
VER-VERIFICATIONMETHODS-009 | Verification of network failover SHALL be performed by physically disconnecting the primary uplink during active vital sign streaming and measuring the interval until data delivery resumes on the redundant path, confirming completion within 3 seconds with zero alarm message loss. | verification-plan | Verification Methods | [5 empty] | verification, network, session-174
VER-VERIFICATIONMETHODS-010 | Verification of wireless roaming handoff SHALL be performed by moving a mobile patient monitor through coverage overlap zones at walking speed while streaming continuous vital signs, confirming no data gap exceeds 100 milliseconds via packet capture analysis. | verification-plan | Verification Methods | [5 empty] | verification, network, session-174
VER-VERIFICATIONMETHODS-011 | The ECG signal bandwidth and sampling rate of the ECG Acquisition Module SHALL be verified by test, applying calibrated sine wave signals at 0.05 Hz, 40 Hz, and 150 Hz through a patient simulator and measuring the digitised output amplitude, confirming frequency response within +/- 3 dB across the specified bandwidth and sampling rate of at least 500 Hz. | verification-plan | Verification Methods | [5 empty] | verification, vital-signs-acquisition, session-175
VER-VERIFICATIONMETHODS-012 | The ECG signal bandwidth and sampling rate of the ECG Acquisition Module SHALL be verified by test, applying calibrated sine wave signals at 0.05 Hz, 40 Hz, and 150 Hz through a patient simulator and measuring the digitised output amplitude, confirming frequency response within plus or minus 3 dB across the specified bandwidth and sampling rate of at least 500 Hz. | verification-plan | Verification Methods | [5 empty] | verification, vital-signs-acquisition, session-175
VER-VERIFICATIONMETHODS-013 | The arterial oxygen saturation accuracy of the Pulse Oximetry Module SHALL be verified by test using a multi-wavelength pulse oximeter tester simulating SpO2 levels at 70%, 80%, 90%, 95%, and 100%, confirming readings within 2 percentage points of the reference value and pulse rate accuracy within 3 bpm across the physiological range of 30 to 240 bpm. | verification-plan | Verification Methods | [5 empty] | verification, vital-signs-acquisition, session-175
VER-VERIFICATIONMETHODS-014 | The NIBP measurement accuracy and cuff safety limits of the NIBP Measurement Module SHALL be verified by test using a calibrated pressure source and adult arm simulator, confirming systolic and diastolic accuracy within 5 mmHg for mean arterial pressures of 60 to 200 mmHg, and verifying that cuff pressure never exceeds 300 mmHg and automatic deflation engages within 120 seconds. | verification-plan | Verification Methods | [5 empty] | verification, vital-signs-acquisition, session-175
VER-VERIFICATIONMETHODS-015 | The mains interference rejection of the Signal Conditioning Unit SHALL be verified by test, superimposing 50 Hz and 60 Hz sinusoidal signals at 1 mV peak-to-peak on a calibrated ECG reference signal and measuring residual mains frequency content at the output, confirming attenuation of at least 40 dB for both frequencies across all sensor input channels. | verification-plan | Verification Methods | [5 empty] | verification, vital-signs-acquisition, session-175
VER-VERIFICATIONMETHODS-016 | The waveform rendering performance of the Display Rendering Engine SHALL be verified by test, streaming continuous ECG waveforms to the display while measuring frame update rate and numeric vital sign update latency, confirming the waveform refresh rate meets or exceeds 25 frames per second and numeric readout updates complete within 1 second of source data change. | verification-plan | Verification Methods | [5 empty] | verification, central-monitoring-station, session-175
VER-VERIFICATIONMETHODS-017 | The patient tile capacity and alarm-driven reordering of the Patient Tile Manager SHALL be verified by test, configuring 32 simultaneous patient streams and triggering alarm state transitions on selected patients, confirming all 32 tiles render correctly and that alarm-priority reordering completes within 2 seconds of alarm state change. | verification-plan | Verification Methods | [5 empty] | verification, central-monitoring-station, session-175
VER-VERIFICATIONMETHODS-018 | The trend analysis buffer retention and deterioration detection of the Trend Analysis Engine SHALL be verified by test, injecting 72 hours of simulated vital sign data with a programmed gradual deterioration pattern, confirming trend buffer retention of the full 72-hour window and detection of the deterioration trend within 15 minutes of onset by comparison against the embedded early warning scoring algorithm. | verification-plan | Verification Methods | [5 empty] | verification, central-monitoring-station, session-175
VER-VERIFICATIONMETHODS-019 | The multi-bed overview capacity of the Multi-bed Overview Processor SHALL be verified by test, simulating a full ward of 32 patients with concurrent vital sign streams, alarm conditions, and demographic data updates, confirming that the consolidated overview refreshes within 3 seconds and correctly reflects the alarm status, bed assignment, and latest vitals for every active patient. | verification-plan | Verification Methods | [5 empty] | verification, central-monitoring-station, session-175

## Trace Links
Source Ref | Target Ref | Link Type | Description | Rationale | Created
SUB-SUBSYSTEMREQUIREMENTS-028 | VER-VERIFICATIONMETHODS-019 | verifies |  |  | 2026-03-15T06:05:11.687Z
SUB-SUBSYSTEMREQUIREMENTS-027 | VER-VERIFICATIONMETHODS-019 | verifies |  |  | 2026-03-15T06:05:11.142Z
SUB-SUBSYSTEMREQUIREMENTS-024 | VER-VERIFICATIONMETHODS-018 | verifies |  |  | 2026-03-15T06:05:10.218Z
SUB-SUBSYSTEMREQUIREMENTS-023 | VER-VERIFICATIONMETHODS-018 | verifies |  |  | 2026-03-15T06:05:09.654Z
SUB-SUBSYSTEMREQUIREMENTS-022 | VER-VERIFICATIONMETHODS-017 | verifies |  |  | 2026-03-15T06:05:08.488Z
SUB-SUBSYSTEMREQUIREMENTS-021 | VER-VERIFICATIONMETHODS-017 | verifies |  |  | 2026-03-15T06:05:07.942Z
SUB-SUBSYSTEMREQUIREMENTS-020 | VER-VERIFICATIONMETHODS-016 | verifies |  |  | 2026-03-15T06:05:06.635Z
SUB-SUBSYSTEMREQUIREMENTS-019 | VER-VERIFICATIONMETHODS-016 | verifies |  |  | 2026-03-15T06:05:06.093Z
SUB-SUBSYSTEMREQUIREMENTS-007 | VER-VERIFICATIONMETHODS-015 | verifies |  |  | 2026-03-15T06:05:00.507Z
SUB-SUBSYSTEMREQUIREMENTS-006 | VER-VERIFICATIONMETHODS-014 | verifies |  |  | 2026-03-15T06:04:59.921Z
SUB-SUBSYSTEMREQUIREMENTS-005 | VER-VERIFICATIONMETHODS-014 | verifies |  |  | 2026-03-15T06:04:59.367Z
SUB-SUBSYSTEMREQUIREMENTS-004 | VER-VERIFICATIONMETHODS-013 | verifies |  |  | 2026-03-15T06:04:58.169Z
SUB-SUBSYSTEMREQUIREMENTS-003 | VER-VERIFICATIONMETHODS-013 | verifies |  |  | 2026-03-15T06:04:57.563Z
SUB-SUBSYSTEMREQUIREMENTS-002 | VER-VERIFICATIONMETHODS-012 | verifies |  |  | 2026-03-15T06:04:56.353Z
SUB-SUBSYSTEMREQUIREMENTS-001 | VER-VERIFICATIONMETHODS-012 | verifies |  |  | 2026-03-15T06:04:55.778Z
SUB-SUBSYSTEMREQUIREMENTS-039 | VER-VERIFICATIONMETHODS-010 | verifies |  |  | 2026-03-15T05:06:52.954Z
SUB-SUBSYSTEMREQUIREMENTS-038 | VER-VERIFICATIONMETHODS-009 | verifies |  |  | 2026-03-15T05:06:52.362Z
SUB-SUBSYSTEMREQUIREMENTS-037 | VER-VERIFICATIONMETHODS-008 | verifies |  |  | 2026-03-15T05:06:51.775Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | IFC-INTERFACEDEFINITIONS-018 | derives |  |  | 2026-03-15T05:06:28.200Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | IFC-INTERFACEDEFINITIONS-017 | derives |  |  | 2026-03-15T05:06:27.638Z
SYS-SYSTEM-LEVELREQUIREMENTS-009 | IFC-INTERFACEDEFINITIONS-016 | derives |  |  | 2026-03-15T05:06:27.094Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | IFC-INTERFACEDEFINITIONS-015 | derives |  |  | 2026-03-15T05:06:26.570Z
SYS-SYSTEM-LEVELREQUIREMENTS-010 | SUB-SUBSYSTEMREQUIREMENTS-046 | derives |  |  | 2026-03-15T05:05:59.516Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-045 | derives |  |  | 2026-03-15T05:05:58.904Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | SUB-SUBSYSTEMREQUIREMENTS-044 | derives |  |  | 2026-03-15T05:05:58.331Z
SYS-SYSTEM-LEVELREQUIREMENTS-009 | SUB-SUBSYSTEMREQUIREMENTS-043 | derives |  |  | 2026-03-15T05:05:57.747Z
SYS-SYSTEM-LEVELREQUIREMENTS-009 | SUB-SUBSYSTEMREQUIREMENTS-042 | derives |  |  | 2026-03-15T05:05:57.129Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-041 | derives |  |  | 2026-03-15T05:05:28.591Z
SYS-SYSTEM-LEVELREQUIREMENTS-009 | SUB-SUBSYSTEMREQUIREMENTS-040 | derives |  |  | 2026-03-15T05:05:28.035Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-039 | derives |  |  | 2026-03-15T05:05:27.505Z
SYS-SYSTEM-LEVELREQUIREMENTS-010 | SUB-SUBSYSTEMREQUIREMENTS-038 | derives |  |  | 2026-03-15T05:05:26.962Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-037 | derives |  |  | 2026-03-15T05:05:26.384Z
IFC-INTERFACEDEFINITIONS-014 | VER-VERIFICATIONMETHODS-007 | verifies |  |  | 2026-03-15T04:08:08.256Z
SUB-SUBSYSTEMREQUIREMENTS-036 | VER-VERIFICATIONMETHODS-006 | verifies |  |  | 2026-03-15T04:08:07.701Z
SUB-SUBSYSTEMREQUIREMENTS-031 | VER-VERIFICATIONMETHODS-005 | verifies |  |  | 2026-03-15T04:08:07.136Z
SYS-SYSTEM-LEVELREQUIREMENTS-009 | IFC-INTERFACEDEFINITIONS-014 | derives |  |  | 2026-03-15T04:07:34.289Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | IFC-INTERFACEDEFINITIONS-013 | derives |  |  | 2026-03-15T04:07:33.678Z
SYS-SYSTEM-LEVELREQUIREMENTS-006 | IFC-INTERFACEDEFINITIONS-012 | derives |  |  | 2026-03-15T04:07:33.100Z
SYS-SYSTEM-LEVELREQUIREMENTS-006 | SUB-SUBSYSTEMREQUIREMENTS-036 | derives |  |  | 2026-03-15T04:06:56.418Z
SYS-SYSTEM-LEVELREQUIREMENTS-006 | SUB-SUBSYSTEMREQUIREMENTS-035 | derives |  |  | 2026-03-15T04:06:55.836Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | SUB-SUBSYSTEMREQUIREMENTS-034 | derives |  |  | 2026-03-15T04:06:55.284Z
SYS-SYSTEM-LEVELREQUIREMENTS-011 | SUB-SUBSYSTEMREQUIREMENTS-033 | derives |  |  | 2026-03-15T04:06:54.745Z
SYS-SYSTEM-LEVELREQUIREMENTS-007 | SUB-SUBSYSTEMREQUIREMENTS-032 | derives |  |  | 2026-03-15T04:06:54.203Z
SYS-SYSTEM-LEVELREQUIREMENTS-007 | SUB-SUBSYSTEMREQUIREMENTS-031 | derives |  |  | 2026-03-15T04:06:53.648Z
SYS-SYSTEM-LEVELREQUIREMENTS-006 | SUB-SUBSYSTEMREQUIREMENTS-030 | derives |  |  | 2026-03-15T04:06:53.069Z
SYS-SYSTEM-LEVELREQUIREMENTS-006 | SUB-SUBSYSTEMREQUIREMENTS-029 | derives |  |  | 2026-03-15T04:06:52.543Z
SYS-SYSTEM-LEVELREQUIREMENTS-008 | IFC-INTERFACEDEFINITIONS-011 | derives |  |  | 2026-03-15T03:07:11.450Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | IFC-INTERFACEDEFINITIONS-010 | derives |  |  | 2026-03-15T03:07:10.885Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | IFC-INTERFACEDEFINITIONS-009 | derives |  |  | 2026-03-15T03:07:10.314Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-028 | derives |  |  | 2026-03-15T03:06:27.129Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-027 | derives |  |  | 2026-03-15T03:06:26.575Z
SYS-SYSTEM-LEVELREQUIREMENTS-008 | SUB-SUBSYSTEMREQUIREMENTS-026 | derives |  |  | 2026-03-15T03:06:26.020Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | SUB-SUBSYSTEMREQUIREMENTS-025 | derives |  |  | 2026-03-15T03:06:25.452Z
SYS-SYSTEM-LEVELREQUIREMENTS-011 | SUB-SUBSYSTEMREQUIREMENTS-024 | derives |  |  | 2026-03-15T03:06:24.897Z
SYS-SYSTEM-LEVELREQUIREMENTS-011 | SUB-SUBSYSTEMREQUIREMENTS-023 | derives |  |  | 2026-03-15T03:06:24.358Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-022 | derives |  |  | 2026-03-15T03:06:23.774Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-021 | derives |  |  | 2026-03-15T03:06:23.241Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-020 | derives |  |  | 2026-03-15T03:06:22.676Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | SUB-SUBSYSTEMREQUIREMENTS-019 | derives |  |  | 2026-03-15T03:06:22.102Z
IFC-INTERFACEDEFINITIONS-005 | VER-VERIFICATIONMETHODS-001 | verifies |  |  | 2026-03-15T02:07:02.585Z
SUB-SUBSYSTEMREQUIREMENTS-018 | VER-VERIFICATIONMETHODS-004 | verifies |  |  | 2026-03-15T02:07:01.960Z
SUB-SUBSYSTEMREQUIREMENTS-014 | VER-VERIFICATIONMETHODS-003 | verifies |  |  | 2026-03-15T02:07:01.375Z
SUB-SUBSYSTEMREQUIREMENTS-011 | VER-VERIFICATIONMETHODS-002 | verifies |  |  | 2026-03-15T02:07:00.768Z
SUB-SUBSYSTEMREQUIREMENTS-009 | VER-VERIFICATIONMETHODS-001 | verifies |  |  | 2026-03-15T02:07:00.185Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | IFC-INTERFACEDEFINITIONS-008 | derives |  |  | 2026-03-15T02:06:25.688Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | IFC-INTERFACEDEFINITIONS-007 | derives |  |  | 2026-03-15T02:06:25.130Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | IFC-INTERFACEDEFINITIONS-006 | derives |  |  | 2026-03-15T02:06:24.564Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | IFC-INTERFACEDEFINITIONS-005 | derives |  |  | 2026-03-15T02:06:24.011Z
SYS-SYSTEM-LEVELREQUIREMENTS-011 | SUB-SUBSYSTEMREQUIREMENTS-018 | derives |  |  | 2026-03-15T02:05:49.267Z
SYS-SYSTEM-LEVELREQUIREMENTS-012 | SUB-SUBSYSTEMREQUIREMENTS-018 | derives |  |  | 2026-03-15T02:05:48.697Z
SYS-SYSTEM-LEVELREQUIREMENTS-012 | SUB-SUBSYSTEMREQUIREMENTS-014 | derives |  |  | 2026-03-15T02:05:48.119Z
SYS-SYSTEM-LEVELREQUIREMENTS-012 | SUB-SUBSYSTEMREQUIREMENTS-011 | derives |  |  | 2026-03-15T02:05:47.567Z
SYS-SYSTEM-LEVELREQUIREMENTS-008 | SUB-SUBSYSTEMREQUIREMENTS-014 | derives |  |  | 2026-03-15T02:05:46.959Z
SYS-SYSTEM-LEVELREQUIREMENTS-008 | SUB-SUBSYSTEMREQUIREMENTS-013 | derives |  |  | 2026-03-15T02:05:46.378Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | SUB-SUBSYSTEMREQUIREMENTS-015 | derives |  |  | 2026-03-15T02:05:45.836Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | SUB-SUBSYSTEMREQUIREMENTS-011 | derives |  |  | 2026-03-15T02:05:45.254Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | SUB-SUBSYSTEMREQUIREMENTS-010 | derives |  |  | 2026-03-15T02:05:44.741Z
SYS-SYSTEM-LEVELREQUIREMENTS-004 | SUB-SUBSYSTEMREQUIREMENTS-009 | derives |  |  | 2026-03-15T02:05:44.160Z
SYS-SYSTEM-LEVELREQUIREMENTS-005 | IFC-INTERFACEDEFINITIONS-004 | derives |  |  | 2026-03-15T01:06:06.611Z
SYS-SYSTEM-LEVELREQUIREMENTS-003 | IFC-INTERFACEDEFINITIONS-003 | derives |  |  | 2026-03-15T01:06:06.070Z
SYS-SYSTEM-LEVELREQUIREMENTS-002 | IFC-INTERFACEDEFINITIONS-002 | derives |  |  | 2026-03-15T01:06:05.482Z
SYS-SYSTEM-LEVELREQUIREMENTS-001 | IFC-INTERFACEDEFINITIONS-001 | derives |  |  | 2026-03-15T01:06:04.961Z
SYS-SYSTEM-LEVELREQUIREMENTS-010 | SUB-SUBSYSTEMREQUIREMENTS-008 | derives |  |  | 2026-03-15T01:05:32.475Z
SYS-SYSTEM-LEVELREQUIREMENTS-012 | SUB-SUBSYSTEMREQUIREMENTS-007 | derives |  |  | 2026-03-15T01:05:31.928Z
SYS-SYSTEM-LEVELREQUIREMENTS-012 | SUB-SUBSYSTEMREQUIREMENTS-006 | derives |  |  | 2026-03-15T01:05:31.392Z
SYS-SYSTEM-LEVELREQUIREMENTS-003 | SUB-SUBSYSTEMREQUIREMENTS-005 | derives |  |  | 2026-03-15T01:05:30.859Z
SYS-SYSTEM-LEVELREQUIREMENTS-002 | SUB-SUBSYSTEMREQUIREMENTS-004 | derives |  |  | 2026-03-15T01:05:30.310Z
SYS-SYSTEM-LEVELREQUIREMENTS-002 | SUB-SUBSYSTEMREQUIREMENTS-003 | derives |  |  | 2026-03-15T01:05:29.771Z
SYS-SYSTEM-LEVELREQUIREMENTS-001 | SUB-SUBSYSTEMREQUIREMENTS-002 | derives |  |  | 2026-03-15T01:05:29.139Z
SYS-SYSTEM-LEVELREQUIREMENTS-001 | SUB-SUBSYSTEMREQUIREMENTS-001 | derives |  |  | 2026-03-15T01:05:28.376Z
STK-STAKEHOLDERNEEDS-006 | SYS-SYSTEM-LEVELREQUIREMENTS-009 | derives |  |  | 2026-03-15T00:06:55.680Z
STK-STAKEHOLDERNEEDS-005 | SYS-SYSTEM-LEVELREQUIREMENTS-008 | derives |  |  | 2026-03-15T00:06:55.116Z
STK-STAKEHOLDERNEEDS-004 | SYS-SYSTEM-LEVELREQUIREMENTS-012 | derives |  |  | 2026-03-15T00:06:54.548Z
STK-STAKEHOLDERNEEDS-003 | SYS-SYSTEM-LEVELREQUIREMENTS-011 | derives |  |  | 2026-03-15T00:06:53.961Z
STK-STAKEHOLDERNEEDS-003 | SYS-SYSTEM-LEVELREQUIREMENTS-007 | derives |  |  | 2026-03-15T00:06:53.390Z
STK-STAKEHOLDERNEEDS-003 | SYS-SYSTEM-LEVELREQUIREMENTS-006 | derives |  |  | 2026-03-15T00:06:52.847Z
STK-STAKEHOLDERNEEDS-002 | SYS-SYSTEM-LEVELREQUIREMENTS-005 | derives |  |  | 2026-03-15T00:06:52.298Z
STK-STAKEHOLDERNEEDS-002 | SYS-SYSTEM-LEVELREQUIREMENTS-004 | derives |  |  | 2026-03-15T00:06:51.744Z
STK-STAKEHOLDERNEEDS-001 | SYS-SYSTEM-LEVELREQUIREMENTS-010 | derives |  |  | 2026-03-15T00:06:51.193Z
STK-STAKEHOLDERNEEDS-001 | SYS-SYSTEM-LEVELREQUIREMENTS-003 | derives |  |  | 2026-03-15T00:06:50.618Z
STK-STAKEHOLDERNEEDS-001 | SYS-SYSTEM-LEVELREQUIREMENTS-002 | derives |  |  | 2026-03-15T00:06:50.023Z
STK-STAKEHOLDERNEEDS-001 | SYS-SYSTEM-LEVELREQUIREMENTS-001 | derives |  |  | 2026-03-15T00:06:49.439Z

## Linksets
Source Document | Target Document | Default Link Type | Link Count
interface-requirements | verification-plan |  | 2
subsystem-requirements | verification-plan |  | 24
system-requirements | interface-requirements |  | 18
system-requirements | subsystem-requirements |  | 46
stakeholder-requirements | system-requirements |  | 12

## Sections
Document | Section | Short Code | Order | Requirements
architecture-decisions | Architecture Decisions |  | 0 | 0
interface-requirements | Interface Definitions |  | 0 | 18
stakeholder-requirements | Stakeholder Needs |  | 0 | 6
subsystem-requirements | Subsystem Requirements |  | 0 | 46
system-requirements | System-Level Requirements |  | 0 | 12
verification-plan | Verification Methods |  | 0 | 19