## Documents
Slug | Name | Short Code | Sections | Requirements | Created
architecture-decisions | Architecture Decisions | ARC | 1 | 14 | 2026-03-19T10:41:32.697Z
interface-requirements | Interface Requirements | IFC | 1 | 55 | 2026-03-19T10:41:32.460Z
stakeholder-requirements | Stakeholder Requirements | STK | 1 | 7 | 2026-03-19T10:41:32.021Z
subsystem-requirements | Subsystem Requirements | SUB | 1 | 100 | 2026-03-19T10:41:32.577Z
system-requirements | System Requirements | SYS | 1 | 12 | 2026-03-19T10:41:32.329Z
verification-requirements | Verification Requirements | VER | 1 | 62 | 2026-03-19T10:41:32.796Z

## Requirements
Ref | Text | Document | Section | Pattern | Verification | QA Score | Compliance | Rationale | Tags
ARC-REQ-001 | ARC: Active Ventilation and Containment System — Twin-bank HEPA with safe-change housings and depression cascade architecture selected over single-bank filtration or positive-pressure glovebox approach. The twin-bank arrangement satisfies defence-in-depth requirements under Nuclear Site Licence Condition 28 (regular and systematic review) by ensuring that a single filter failure or breach does not result in uncontrolled release. Depression cascade (negative pressure increasing toward highest activity zones) was chosen over local containment-only because the laboratory handles diverse alpha-emitting materials across multiple workstations — centralised extract through a cascade provides whole-building containment assurance rather than relying on individual device integrity. Seven-component decomposition: Supply AHU, Extract Fan System, HEPA Filtration Assembly, Depression Cascade Control System, Stack Monitoring and Discharge System, Fume Cupboard and Glove Box Extract Network, and Standby Ventilation and Emergency Isolation System. The control system uses 1oo2D PLC architecture because cascade control is a safety function (SIL 2) — loss of depression cascade during alpha handling could result in personnel inhalation doses exceeding 20 mSv. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Documents the key design trade-offs for the primary engineered safety barrier. Twin-bank HEPA and depression cascade are standard nuclear practice (IAEA Safety Guide RS-G-1.7, ONR TAG 049) but the specific rationale for this facility — multiple alpha-handling workstations with varying activity levels — justifies recording the engineering judgement. | architecture, ventilation, session-326
ARC-REQ-003 | ARC: Radiation Protection and Health Physics Monitoring System — Centralised star topology with dedicated HP Central Alarm and Display System as single aggregation point for all monitoring instruments. Chose centralised SCADA-grade architecture over distributed peer-to-peer because: (a) ONR SAP EKP.3 requires demonstrable single-point alarm awareness for the Radiation Protection Supervisor; (b) SIL 1 alarm function classification per IEC 61508 requires defined reliability architecture with known common-cause failure modes; (c) 10-year historian with statutory returns generation requires single database rather than federated stores. Dual-server active-standby with 4-hour UPS eliminates single points of failure at the aggregation layer while maintaining the simplicity advantage of star topology. Component-level instruments retain local alarm capability as defence-in-depth — loss of central system degrades trending and recording but does not eliminate local audible/visual warnings. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Nuclear site licence condition 26 (control and supervision of operations) and ONR SAP EKP.3 (single competent person awareness) mandate centralised visibility of all radiation monitoring. Star topology with local alarm fallback satisfies both the centralisation requirement and the defence-in-depth principle. | architecture, radiation-protection, session-327
ARC-REQ-004 | ARC: Radiation Protection and HP Monitoring — Centralised star topology with HP Central Alarm and Display as single aggregation point. Chose centralised SCADA over distributed peer-to-peer because ONR SAP EKP.3 requires single-point alarm awareness for RPS, SIL 1 alarm classification requires defined reliability architecture, and statutory returns need single historian. Dual-server active-standby with 4-hour UPS at aggregation layer. Instruments retain local alarm as defence-in-depth. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | LC26 and ONR SAP EKP.3 mandate centralised radiation monitoring visibility. Star topology with local alarm fallback satisfies centralisation and defence-in-depth. | architecture, radiation-protection, session-327
ARC-REQ-005 | ARC: Active Effluent Treatment and Discharge System — Batch-mode processing with four delay tanks in a fill-sample-hold-discharge cycle, rather than continuous flow-through treatment. This topology ensures every batch is characterised before discharge, satisfying Environment Agency permit conditions requiring pre-discharge analysis. Chemical treatment (floc co-precipitation + ion exchange) was chosen over evaporation due to lower capital cost and simpler ILW sludge management, accepting higher per-batch reagent costs. The Discharge Monitoring and Control System is designated SIL 2 because an uncontrolled release of above-limit effluent is a credible environmental hazard, and the single discharge point allows a simple fail-safe isolation architecture. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Batch processing provides a hold point for regulatory compliance that continuous systems cannot. SIL 2 assignment follows IEC 61511 risk graph for environmental consequence category. | architecture, aetds, session-328
ARC-REQ-006 | ARC: Facility Safety and Emergency Response System — Six-component architecture separating criticality warning, fire detection, safety interlocks, emergency communications, emergency power, and spill containment into independent subsystems with dedicated power paths. The criticality warning and safety interlock systems are hardwired (relay-based) rather than PLC-based, ensuring deterministic response times and immunity to common-cause software failures. The emergency power system uses a three-tier architecture (UPS, diesel, dockyard ring main) because loss of safety monitoring power in a facility handling submarine-origin fissile materials would create an unacceptable criticality risk gap. Fire suppression uses zone-differentiated media (water mist vs IG-541 inert gas) because water application in alpha-contaminated areas would spread plutonium contamination, converting a fire event into a combined fire-contamination incident. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Architecture driven by ONR SAPs (EHA.7 fire, ECS.3 criticality), IEC 61508/61511 safety integrity requirements, and the specific hazard profile of a nuclear dockyard radiochemistry laboratory handling Pu-239 and U-235. Independence between subsystems prevents common-cause failure cascading across safety functions. | architecture, safety, session-329
ARC-REQ-008 | ARC: Radioactive Waste Management Facility — Separated solid and liquid waste streams with centralised records. The decomposition separates characterization from packaging because waste categorisation (VLLW/LLW/ILW) must be completed and recorded BEFORE packaging decisions are made — mixing these functions would create a single point of failure in the waste acceptance criteria compliance chain. Liquid waste conditioning is isolated because alpha-bearing liquors require shielded cementation in a different containment zone from dry solid waste handling. The Interim Waste Store is a distinct component rather than an extension of the packaging area because criticality-safe storage geometry, environmental monitoring, and multi-decade package surveillance require dedicated infrastructure and separate safety case arguments. The Waste Records and Consignment System is centralised across all waste streams to ensure a single authoritative nuclear material accountancy record, avoiding reconciliation errors between parallel databases. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Nuclear waste management at a dockyard radiochemistry lab must comply with Environmental Permitting Regulations, Radioactive Substances Regulation, and ONR Licence Conditions (LC32 accumulation, LC33 disposal). The separation of characterization, conditioning, packaging, storage, and records reflects regulatory requirements for independent verification at each stage of the waste lifecycle, and the need for separate safety case arguments for storage criticality and liquid waste immobilisation. | architecture, waste-management, session-330
ARC-REQ-009 | ARC: Sample Receipt — Linear workflow with screening gate. Samples follow a strict one-way path: Receiving Bay to Screening to Registration to Preparation to Storage/Analysis. The screening station is placed BEFORE registration to prevent uncharacterised high-activity samples from entering the preparation lab. Alternative considered: parallel screening and registration — rejected because dose rate category determines required handling procedures and must be known before LIMS work order generation. The chain-of-custody system is logically separate from LIMS to allow standalone operation if LIMS is unavailable, with batch synchronisation when connectivity is restored. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Architectural rationale for the sequential screening-first workflow in the Sample Receipt facility. | architecture, sample-receipt, session-331
ARC-REQ-010 | ARC: Radiochemical Separations Laboratory — Analyte-driven workflow topology. The laboratory is organised by analyte group rather than by technique, with dedicated stations for actinides, fission products (Sr/Cs), and volatile radionuclides (H-3/C-14). This topology minimises cross-contamination between high-alpha actinide work and low-activity fission product separations, and allows independent ventilation risk assessment per station. Alternative considered: technique-based layout (all precipitation in one area, all chromatography in another) — rejected because different analyte groups require different containment levels and waste streams, and mixing actinide and fission product work at shared stations creates unacceptable cross-contamination risk for low-level measurements. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Records the fundamental architectural choice of analyte-group organisation for the separations laboratory. | architecture, radchem-sep, session-331
ARC-REQ-011 | ARC: Gamma Spectrometry Suite — Electric cryocoolers selected over LN2 dewars despite higher unit cost and slightly higher microphonic noise floor. LN2 handling in a nuclear-licensed facility introduces manual handling risks, supply chain dependency, and oxygen displacement hazards in confined counting rooms. Cryocoolers eliminate liquid cryogen inventory, reduce operator dose from routine dewar filling near active samples, and enable unattended 24/7 operation essential for high-throughput submarine refit sample campaigns. Vibration isolation mounts and digital signal processing with anti-microphonic filtering mitigate the noise penalty to <0.1 keV degradation at 1332 keV. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Nuclear safety and operational availability drove the LN2 elimination decision. Dewar filling near active samples was the highest routine dose contributor in legacy gamma spec labs. Electric cryocoolers with vibration isolation achieve comparable spectral performance while eliminating cryogen logistics. | architecture, gamma-spectrometry, session-333
ARC-REQ-013 | ARC: Alpha Spectrometry Laboratory — Dual source preparation methods retained (electrodeposition and NdF3 microprecipitation) despite duplication of equipment and training burden. Electrodeposition produces higher-quality thin sources (better resolution, lower self-absorption) essential for resolving Pu-238/Am-241 peak overlap and U-234/U-235 separation. NdF3 microprecipitation provides a rapid 30-minute preparation for urgent submarine defuelling samples where 90-minute electrodeposition turnaround is unacceptable. Retaining both methods provides operational resilience and covers the full range of analytical urgency encountered in dockyard operations. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Single-method labs are vulnerable to method failure halting all alpha analysis. The dockyard has routine and urgent analysis streams requiring different source preparation turnaround times. Cost of maintaining dual capability is offset by avoiding analytical delays during submarine refit critical path. | architecture, alpha-spectrometry, session-333
ARC-REQ-014 | ARC: Liquid Scintillation Counting Facility — Ultra-low-background counter selection with pulse shape analysis alpha/beta discrimination chosen over gas proportional counting. Justification: LS counting provides superior sensitivity for H-3 (18.6 keV beta max) with lower minimum detectable activity, handles aqueous reactor coolant matrices directly after distillation, and modern PSA achieves alpha rejection factors exceeding 99.95%. Gas proportional counting was rejected due to higher background rates, requirement for thin-source preparation, and poorer H-3 sensitivity. Dual-counter configuration selected for redundancy and throughput during high-activity dockyard campaigns. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Counter technology selection drives detection limits, sample throughput, and operating costs. Ultra-low-background LSC with PSA is the current industry standard for environmental H-3 and C-14 at nuclear sites per EA RIFE methodology. | architecture, liquid-scintillation, session-334
ARC-REQ-015 | ARC: ICP-MS and Elemental Analysis Suite — Quadrupole ICP-MS with collision/reaction cell (CRC) selected over sector-field (SF-ICP-MS) for routine actinide analysis. While SF-ICP-MS offers superior abundance sensitivity for U-236/U-238 and Pu-240/Pu-239 measurements, the CRC quadrupole handles routine Tc-99, Np-237, and total actinide screening at sufficient sensitivity (sub-pg/L) at approximately one-third the capital and maintenance cost. Helium KED mode resolves ArCa+ on Sr-80 and UH+ on Pu-239 interferences adequately for regulatory compliance measurements. SF-ICP-MS capability can be procured as a service for isotope dilution measurements requiring abundance sensitivity exceeding 1E7. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | Technology selection drives capital cost, maintenance burden, and analytical capability. The CRC quadrupole meets EA environmental monitoring requirements for Tc-99 at 0.05 Bq/L MDA while costing significantly less than sector-field instrumentation. This is a proportionate choice for a dockyard lab that primarily needs screening and confirmation rather than research-grade isotope ratios. | architecture, icp-ms, session-334
ARC-REQ-016 | ARC: Laboratory Information Management System — Centralised LIMS with instrument interface middleware selected over distributed spreadsheet-based tracking. The LIMS provides single source of truth for sample chain-of-custody, eliminates transcription errors in result reporting, enables automated regulatory discharge calculations, and satisfies ONR SYAPS requirements for software used in nuclear safety-related measurement. Active-passive database failover selected over active-active clustering to reduce complexity while meeting 99.5% availability target. | architecture-decisions | Architecture Decisions |  | Analysis |  |  | A nuclear dockyard radiochemistry laboratory handling thousands of samples annually requires automated data management to prevent transcription errors that could affect discharge compliance reporting. UKAS accreditation (ISO 17025) mandates traceable data management. ONR expects LIMS for nuclear site laboratories. The failover architecture trades maximum availability for operational simplicity appropriate to a laboratory (not a real-time control system). | architecture, lims, session-334
ARC-REQ-017 | ARC: Laboratory Utilities and Building Services — decomposed into six components: Electrical Power Distribution, Laboratory Gas Supply, DI Water, BMS, Physical Security, and Chilled Water/Heating. The BMS acts as the central supervisory node, monitoring all other utility components. Power distribution is separated from the Emergency Power System (under Facility Safety) because normal supply reliability requirements differ from safety-classified emergency supply — normal distribution is availability-driven (99.5% uptime), while emergency power is safety-integrity-driven (SIL 2). Physical Security is placed here rather than under Safety because NISR compliance is an access governance concern, not a process safety function. Gas supply is centralised rather than per-laboratory to enable bulk storage, manifold changeover, and consistent pressure regulation, reducing per-instrument gas handling complexity. | architecture-decisions | Architecture Decisions |  | Inspection |  |  | Separation of normal utilities from safety systems follows ONR Safety Assessment Principles expectation that safety-classified systems remain independent of general building services. BMS centralisation enables single-point monitoring required by nuclear site licence conditions for environmental compliance reporting. | architecture, lab-utilities, session-335
IFC-REQ-001 | The interface between the Extract Fan System and the HEPA Filtration Assembly SHALL comprise 316L stainless steel ductwork rated for minus 3 kPa, with the fan positioned downstream of both filter banks to ensure the entire extract path from containment devices to fan inlet operates at negative pressure relative to surrounding spaces. | interface-requirements | Interface Requirements |  | Inspection |  |  | Fan-downstream-of-filters is a fundamental nuclear ventilation design principle. If the fan were upstream, any ductwork leak between fan and filters would release unfiltered contaminated air. With fan downstream, all leaks are inward, maintaining containment. 316L stainless steel is specified because extract ductwork may accumulate surface contamination requiring periodic decontamination with acidic solutions. | interface, ventilation, session-326, idempotency:ifc-efs-hepa-326
IFC-REQ-002 | The interface between the Depression Cascade Control System and the Extract Fan System SHALL provide variable speed drive setpoint signals (4-20 mA analogue or Modbus RTU) with fan speed feedback and status returned at update rates not exceeding 500 milliseconds. | interface-requirements | Interface Requirements |  | Test |  |  | The 500 ms update rate supports the 2-second disturbance response in SUB-REQ-004. The control loop needs at least 4 measurement-actuation cycles within the response window. Dual interface option provides resilience against communication bus failure. | interface, ventilation, session-326, idempotency:ifc-dccs-efs-326
IFC-REQ-003 | The interface between the Fume Cupboard and Glove Box Extract Network and the HEPA Filtration Assembly SHALL comprise a common extract manifold with total airflow capacity of not less than 12000 cubic metres per hour and static pressure at manifold entry of not less than minus 500 Pa. | interface-requirements | Interface Requirements |  | Test |  |  | The manifold must accommodate simultaneous extract from all fume cupboards (8-12 at 500-800 m3/h each) plus glove boxes (4-6 at 50-100 m3/h). The 12000 m3/h capacity provides margin for future expansion. The minus 500 Pa ensures adequate flow to the most remote containment device. | interface, ventilation, session-326, idempotency:ifc-fcgb-hepa-326
IFC-REQ-004 | The interface between the HEPA Filtration Assembly and the Stack Monitoring and Discharge System SHALL comprise the filtered extract duct with an isokinetic sampling point located at not less than 8 duct diameters downstream of the last flow disturbance and 2 duct diameters upstream of the stack exit, per ISO 2889. | interface-requirements | Interface Requirements |  | Inspection |  |  | ISO 2889 specifies isokinetic sampling locations for representative radioactive effluent monitoring. Inadequate straight-run distances cause non-representative sampling due to particle concentration gradients, resulting in either under-reporting (regulatory non-compliance) or over-reporting (spurious stack isolation). | interface, ventilation, session-326, idempotency:ifc-hepa-stack-326
IFC-REQ-005 | The interface between the Depression Cascade Control System and the Supply Air Handling Unit SHALL provide supply fan speed setpoint and supply damper position commands, with supply airflow measurement feedback (pitot array or orifice plate, accuracy plus or minus 5 percent of reading) to enable cascade balancing between supply and extract volumes. | interface-requirements | Interface Requirements |  | Test |  |  | Cascade control requires coordinated supply and extract flow modulation. If supply volume is not reduced proportionally when extract is reduced (e.g. during fume cupboard closure), over-pressurisation of C1/C2 zones can result. The 5 percent accuracy ensures the supply-extract flow balance remains within the control system deadband. | interface, ventilation, session-326, idempotency:ifc-dccs-sahu-326
IFC-REQ-006 | The interface between the Standby Ventilation and Emergency Isolation System and the Depression Cascade Control System SHALL provide hardwired override signals for fire damper closure and smoke extract mode activation, bypassing the normal cascade control logic via dedicated SIL 2 rated safety relay circuits independent of the PLC. | interface-requirements | Interface Requirements |  | Test |  |  | Emergency isolation must operate independently of the cascade PLC to satisfy common-cause failure requirements under IEC 61511. A PLC failure must not prevent fire damper closure or smoke extract activation. Hardwired safety relays provide the required independence and achieve the SIL 2 target for the containment preservation safety function. | interface, ventilation, safety, session-326, idempotency:ifc-standby-dccs-326
IFC-REQ-007 | The interface between the Continuous Air Monitor Network and the Health Physics Central Alarm and Display System SHALL comprise RS-485 Modbus RTU at 9600 baud with polling interval not exceeding 5 seconds, transmitting DAC-fraction reading, alarm state, instrument status, and filter paper accumulation for each CAM station. | interface-requirements | Interface Requirements |  | Test |  |  | RS-485 Modbus is the established industrial standard for radiation monitoring instruments and provides reliable multi-drop communication over distances up to 1200m. 5-second polling ensures the central display updates within the 2-second alarm annunciation window when combined with local alarm relay. Data fields cover all information needed for HP assessment and statutory reporting. | interface, radiation-protection, cam-network, hp-central, session-327, idempotency:ifc-cam-hpcentral-327
IFC-REQ-008 | The interface between the Area Gamma Dose Rate Monitoring Array and the Health Physics Central Alarm and Display System SHALL provide dual-path data: a 4-20mA analogue signal proportional to log dose rate for hardwired alarm relay, and Modbus TCP/IP digital data at 10-second intervals carrying dose rate, energy-compensated H*(10), cumulative dose, and instrument health status. | interface-requirements | Interface Requirements |  | Test |  |  | Dual-path interface provides diversity against common-mode communication failure. Hardwired 4-20mA ensures alarm relay even if digital network fails — this is the SIL 1 safety function path. Digital path carries the full data set needed for trending, statutory returns, and dose assessment. 10-second interval matches the area gamma update rate requirement. | interface, radiation-protection, gamma-array, hp-central, session-327, idempotency:ifc-gamma-hpcentral-327
IFC-REQ-009 | The interface between the Contamination Monitoring Stations and the Health Physics Central Alarm and Display System SHALL transmit measurement results (pass/fail status, alpha and beta-gamma surface activity in Bq/cm2, measurement timestamp, and monitor identity) via Ethernet TCP/IP within 2 seconds of measurement completion. | interface-requirements | Interface Requirements |  | Test |  |  | Real-time contamination monitoring results at the central display enable the RPS to track contamination trends across the facility and identify developing contamination events before they reach clearance levels. 2-second transmission requirement ensures the mimic display reflects current facility status for shift handover and incident response. | interface, radiation-protection, contamination, hp-central, session-327, idempotency:ifc-contam-hpcentral-327
IFC-REQ-010 | The interface between the Personal Dosimetry and Dose Record System and the Health Physics Central Alarm and Display System SHALL transmit EPD dose data (current dose rate, integrated dose, dose constraint percentage, and personnel identifier) via OPC-UA at 60-second intervals, with immediate event notification when any dose alarm threshold is exceeded. | interface-requirements | Interface Requirements |  | Test |  |  | OPC-UA interface aligns with the HP Central system SCADA architecture and provides structured, secure data exchange. 60-second routine interval is adequate for dose trend monitoring. Immediate event notification ensures the RPS is aware of dose alarm conditions in real-time for personnel protection decisions. Personnel identifier enables dose tracking by individual on the mimic display. | interface, radiation-protection, dosimetry, hp-central, session-327, idempotency:ifc-dosim-hpcentral-327
IFC-REQ-011 | The interface between the Health Physics Central Alarm and Display System and the Laboratory Information Management System SHALL provide OPC-UA read access to real-time zone radiation status, current area dose rates, and airborne contamination levels, enabling LIMS to associate radiological conditions with sample processing records. | interface-requirements | Interface Requirements |  | Test |  |  | LIMS integration enables correlation of radiological conditions with sample analysis activities, supporting dose assessment for specific operations and investigation of any anomalous results that may be attributable to elevated background. OPC-UA read-only access ensures LIMS cannot inadvertently modify radiation monitoring data or alarm states. | interface, radiation-protection, hp-central, lims, session-327, idempotency:ifc-hpcentral-lims-327
IFC-REQ-012 | The interface between the Health Physics Central Alarm and Display System and the Facility Safety and Emergency Response System SHALL provide hardwired relay contacts for evacuation-level radiation alarms from any CAM or area gamma monitor, activating the facility evacuation alarm within 1 second of the radiation alarm condition being confirmed. | interface-requirements | Interface Requirements |  | Test |  |  | Hardwired relay interface for evacuation-level alarms provides deterministic, network-independent triggering of the facility emergency response. 1-second activation ensures evacuation alarm sounds within the 3-second total response target from radiation detection to audible warning. This is the most safety-critical interface in the RP subsystem — it must not depend on software or network availability. | interface, radiation-protection, hp-central, facility-safety, session-327, idempotency:ifc-hpcentral-safety-327
IFC-REQ-013 | The interface between the Active Drain Collection Network and the Effluent Collection and Delay Tanks SHALL transfer liquid waste via gravity and sump pump through 50mm bore 316L stainless steel pipework at a maximum flow rate of 20 L/min, with isolation valves at each tank inlet and flow totalisation at the collection header. | interface-requirements | Interface Requirements |  | Test |  |  | 50mm bore sized for peak laboratory waste generation rate (multiple fume cupboard drains simultaneously) with margin. Flow totalisation needed for mass balance accounting per LC34 arrangements. | interface, aetds, session-328, idempotency:ifc-drain-tanks-328
IFC-REQ-014 | The interface between the Effluent Collection and Delay Tanks and the Batch Sampling and Analysis Station SHALL provide dedicated 12mm bore 316L stainless steel sample lines from each tank with peristaltic pump sampling capability, including a 5-minute line purge cycle before each sample draw to clear stagnant liquid from the sample line. | interface-requirements | Interface Requirements |  | Test |  |  | Dedicated sample lines per tank prevent cross-contamination between batches. 5-minute purge at 12mm bore clears approximately 3 line volumes at typical pump rate, ensuring sample represents current tank contents. | interface, aetds, session-328, idempotency:ifc-tanks-sampling-328
IFC-REQ-015 | The interface between the Effluent Collection and Delay Tanks and the Chemical Treatment Plant SHALL transfer effluent batches at a controlled flow rate of 5-15 L/min via positive displacement pump, with return of treated effluent to a clean receiving tank for post-treatment sampling before discharge authorisation. | interface-requirements | Interface Requirements |  | Test |  |  | Controlled flow rate ensures adequate residence time in precipitation reactor and ion exchange columns. Return to receiving tank enables post-treatment verification sampling — if treatment is inadequate, the batch can be re-treated rather than discharged. | interface, aetds, session-328, idempotency:ifc-tanks-treatment-328
IFC-REQ-016 | The interface between the Discharge Monitoring and Control System and the site active drainage system SHALL include continuous proportional sampling during discharge, with in-line ZnS scintillation alpha monitor (MDA 0.05 Bq/mL) and plastic scintillation beta-gamma monitor (MDA 0.5 Bq/mL), electromagnetic flowmeter (±1% accuracy), and a fail-safe pneumatic discharge valve with less than 5-second closure time. | interface-requirements | Interface Requirements |  | Test |  |  | Continuous proportional sampling provides a composite sample for post-discharge laboratory confirmation. MDA values set at half the Environment Agency discharge limits to ensure reliable detection before limits are reached. Pneumatic valve chosen for fail-safe closure on air supply loss. | interface, aetds, safety, session-328, idempotency:ifc-discharge-site-328
IFC-REQ-017 | The interface between the Batch Sampling and Analysis Station and the Laboratory Information Management System SHALL transmit batch screening results (pH, conductivity, temperature, gross gamma) as structured data records within 30 seconds of measurement completion, using the facility OPC UA data bus with authentication and integrity verification. | interface-requirements | Interface Requirements |  | Test |  |  | 30-second latency ensures batch status is current in LIMS for discharge decisions. OPC UA provides standard industrial interoperability with built-in security. Integrity verification prevents corrupted data driving incorrect discharge authorisation. | interface, aetds, session-328, idempotency:ifc-sampling-lims-328
IFC-REQ-018 | The interface between the Chemical Treatment Plant and the Radioactive Waste Management Facility SHALL transfer ILW sludge in sealed 200-litre drums with dose rate not exceeding 2 mSv/h at contact, accompanied by a waste characterisation record including activity inventory, chemical composition, and package weight. | interface-requirements | Interface Requirements |  | Inspection |  |  | 2 mSv/h contact dose rate limit aligned with transport regulations and RWMF receipt acceptance criteria. Waste characterisation record required for ILW inventory accounting under LC32 and future disposal case development. | interface, aetds, session-328, idempotency:ifc-treatment-waste-328
IFC-REQ-019 | The interface between the Criticality Warning System and the Safety Interlock and Trip System SHALL transmit criticality trip signals via hardwired relay contacts (volt-free, fail-safe energise-to-run) with end-to-end signal propagation latency not exceeding 10 milliseconds and electrical isolation between the two systems maintained by galvanic isolation barriers rated to 2.5 kV. | interface-requirements | Interface Requirements |  | Test |  |  | Hardwired relay contacts ensure deterministic signal path without software dependency. Fail-safe energise-to-run means loss of signal (wire break, power loss) defaults to trip state. 10 ms latency budget allocated from the 500 ms total trip initiation budget. 2.5 kV galvanic isolation prevents fault propagation between safety-critical subsystems. | interface, safety, criticality, session-329, idempotency:ifc-cws-sit-329
IFC-REQ-020 | The interface between the Fire Detection and Suppression System and the Standby Ventilation and Emergency Isolation System SHALL transmit zone-specific fire damper closure commands within 2 seconds of fire confirmation, using dedicated fire-rated cabling (BS 8519) maintaining circuit integrity for a minimum of 120 minutes. | interface-requirements | Interface Requirements |  | Test |  |  | Zone-specific commands prevent whole-facility ventilation shutdown on a single-zone fire, maintaining containment in unaffected zones. 2-second command latency ensures damper closure completes before fire growth compromises ductwork integrity. 120-minute fire-rated cabling per BS 8519 ensures damper commands remain available throughout a fire event. | interface, fire, ventilation, session-329, idempotency:ifc-fds-sveis-329
IFC-REQ-021 | The interface between the Emergency Power System and the Criticality Warning System SHALL provide dual-redundant UPS-backed power feeds at 24 VDC with automatic changeover in less than 1 millisecond, each feed capable of sustaining the full criticality detection load independently. | interface-requirements | Interface Requirements |  | Test |  |  | Criticality detection must operate continuously without any monitoring gap. Dual-redundant feeds ensure single UPS failure does not interrupt detection. 1 ms changeover maximum prevents detector reset or transient loss of counting that could mask a neutron burst. 24 VDC selected as standard for nuclear instrumentation per IEC 61225. | interface, power, criticality, session-329, idempotency:ifc-eps-cws-329
IFC-REQ-022 | The interface between the Safety Interlock and Trip System and the Depression Cascade Control System SHALL provide ventilation isolation trip commands on confirmed fire or seismic event, with the depression cascade controller maintaining minimum 10 Pa containment depression on non-affected zones during partial isolation. | interface-requirements | Interface Requirements |  | Test |  |  | Partial isolation allows ventilation to continue in unaffected zones, maintaining containment where there is no fire. 10 Pa minimum depression during partial isolation is the lower bound required to prevent contamination migration through door seals and penetrations under wind loading conditions at the dockyard coastal site. | interface, safety, ventilation, session-329, idempotency:ifc-sit-dccs-329
IFC-REQ-023 | The interface between the Spill Containment and Emergency Decontamination System and the Active Drain Collection Network SHALL route all spill drainage and decontamination shower effluent via sealed stainless steel pipework to the active drain collection system, with no cross-connection to inactive drainage, and floor drain traps maintaining a minimum 50 mm water seal. | interface-requirements | Interface Requirements |  | Inspection |  |  | Cross-connection to inactive drains would bypass the active effluent treatment system, risking unmonitored radioactive discharge. 50 mm water seal prevents airborne contamination backflow through the drain system, which would compromise the depression cascade containment regime. | interface, containment, effluent, session-329, idempotency:ifc-scd-adn-329
IFC-REQ-024 | The interface between the Solid Waste Characterization and Segregation Station and the Solid Waste Packaging and Compaction System SHALL transfer categorised waste items with an accompanying electronic categorisation certificate containing: waste category (VLLW/LLW/ILW), measured activity inventory, surface contamination results, and approved packaging route, transmitted via the Waste Records and Consignment System within 5 minutes of categorisation completion. | interface-requirements | Interface Requirements |  | Test |  |  | Waste must not be packaged without a completed categorisation record — packaging the wrong category of waste (e.g., ILW into an LLW drum) is a regulatory non-compliance that could result in incorrectly disposed packages. The 5-minute transfer time ensures real-time data availability for packaging operators and prevents orphaned waste items sitting without categorisation records. | interface, waste-management, characterization, packaging, session-330, idempotency:ifc-char-pack-330
IFC-REQ-025 | The interface between the Solid Waste Packaging and Compaction System and the Interim Waste Store SHALL transfer sealed waste drums with verified surface contamination below transport limits, each drum bearing a unique barcode label cross-referenced to its waste package data record, with a maximum of 4 hours between sealing and placement in the designated store position. | interface-requirements | Interface Requirements |  | Test |  |  | The 4-hour window between sealing and store placement limits the time drums spend in the uncontrolled transfer area, reducing the risk of accidental damage or loss of traceability. Barcode labelling provides positive identification that does not degrade during multi-decade storage, unlike paper labels. Surface contamination verification at this boundary prevents contamination entering the store. | interface, waste-management, packaging, interim-store, session-330, idempotency:ifc-pack-store-330
IFC-REQ-026 | The interface between the Liquid Waste Conditioning System and the Interim Waste Store SHALL transfer cemented wasteform packages only after a minimum 7-day curing period and confirmation that compressive strength sampling indicates the batch will meet the 4 MPa 28-day target, with each package assigned a storage position that respects criticality spacing requirements for its fissile content. | interface-requirements | Interface Requirements |  | Test |  |  | The 7-day hold before transfer allows early-age compressive strength testing (typically 7-day cubes) to verify the cement mix is developing strength correctly before committing the package to long-term storage. Packages failing early strength checks can be remediated before placement in the store, where retrieval for rework is operationally expensive. Criticality spacing assignment at transfer time prevents ad-hoc placement that could violate keff limits. | interface, waste-management, liquid-conditioning, interim-store, session-330, idempotency:ifc-liquid-store-330
IFC-REQ-027 | The interface between the Liquid Waste Conditioning System and the Active Drain Collection Network SHALL return evaporator distillate to the active drain system at a maximum activity concentration of 100 Bq/L total alpha and 1000 Bq/L total beta/gamma, at a flow rate not exceeding 0.5 L/min, via a dedicated sampling point that enables batch hold-and-release before drain entry. | interface-requirements | Interface Requirements |  | Test |  |  | Distillate activity limits are derived from the AETDS input assumptions — exceeding these values would overload the Chemical Treatment Plant and risk breaching discharge authorisation limits. The hold-and-release sampling point provides an additional verification barrier, consistent with the ALARP principle, before returning potentially active liquid to the site drainage system. | interface, waste-management, liquid-conditioning, effluent, session-330, idempotency:ifc-liquid-drain-330
IFC-REQ-028 | The interface between the Waste Records and Consignment System and the Laboratory Information Management System SHALL exchange waste characterisation analytical results and sample chain-of-custody data via a bidirectional API with mutual TLS authentication, with data synchronisation completing within 60 seconds of record creation in either system. | interface-requirements | Interface Requirements |  | Test |  |  | Analytical results from the characterization station are initially recorded in LIMS (as the authoritative analytical record under ISO 17025) and must flow to the waste records system for waste package data records. Bidirectional exchange ensures traceability in both directions. Mutual TLS prevents unauthorised data injection into either system. The 60-second synchronisation window ensures near-real-time consistency between the two databases. | interface, waste-management, records, lims, session-330, idempotency:ifc-records-lims-330
IFC-REQ-029 | The interface between the Chemical Treatment Plant and the Liquid Waste Conditioning System SHALL transfer treatment sludges and concentrates via a shielded pipeline with a maximum throughput of 50 L/batch, with each batch accompanied by an analytical certificate confirming radionuclide inventory and chemical composition, and a physical interlock preventing transfer if the conditioning system is not in a ready state. | interface-requirements | Interface Requirements |  | Test |  |  | The Chemical Treatment Plant generates sludges and concentrates that exceed discharge authorisation limits and must be conditioned as solid waste. The 50 L/batch limit is sized to the cementation vessel capacity. The analytical certificate is required because the cement formulation must be adjusted for different waste chemistries. The physical interlock prevents overfilling or transfer to an unprepared system, both of which are credible initiating events in the liquid waste safety case. | interface, waste-management, liquid-conditioning, aetds, session-330, idempotency:ifc-ctp-liquid-330
IFC-REQ-030 | The interface between the Sample Receiving Bay and the Dose Rate and Contamination Screening Station SHALL transfer sample containers via a pass-through hatch with interlocked doors, with each container accompanied by a physical transfer slip recording sender identity, sample origin, and declared isotopes, and the hatch SHALL be sized to accommodate containers up to 500 mm diameter and 800 mm height including shielded transport pigs. | interface-requirements | Interface Requirements |  | Test |  |  | The pass-through hatch maintains the C2/C3 containment boundary while enabling sample transfer. Physical transfer slips provide backup chain-of-custody when electronic systems are unavailable. Container sizing is based on the largest transport pig used for reactor coolant samples at UK naval dockyards. | interface, sample-receipt, session-331, idempotency:ifc-recv-screen-331
IFC-REQ-031 | The interface between the Dose Rate and Contamination Screening Station and the Sample Preparation Laboratory SHALL communicate sample handling category (contact-handled or remote-handled), measured dose rate, and contamination screening results electronically to the Sample Registration System within 30 seconds of measurement completion, and SHALL physically route contact-handled samples via a direct transfer corridor and remote-handled samples via a shielded transfer route with manipulator access. | interface-requirements | Interface Requirements |  | Test |  |  | Dual routing based on activity level is a fundamental ALARP measure — contact-handled samples can be processed at open bench workstations while remote-handled samples require shielded fume cupboards or glove boxes. Electronic data transfer ensures screening results are captured in the sample record before preparation begins, preventing preparation without characterisation. | interface, sample-receipt, session-331, idempotency:ifc-screen-prep-331
IFC-REQ-032 | The interface between the Sample Registration and Chain-of-Custody System and the Laboratory Information Management System SHALL transmit sample registration records including unique identifier, origin, requested analyses, handling category, and screening results via a secure authenticated API, with message queuing to buffer transmissions during LIMS unavailability and automatic resynchronisation within 15 minutes of LIMS recovery. | interface-requirements | Interface Requirements |  | Demonstration |  |  | The Registration System must operate independently of LIMS to prevent sample receipt being blocked by LIMS outage. Message queuing with guaranteed delivery ensures no registration data is lost. The 15-minute resynchronisation window ensures LIMS work orders are generated promptly after recovery, maintaining laboratory throughput. | interface, sample-receipt, session-331, idempotency:ifc-reg-lims-331
IFC-REQ-033 | The interface between the Acid Digestion and Dissolution Station and the Actinide Separation Chemistry Station SHALL transfer dissolved sample aliquots in capped PTFE or borosilicate glass vials labelled with sample ID barcode, with a maximum transfer volume of 100 mL per aliquot, and the transferred solution SHALL be in a 2-8 M HNO3 matrix compatible with direct loading onto TEVA/TRU extraction chromatography columns without further matrix adjustment. | interface-requirements | Interface Requirements |  | Test |  |  | Standardising the dissolution matrix to 2-8 M HNO3 eliminates a manual acid adjustment step before column loading, reducing both operator dose and the risk of cross-contamination from additional reagent additions. TEVA and TRU resins require nitric acid feed at this molarity range for quantitative actinide uptake. PTFE vials prevent HF etching of glass that would release interfering elements. | interface, radchem-sep, session-331, idempotency:ifc-acid-actinide-331
IFC-REQ-034 | The interface between the Actinide Separation Chemistry Station and the Alpha Spectrometry Laboratory SHALL deliver purified actinide fractions electrodeposited onto 25 mm diameter polished stainless steel discs, with a maximum deposited mass of 100 micrograms per disc to maintain alpha spectral resolution, and each disc SHALL be labelled with sample ID, fraction identity (Pu, Am/Cm, U), and electrodeposition date. | interface-requirements | Interface Requirements |  | Inspection |  |  | Electrodeposition onto thin-layer stainless steel discs is the standard source preparation for alpha spectrometry. The 100 microgram mass limit prevents self-absorption degradation of alpha peak resolution — thicker deposits broaden peaks below the 40 keV FWHM required to resolve Pu-238 from Pu-239/240. This interface is the hand-off from wet chemistry to counting and defines the source quality that drives measurement performance. | interface, radchem-sep, session-331, idempotency:ifc-actinide-alpha-331
IFC-REQ-035 | The interface between the Separations Waste Segregation Point and the Radioactive Waste Management Facility SHALL segregate waste into at least four streams: aqueous radioactive liquids routed via active drain (pH 1-13, max 1E6 Bq/L alpha), organic solvent waste collected in 5-litre containers with flash point and activity labelling, compactible dry active waste in 60-litre bags, and non-compactible items in dedicated bins, with each waste container identified by barcode linking to the originating sample work order. | interface-requirements | Interface Requirements |  | Inspection |  |  | Waste segregation at source is mandatory under the UK waste hierarchy and prevents downstream processing problems — mixing organic solvents with aqueous waste produces flammable mixtures in the effluent treatment plant, and unsegregated alpha-bearing waste contaminates low-level waste streams causing reclassification as ILW. Traceability to originating sample work order is required for nuclear material accountancy of fissile material in waste. | interface, radchem-sep, session-331, idempotency:ifc-waste-seg-rwm-331
IFC-REQ-036 | The interface between the HPGe Detector Array and the Digital MCA and Signal Processing System SHALL carry charge-sensitive preamplifier output pulses via RG-62 coaxial cable with BNC connectors, with signal amplitude range 0 to 10 V, rise time of less than 100 ns, and maximum cable length of 15 m to limit noise pickup in the nuclear-licensed facility environment. | interface-requirements | Interface Requirements |  | Test |  |  | Coaxial cable length and type determine signal-to-noise ratio at the MCA input. 15 m maximum allows placement of MCA electronics outside the counting room to reduce electronic noise contribution to background. BNC standardises connector interface across detector vendors. | interface, gamma-spectrometry, session-333, idempotency:ifc-hpge-mca-signal-333
IFC-REQ-037 | The interface between the Digital MCA and Signal Processing System and the Gamma Spectroscopy Analysis and QA Workstation SHALL transfer spectral data via dedicated Gigabit Ethernet on a physically isolated VLAN, using the vendor spectroscopy protocol, with maximum spectral transfer latency of 2 seconds for a 16384-channel spectrum. | interface-requirements | Interface Requirements |  | Test |  |  | Physical VLAN isolation prevents general network traffic from disrupting acquisitions during long counting periods. 2-second transfer latency is needed to support automated sequential sample counting without operator intervention. | interface, gamma-spectrometry, session-333, idempotency:ifc-mca-analysis-network-333
IFC-REQ-038 | The interface between the Gamma Spectroscopy Analysis and QA Workstation and the Laboratory Information Management System SHALL transmit analysis results in a structured format including nuclide identity, activity concentration, combined uncertainty at k equals 2, MDA, counting parameters, and QC status, with results posted automatically upon analyst approval. | interface-requirements | Interface Requirements |  | Demonstration |  |  | Automated result transfer eliminates transcription errors which are the largest source of non-measurement errors in accredited nuclear laboratories. Structured format enables LIMS to perform automated regulatory limit checking before results release. | interface, gamma-spectrometry, session-333, idempotency:ifc-analysis-lims-results-333
IFC-REQ-039 | The interface between the Electrodeposition and Source Preparation Station and the Alpha Detector Chamber Array SHALL use standard 25 mm diameter stainless steel planchets loaded into spring-clip sample holders compatible with the vacuum chamber sample tray, with each planchet uniquely identified by engraved number cross-referenced to the chain-of-custody system. | interface-requirements | Interface Requirements |  | Inspection |  |  | Standardised planchet geometry ensures reproducible source-detector distance and counting geometry across all 8 chambers. Engraved identification prevents sample mix-up which would invalidate nuclear material accountancy records. | interface, alpha-spectrometry, session-333, idempotency:ifc-electrodep-detector-333
IFC-REQ-040 | The interface between the Alpha Detector Chamber Array and the Alpha MCA and Spectral Processing Unit SHALL carry PIPS detector bias voltage (0 to 100 V) and preamplifier output pulses via multipin vacuum feedthrough connectors, with each chamber providing a vacuum status signal (good or bad) that enables or inhibits MCA acquisition for that channel. | interface-requirements | Interface Requirements |  | Test |  |  | Vacuum interlock prevents acquisition of degraded spectra that would waste counting time and potentially generate misleading results. The multipin feedthrough maintains vacuum integrity while passing electrical signals. | interface, alpha-spectrometry, session-333, idempotency:ifc-alpha-detector-mca-333
IFC-REQ-041 | The interface between Radiochemical Separations Laboratory and LSC Sample Preparation Station SHALL transfer purified H-3 fractions as aqueous distillate in sealed 20 mL HDPE vials with chain-of-custody labels, and C-14 fractions as benzene or aqueous solution in sealed glass vials, with transfer batch documentation recording fraction identity, source sample ref, separation date, expected activity range, and any chemical interferents. | interface-requirements | Interface Requirements |  | Test |  |  | Physical and chemical form of separated fractions determines cocktail compatibility and quench behaviour. H-3 as aqueous distillate is standard for direct cocktail mixing. C-14 may arrive as benzene (from benzene synthesis) or aqueous (from wet oxidation), requiring different cocktail types. Sealed vials prevent tritiated water vapour loss (H-3 has high vapour pressure). Documentation enables sample traceability and alerts operators to high-activity samples requiring dilution. | interface, liquid-scintillation, session-334, idempotency:ifc-sep-lsc-prep-334
IFC-REQ-042 | The interface between LSC Sample Preparation Station and Liquid Scintillation Counter Array SHALL deliver prepared counting vials in standard 20 mL polyethylene or low-potassium glass vials compatible with the counter sample changer, with each vial barcode-labelled linking to the LIMS work order, and loaded into counting cassettes in the sequence defined by the counting protocol (standards, blanks, samples, spikes). | interface-requirements | Interface Requirements |  | Test |  |  | Vial type affects counting geometry and background. Polyethylene vials have lower background than glass but higher permeability. The counter sample changer requires standard 20 mL vial geometry. Barcode labelling enables automatic protocol assignment and result association. Counting sequence (standard-blank-sample-spike) is essential for QA/QC batch validation per ISO 17025. | interface, liquid-scintillation, session-334, idempotency:ifc-lsc-prep-counter-334
IFC-REQ-043 | The interface between Liquid Scintillation Counter Array and LSC Data Analysis and QA Workstation SHALL transfer raw spectral data in vendor-compatible format over Ethernet TCP/IP at completion of each counting cycle, including full pulse height spectrum (0-2000 keV equivalent), tSIE quench indicator, total and windowed count rates, PSA histogram, counting time, and instrument status flags. | interface-requirements | Interface Requirements |  | Test |  |  | Full spectral data transfer enables offline re-analysis with different windows or quench correction methods. The tSIE quench indicator is required for external standard quench correction curve application. PSA histograms enable QA verification of alpha/beta separation quality. Instrument status flags alert to anomalies (power interruptions, temperature excursions, detector malfunction). Ethernet TCP/IP is standard for modern LSC-to-workstation communication. | interface, liquid-scintillation, session-334, idempotency:ifc-lsc-counter-analysis-334
IFC-REQ-044 | The interface between LSC Data Analysis and QA Workstation and Laboratory Information Management System SHALL transfer validated analysis results including nuclide activity concentration, combined standard uncertainty, expanded uncertainty at k=2, MDA, QC batch acceptance status, and analysis certificate reference, via authenticated API or file transfer with data integrity verification, within 30 minutes of batch QA approval. | interface-requirements | Interface Requirements |  | Test |  |  | LIMS is the system of record for all analytical results. Automated transfer eliminates transcription errors (a key concern in regulated nuclear analysis). The 30-minute transfer target supports same-day reporting for urgent dockyard operational samples. Data integrity verification prevents corruption during transfer. Authentication prevents unauthorised result modification. All listed data fields are required for EA regulatory reporting and UKAS accreditation records. | interface, liquid-scintillation, session-334, idempotency:ifc-lsc-analysis-lims-334
IFC-REQ-045 | The interface between Radiochemical Separations Laboratory and ICP-MS Sample Introduction System SHALL deliver purified actinide fractions in 15 mL polypropylene centrifuge tubes in 2-5% HNO3 matrix, with fraction identity label, expected concentration range, and carrier chemistry documentation, enabling the autosampler to process fractions without manual matrix adjustment. | interface-requirements | Interface Requirements |  | Test |  |  | Matrix compatibility between separation chemistry and ICP-MS is critical. The 2-5% HNO3 matrix is standard for ICP-MS analysis; higher acid concentrations cause signal suppression and torch damage. Polypropylene tubes are compatible with the autosampler probe. Concentration range documentation enables method selection (standard vs. dilution protocol) and prevents detector saturation from unexpectedly high-concentration samples. | interface, icp-ms, session-334, idempotency:ifc-sep-icpms-334
IFC-REQ-046 | The interface between ICP-MS Instrument and ICP-MS Data Acquisition and Processing Workstation SHALL transfer raw mass spectral data including counts per second at each measured mass, internal standard signal, analyte-to-internal-standard ratios, and instrument tuning parameters, via vendor data format over Ethernet TCP/IP, with automatic data backup to network storage. | interface-requirements | Interface Requirements |  | Test |  |  | Complete mass spectral data transfer enables offline re-processing with alternative quantification methods (standard addition vs. external calibration). Internal standard ratios are needed for matrix effect correction. Instrument tuning parameters provide audit trail for method validation. Network backup protects against workstation failure — loss of raw data would require re-analysis of potentially irreplaceable dockyard samples. | interface, icp-ms, session-334, idempotency:ifc-icpms-data-334
IFC-REQ-047 | The interface between LIMS Instrument Interface Module and all analytical instrument workstations SHALL use standardised message formats (HL7 LIS2-A2 or ASTM E1394 for structured results, vendor API where available), with message queuing to buffer results during LIMS maintenance windows and automatic retry with exponential backoff on communication failure, supporting minimum 200 result transactions per hour aggregate throughput. | interface-requirements | Interface Requirements |  | Test |  |  | Standardised messaging protocols ensure maintainability as instruments are replaced over the laboratory lifecycle (15-20 years). Message queuing prevents result loss during planned LIMS maintenance. The 200 transactions/hour throughput handles peak campaign periods when multiple instruments are running simultaneously. Exponential backoff prevents network flooding during recovery from outage. | interface, lims, session-334, idempotency:ifc-lims-instruments-334
IFC-REQ-048 | The interface between the Building Management System and the Electrical Power Distribution System SHALL provide real-time monitoring of distribution board loads, transformer temperature, and power quality parameters (voltage, frequency, THD) via Modbus TCP at 1-second polling interval, with automatic load shedding commands on transformer overtemperature. | interface-requirements | Interface Requirements |  | Test |  |  | Transformer overtemperature protection requires fast response to prevent winding damage. 1-second polling enables BMS to detect rising trends and initiate non-essential load shedding before protection relay trip, maintaining supply to essential circuits. | interface, lab-utilities, session-335, idempotency:ifc-bms-epds-335
IFC-REQ-049 | The interface between the Building Management System and the Chilled Water and Heating System SHALL transmit chiller plant status, flow and return temperatures, pump speeds, and valve positions via BACnet/IP, with the BMS providing temperature setpoint control and automatic duty/standby changeover on chiller fault. | interface-requirements | Interface Requirements |  | Test |  |  | Laboratory temperature stability depends on closed-loop control of chilled and heating water. BACnet/IP is the standard protocol for HVAC plant integration, enabling the BMS to coordinate chiller operation with AHU demand and implement optimal start/stop sequences. | interface, lab-utilities, session-335, idempotency:ifc-bms-chw-335
IFC-REQ-050 | The interface between the Laboratory Gas Supply System and the Fire Detection and Suppression System SHALL receive zone fire alarm signals as volt-free contacts, triggering automatic closure of flammable gas isolation valves within the affected zone. Valve closure status SHALL be reported back to the fire panel within 3 seconds. | interface-requirements | Interface Requirements |  | Test |  |  | Fire safety case requires confirmed gas isolation as part of the fire response strategy. Volt-free contacts provide electrical isolation between safety-classified fire system and utility gas controls. 3-second feedback confirms isolation before fire suppression activation in gas-supplied areas. | interface, lab-utilities, session-335, idempotency:ifc-gas-fire-335
IFC-REQ-051 | The interface between the Chilled Water and Heating System and the Supply Air Handling Unit SHALL provide chilled water at 6 plus or minus 0.5 degrees C flow temperature and LTHW at 82 plus or minus 2 degrees C to AHU cooling and heating coils respectively, with two-port control valves modulated by the BMS to maintain supply air temperature setpoint. | interface-requirements | Interface Requirements |  | Test |  |  | The AHU conditions fresh supply air before delivery to laboratories. Temperature tolerance on chilled water flow directly impacts achievable supply air dewpoint and dehumidification capacity. The ventilation system depression cascade cannot function correctly if supply air temperature is uncontrolled. | interface, lab-utilities, session-335, idempotency:ifc-chw-ahu-335
IFC-REQ-052 | The interface between the Physical Security and Access Control System and the Health Physics Central Alarm and Display System SHALL exchange access zone occupancy data so that HP monitoring knows which personnel are present in each radiological zone, enabling dose tracking per individual and supporting emergency muster accountability. | interface-requirements | Interface Requirements |  | Test |  |  | Regulation 19 of IRR 2017 requires employers to investigate doses that may exceed dose limits. Correlating zone occupancy with area dose rates enables real-time estimated dose calculation. Emergency muster requires knowing who is in controlled areas at the time of an incident. | interface, lab-utilities, session-335, idempotency:ifc-security-hp-335
IFC-REQ-053 | The interface between the Deionised Water Treatment and Distribution System and the Acid Digestion and Dissolution Station SHALL deliver DI water at point-of-use dispensers with a maximum delivery distance of 3 metres from the treatment polishing cartridge, maintaining resistivity above 18.0 megaohm-cm at the dispenser outlet. | interface-requirements | Interface Requirements |  | Test |  |  | Ultra-pure water resistivity degrades with pipe length due to CO2 absorption and material leaching. 3-metre maximum distance from final polishing ensures point-of-use quality meets ASTM D1193 Type I without additional treatment. Critical for acid digestion blanks in trace analysis. | interface, lab-utilities, session-335, idempotency:ifc-diw-digestion-335
IFC-REQ-054 | The interface between the Electrical Power Distribution System and the Emergency Power System SHALL provide automatic changeover via a motorised changeover switch that detects normal supply failure (voltage below 85 percent of nominal for more than 500ms) and transfers essential loads to the emergency generator within 15 seconds of generator reaching rated voltage and frequency. | interface-requirements | Interface Requirements |  | Test |  |  | 15-second transfer time is derived from the UPS autonomy margin: essential loads on UPS bridge the gap during generator start (typically 8-10 seconds to reach rated speed). The 500ms detection delay prevents nuisance transfers on voltage transients while ensuring genuine failures trigger changeover before UPS battery depletion. | interface, lab-utilities, session-335, idempotency:ifc-power-emergency-335
IFC-REQ-055 | The interface between the Building Management System and the Safety Interlock and Trip System SHALL be implemented as a unidirectional data diode or hardware-enforced one-way gateway permitting status data transfer from the safety system to the BMS for display purposes only, with no data path from the BMS network to the safety system logic solver. | interface-requirements | Interface Requirements |  | Test |  |  | The BMS is a general-purpose SCADA system connected to multiple building services networks and potentially to wider site networks. The Safety Interlock and Trip System is a SIL-rated safety system whose integrity must be protected from cyber threats and accidental corruption. A unidirectional gateway per IEC 62443 zone and conduit model ensures that compromise of the BMS cannot affect safety system operation. This is consistent with ONR guidance on independence of safety-related and non-safety systems. | interface, cybersecurity, safety, validation, session-337
STK-REQ-001 | The Radiochemistry Laboratory SHALL provide analysis results for reactor primary coolant samples within 4 hours of sample receipt for safety-critical parameters (gross alpha, gross beta, H-3, dissolved gases) to support reactor operations decision-making. | stakeholder-requirements | Stakeholder Requirements |  | Test |  |  | Reactor operations require timely coolant chemistry data to confirm fuel clad integrity and authorise continued power operation. A 4-hour turnaround enables the duty chemist to report results within the same watch period, supporting the submarine's operational programme. Delayed results could force precautionary reactor shutdown with significant fleet availability impact. | stakeholder, session-325
STK-REQ-002 | The Radiochemistry Laboratory SHALL demonstrate compliance with the nuclear site licence conditions, particularly LC14 (safety documentation), LC23 (operating rules), LC27 (safety mechanisms), and LC28 (examination, maintenance, inspection and testing) as assessed by the ONR. | stakeholder-requirements | Stakeholder Requirements |  | Inspection |  |  | ONR is the statutory nuclear safety regulator for UK defence nuclear sites. Non-compliance with licence conditions can result in enforcement action including improvement notices, prohibition notices, or prosecution. The laboratory safety case and its compliance arrangements must satisfy ONR inspectors during routine and reactive inspections. | stakeholder, session-325
STK-REQ-003 | The Radiochemistry Laboratory SHALL ensure that all liquid and gaseous radioactive discharges remain within the limits and conditions specified in the current environmental permit (EPR or RSA authorisation) and SHALL provide verifiable discharge monitoring data to the Environment Agency or SEPA. | stakeholder-requirements | Stakeholder Requirements |  | Analysis |  |  | The laboratory generates liquid effluent (from sample preparation, washing, decontamination) and gaseous discharges (from fume cupboard and glovebox extracts). EA/SEPA discharge permits set annual limits on total alpha, total beta, tritium, and specific nuclides. Exceeding these limits is a criminal offence and risks permit revocation, which would shut down the laboratory. | stakeholder, session-325
STK-REQ-004 | The Radiochemistry Laboratory SHALL maintain individual annual effective dose to all laboratory personnel as low as reasonably practicable (ALARP) and in all cases below a dose constraint of 10 mSv per year, with investigation levels at 1 mSv per quarter. | stakeholder-requirements | Stakeholder Requirements |  | Analysis |  |  | Laboratory staff handle open radioactive sources daily, creating external irradiation and internal contamination hazards. The 10 mSv/year constraint provides margin below the 20 mSv/year IRR17 legal limit. ALARP is a legal requirement under IRR17 Regulation 9 and is actively enforced by ONR. The 1 mSv/quarter investigation level ensures trends are caught early before approaching annual limits. | stakeholder, session-325
STK-REQ-005 | The Radiochemistry Laboratory SHALL provide nuclear material assay results with measurement uncertainties meeting IAEA International Target Values for destructive assay of uranium and plutonium, to support the site nuclear material accountancy system. | stakeholder-requirements | Stakeholder Requirements |  | Test |  |  | As a defence nuclear site, the dockyard must account for all special nuclear material to IAEA safeguards-equivalent standards. The laboratory's mass spectrometric and radiometric assay results feed directly into the nuclear material balance. Measurement uncertainty must meet ITVs (e.g. 0.1% relative for U mass, 0.5% for Pu isotopic composition by IDMS) to avoid material unaccounted for (MUF) triggers. | stakeholder, session-325
STK-REQ-006 | The Radiochemistry Laboratory SHALL be designed to facilitate future decommissioning and decontamination, with all active areas employing smooth, non-porous, sealed surface finishes and modular service routing that enables progressive zone-by-zone clearance without compromising containment of remaining active areas. | stakeholder-requirements | Stakeholder Requirements |  | Inspection |  |  | Nuclear site licence condition LC35 requires the licensee to make and implement adequate arrangements for decommissioning. ONR Safety Assessment Principles EKP.2 requires that design facilitates decommissioning from the outset. A radiochemistry laboratory handling open sources creates surface contamination that, if embedded in porous materials, would generate significant quantities of intermediate-level waste during decommissioning. Smooth sealed surfaces and modular services are established good practice (IAEA SSG-47) to minimise decommissioning waste volumes and worker dose. | stakeholder, decommissioning, validation, session-337
STK-REQ-007 | The Radiochemistry Laboratory SHALL protect nuclear material accountancy data, safeguards-relevant analytical results, and facility security information in accordance with the Classification Policy Framework and the site Nuclear Industries Security Regulations 2003 (NISR) security plan, with electronic systems meeting NCSC Cyber Essentials Plus as a minimum baseline. | stakeholder-requirements | Stakeholder Requirements |  | Inspection |  |  | The dockyard is a defence nuclear site where nuclear material accountancy data (Pu/U assay results, isotopic compositions) and facility vulnerability information are classified under the government security classification policy. NISR 2003 requires approved security plans for nuclear premises. ONR Civil Nuclear Security division assesses cyber security under the NIS Regulations 2018. LIMS contains safeguards-sensitive data that must be protected from exfiltration or manipulation — loss of integrity in NMA results could mask material diversion. | stakeholder, security, validation, session-337
SUB-REQ-001 | The Supply Air Handling Unit SHALL deliver conditioned air to C1 and C2 zones at 20 plus or minus 2 degrees Celsius and 40 to 60 percent relative humidity at a supply rate of not less than 8000 cubic metres per hour. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Temperature and humidity stability is required for instrument calibration consistency in the gamma spectrometry and alpha spectrometry suites. A 2-degree-C band ensures detector gain drift remains within calibration tolerance. Humidity control prevents condensation on cooled detector surfaces and static discharge in low-background counting areas. | subsystem, ventilation, session-326, idempotency:sub-sahu-temp-326
SUB-REQ-002 | The Extract Fan System SHALL provide automatic changeover from duty to standby fan within 10 seconds of duty fan failure, maintaining extract airflow at not less than 80 percent of nominal during the changeover transient. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | A 10-second changeover limit ensures the depression cascade does not collapse during fan failure. At 10000 m3/h extract rate, a 10-second interruption allows approximately 28 cubic metres of unextracted air — insufficient to pressurise C3/C4 zones beyond the 15 Pa minimum differential given room volumes of approximately 200 m3. The 80 percent flow floor prevents cascade reversal during transient. | subsystem, ventilation, session-326, idempotency:sub-efs-changeover-326
SUB-REQ-003 | The HEPA Filtration Assembly SHALL achieve a minimum decontamination factor of 10000 (99.99 percent retention efficiency) on each filter bank when tested in situ using DOP/PAO aerosol challenge at rated airflow, with leak-tight safe-change housings on the primary bank preventing operator exposure exceeding 1 microsievert during filter replacement. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | H14-grade HEPA filters are rated at 99.995 percent but in-situ performance must achieve at least 99.99 percent to account for housing seal leakage and installation imperfections. The 1-microsievert dose constraint during filter change drives the safe-change (bag-in/bag-out) housing requirement — used primary HEPA filters in a radiochemistry lab accumulate alpha activity that would present significant skin and inhalation dose if exposed during manual removal. | subsystem, ventilation, session-326, idempotency:sub-hepa-dop-326
SUB-REQ-004 | The Depression Cascade Control System SHALL restore zone differential pressures to within 5 Pa of setpoint within 2 seconds of a step disturbance of 10 Pa, for all zone boundaries (C1/C2 at minus 15 Pa, C2/C3 at minus 30 Pa, C3/C4 at minus 50 Pa). | subsystem-requirements | Subsystem Requirements |  | Test |  |  | A 2-second response time prevents sustained cascade reversal during door openings or fume cupboard sash movements. At typical room volumes (150-200 m3) and extract rates, a 10 Pa step disturbance represents a major perturbation (e.g. door opening). Failure to restore within 2 seconds risks airflow from higher-activity to lower-activity zones, potentially spreading alpha contamination. The 5 Pa tolerance band prevents continuous hunting while maintaining containment integrity. | subsystem, ventilation, session-326, idempotency:sub-dccs-response-326
SUB-REQ-005 | The Stack Monitoring and Discharge System SHALL automatically close the stack isolation damper within 5 seconds when the measured airborne alpha activity exceeds one-tenth of the derived air concentration limit for the most restrictive isotope handled (Pu-239, DAC 0.08 Bq/m3, trigger at 0.008 Bq/m3). | subsystem-requirements | Subsystem Requirements |  | Test |  |  | One-tenth DAC trigger provides margin before the regulatory discharge limit is approached. Pu-239 is the most restrictive alpha emitter handled in the laboratory (lowest ALI). The 5-second closure time limits the total release during the transient to approximately 0.014 Bq at 10000 m3/h stack flow — negligible compared to quarterly discharge limits. Damper closure faster than 5 seconds risks pressure surge in the extract ductwork. | subsystem, ventilation, session-326, idempotency:sub-stack-isolation-326
SUB-REQ-006 | The Fume Cupboard and Glove Box Extract Network SHALL maintain a minimum face velocity of 0.5 metres per second at each C3 zone fume cupboard with the sash at the 500 mm working opening, and shall maintain each C4 zone glove box at minus 250 Pa relative to the surrounding room. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | The 0.5 m/s face velocity is the minimum specified by BS EN 14175-3 for containment of hazardous substances and is the standard adopted by UKAS-accredited laboratories handling alpha-active materials. Below 0.5 m/s, turbulent room air currents can defeat the fume cupboard air curtain, allowing alpha particulate to escape to the operator breathing zone. Glove box negative pressure of 250 Pa ensures that any glove breach results in inward airflow, preventing release of loose alpha contamination. | subsystem, ventilation, session-326, idempotency:sub-fc-facevel-326
SUB-REQ-007 | The Standby Ventilation and Emergency Isolation System SHALL maintain containment integrity (depression cascade not less than 10 Pa at all zone boundaries) for a minimum of 72 hours following loss of normal mains power, achieving Safety Integrity Level 2 for the containment preservation function per IEC 61511. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | The 72-hour duration covers the maximum credible mains restoration time for a UK nuclear licensed site (National Grid ESO restoration planning assumption). SIL 2 is assigned because loss of containment during alpha handling could result in individual effective doses of 20-50 mSv — ALARP assessment shows that the cost of SIL 2 (redundant components, proof testing) is grossly disproportionate to the risk reduction foregone. The 10 Pa minimum (versus normal 15-50 Pa) is the degraded-mode threshold below which cascade reversal becomes possible with door openings. | subsystem, ventilation, safety, session-326, idempotency:sub-standby-sil2-326
SUB-REQ-008 | The HEPA Filtration Assembly SHALL provide continuous differential pressure monitoring across each filter bank with alarm at 250 Pa (advisory, approaching loading limit) and 500 Pa (mandatory filter change), transmitting readings to the Depression Cascade Control System at intervals not exceeding 5 seconds. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | HEPA filter loading increases pressure drop from approximately 250 Pa (clean) to over 750 Pa (blocked). At 500 Pa, extract fan power consumption increases by approximately 40 percent and flow reduction risks cascade collapse. The 250 Pa advisory threshold provides 2-4 weeks warning (typical loading rate in a radiochemistry lab handling microgramme quantities of alpha emitters) for planned filter replacement, avoiding emergency changes that carry higher dose risk. 5-second update rate matches the cascade control loop bandwidth. | subsystem, ventilation, session-326, idempotency:sub-hepa-dp-326
SUB-REQ-009 | The Extract Fan System SHALL include vibration monitoring on all fan bearings with automatic trip at 11 mm/s RMS velocity (ISO 10816-3 Zone D boundary for Group 2 machines) and alarm at 7.1 mm/s RMS (Zone C/D boundary), initiating automatic changeover to standby fan on trip. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Extract fans operate continuously and bearing failure is the dominant failure mode. ISO 10816-3 vibration severity zones provide well-established thresholds for rotating machinery condition monitoring. Zone D (11 mm/s) indicates damage is occurring and continued operation risks catastrophic bearing seizure — which would halt extract ventilation entirely. Early warning at Zone C/D (7.1 mm/s) allows planned maintenance during low-activity periods. | subsystem, ventilation, session-326, idempotency:sub-efs-vibration-326
SUB-REQ-010 | The Continuous Air Monitor Network SHALL detect airborne alpha contamination at a sensitivity of 0.2 DAC-hours for Pu-239 (corresponding to 0.037 Bq/m3 DAC fraction) with a false alarm rate not exceeding 1 per 1000 operating hours per monitor. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Derived from SYS-REQ-006 airborne alpha detection requirement and IRR17 Schedule 3 dose assessment obligations. The 0.2 DAC-hour threshold ensures personnel dose from inhalation is assessed before exceeding 1/10 of the annual dose constraint. False alarm rate limit prevents alarm fatigue which degrades response reliability. | subsystem, radiation-protection, cam-network, session-327, idempotency:sub-cam-sensitivity-327
SUB-REQ-011 | The Continuous Air Monitor Network SHALL provide a three-level alarm cascade: investigation at 0.5 DAC, action at 1.0 DAC, and evacuation at 3.0 DAC, with alarm annunciation at both the local CAM unit and the Health Physics Central Alarm and Display System within 2 seconds of threshold exceedance. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Three-level cascade aligns with HSE Approved Code of Practice for IRR17 Regulation 18 (contingency plans). Investigation level triggers RPS assessment, action level triggers area evacuation and HP survey, evacuation level triggers emergency response. The 2-second latency ensures real-time awareness for time-critical contamination events. | subsystem, radiation-protection, cam-network, session-327, idempotency:sub-cam-alarm-327
SUB-REQ-012 | The Area Gamma Dose Rate Monitoring Array SHALL measure ambient dose equivalent rate H*(10) from 0.05 microSv/h to 10 Sv/h with energy response within plus or minus 30 percent across the range 50 keV to 3 MeV, calibrated against NPL-traceable Cs-137 and Co-60 reference fields. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Lower range of 0.05 microSv/h is needed to resolve natural background variations and detect early onset contamination in supervised areas. Upper range of 10 Sv/h covers credible accident scenarios including source handling incidents. Energy response specification ensures accurate dose assessment for the mixed gamma fields encountered from activated corrosion products and fission products in reactor coolant samples. | subsystem, radiation-protection, gamma-array, session-327, idempotency:sub-gamma-range-327
SUB-REQ-013 | The Area Gamma Dose Rate Monitoring Array SHALL alarm at three thresholds: investigation at 7.5 microSv/h, action at 25 microSv/h, and evacuation at 100 microSv/h, with dose rate update interval not exceeding 10 seconds. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Investigation level set at 3x typical background (2.5 microSv/h) to reliably distinguish source movement from statistical fluctuation. Action level corresponds to dose rate that would deliver 200 microSv in an 8-hour shift (daily investigation level). Evacuation level limits potential acute dose. 10-second update ensures timely detection of rapidly changing fields during sample handling. | subsystem, radiation-protection, gamma-array, session-327, idempotency:sub-gamma-alarm-327
SUB-REQ-014 | The Contamination Monitoring Stations SHALL achieve alpha detection sensitivity of 0.04 Bq/cm2 and beta-gamma detection sensitivity of 0.4 Bq/cm2 on hand-foot-clothing monitors, with total measurement time not exceeding 10 seconds and throughput of at least 6 persons per monitor per hour. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Detection sensitivities are set at 1/10 of the ONR-accepted surface contamination clearance levels (0.4 Bq/cm2 alpha, 4 Bq/cm2 beta-gamma) to provide adequate margin for reliable detection at the clearance boundary. 10-second measurement time and throughput requirement prevent queuing at shift changeover which would incentivise personnel to bypass monitoring. | subsystem, radiation-protection, contamination, session-327, idempotency:sub-contam-hfc-327
SUB-REQ-015 | The Contamination Monitoring Stations SHALL include portal monitors at the controlled area final exit capable of detecting 400 Bq of Cs-137 distributed on a person within a 5-second walk-through measurement, with automatic barrier lock on alarm. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Portal monitors serve as the final barrier against spread of contamination outside the controlled area. 400 Bq Cs-137 detection threshold corresponds to approximately 1/10 of the body contamination level requiring decontamination intervention. Automatic barrier lock prevents contaminated personnel from leaving without HP assessment. | subsystem, radiation-protection, contamination, session-327, idempotency:sub-contam-portal-327
SUB-REQ-016 | The Personal Dosimetry and Dose Record System SHALL issue EPDs with audible dose rate alarm at 10 microSv/h and integrated dose alarm at 200 microSv per entry, and SHALL automatically flag any individual whose running annual dose exceeds 80 percent of the 6 mSv dose constraint, restricting controlled area access until RPS authorisation. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | EPD dose rate alarm set below area gamma investigation level to provide personal warning before area alarm triggers. 200 microSv daily investigation level aligns with IRR17 Regulation 8 dose investigation requirements. 80 percent dose constraint trigger provides margin before reaching the 6 mSv investigation level and the 15 mSv statutory limit, preventing inadvertent overexposure. | subsystem, radiation-protection, dosimetry, session-327, idempotency:sub-dosim-alarm-327
SUB-REQ-017 | The Personal Dosimetry and Dose Record System SHALL submit dose records to the Central Approved Dosimetry Organisation Record (CADOR) quarterly, with reconciliation between EPD readings, TLD results, and CADOR entries completed within 10 working days of TLD processing. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | CADOR submission is a legal requirement under IRR17 Regulation 21 for classified persons. Reconciliation window of 10 working days ensures discrepancies between EPD and TLD are investigated promptly. TLD remains the legal dose of record; EPD provides operational dose management. | subsystem, radiation-protection, dosimetry, session-327, idempotency:sub-dosim-cador-327
SUB-REQ-018 | The Health Physics Central Alarm and Display System SHALL achieve 99.9 percent availability (less than 8.76 hours unplanned downtime per year) through active-standby redundant server architecture, with automatic failover completing within 30 seconds and no loss of alarm state or historian data. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Continuous radiation monitoring is an ONR licence condition requirement. 99.9 percent availability target balances cost against the risk that loss of central monitoring forces reliance on local instrument alarms alone, which lack trending and remote awareness capability. 30-second failover ensures alarm continuity during server switchover. | subsystem, radiation-protection, hp-central, session-327, idempotency:sub-hpcentral-avail-327
SUB-REQ-019 | The Health Physics Central Alarm and Display System SHALL retain all radiation monitoring data in a 10-year historian database with tamper-evident audit trail, and SHALL automatically generate quarterly EA/SEPA statutory discharge returns from stack and effluent monitoring data. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | 10-year retention satisfies ONR guidance on records retention for nuclear sites and supports epidemiological follow-up obligations. Tamper-evident audit trail required for regulatory inspection credibility. Automatic statutory returns generation reduces manual transcription errors in discharge reporting which could result in regulatory enforcement action. | subsystem, radiation-protection, hp-central, session-327, idempotency:sub-hpcentral-historian-327
SUB-REQ-020 | When mains power is lost, the Health Physics Central Alarm and Display System SHALL maintain full alarm and display functionality on UPS power for a minimum of 4 hours, with automatic graceful shutdown and data preservation if UPS capacity falls below 10 percent. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 4-hour UPS duration covers credible mains outage scenarios including time for diesel generator start and load transfer. Matches the site emergency plan assumption for restoration of essential supplies. Graceful shutdown protects historian database integrity — abrupt shutdown risks data corruption requiring manual recovery. | subsystem, radiation-protection, hp-central, session-327, idempotency:sub-hpcentral-ups-327
SUB-REQ-021 | The Radioactive Source Inventory and Calibration System SHALL maintain a complete register of all sealed radioactive sources held on site with real-time location tracking via barcode scanning, and SHALL generate automated alerts at least 30 days before each source leak test is due under IRR17 Schedule 4. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | Source inventory is a legal requirement under IRR17 Regulation 28. Real-time location tracking prevents source loss which is a nuclear security reportable event. 30-day advance alert for leak tests provides adequate scheduling margin to avoid regulatory non-compliance. Barcode scanning provides auditable source movement records for ONR inspection. | subsystem, radiation-protection, source-management, session-327, idempotency:sub-source-tracking-327
SUB-REQ-022 | The Effluent Collection and Delay Tanks SHALL hold each batch for a minimum of 24 hours before discharge authorisation to allow decay of short-lived isotopes (Na-24 half-life 15h, Mn-56 half-life 2.6h) generated during reactor coolant sample processing. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 24-hour hold ensures Na-24 decays by at least one half-life and Mn-56 by nine half-lives, reducing gross gamma activity substantially and avoiding unnecessary chemical treatment for short-lived nuclides. | subsystem, aetds, delay-tanks, session-328, idempotency:sub-delay-tanks-hold-time-328
SUB-REQ-023 | The Effluent Collection and Delay Tanks SHALL provide a minimum of four tanks of 5000 litres each, such that at least one tank is always available for filling while others are in hold, sampling, or discharge states. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | Four-tank configuration ensures continuous waste receipt from laboratory operations without interruption. Weekly laboratory throughput of approximately 8000L requires at least two tanks in parallel fill/hold, with remaining tanks for treatment and discharge. | subsystem, aetds, delay-tanks, session-328, idempotency:sub-delay-tanks-capacity-328
SUB-REQ-024 | The Effluent Collection and Delay Tanks SHALL be double-contained within bunded enclosures sized to at least 110% of the volume of the largest single tank, with leak detection sensors providing alarm within 60 seconds of bund water ingress. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 110% bund volume per nuclear site licence LC34 (leakage and escape of radioactive material). 60-second detection limit ensures rapid response before significant environmental release pathway is established. | subsystem, aetds, delay-tanks, session-328, idempotency:sub-delay-tanks-containment-328
SUB-REQ-025 | The Chemical Treatment Plant SHALL reduce total alpha activity in treated effluent to below 0.1 Bq/mL and total beta-gamma activity to below 1.0 Bq/mL, achieving a decontamination factor of at least 100 for actinides and at least 10 for Cs-137 and Sr-90. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Discharge limits derived from Environment Agency RSR permit schedule. DF of 100 for actinides required because incoming effluent from plutonium handling areas may reach 10 Bq/mL alpha. DF of 10 for fission products covers worst-case reactor coolant sample waste. | subsystem, aetds, chemical-treatment, session-328, idempotency:sub-chem-treatment-df-328
SUB-REQ-026 | The Batch Sampling and Analysis Station SHALL draw representative 500 mL aliquots from agitated delay tanks such that the sample activity concentration is within ±10% of the true batch mean, verified by triplicate sampling during commissioning. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | ±10% representativeness ensures discharge decisions are based on accurate batch characterisation. Triplicate verification during commissioning establishes confidence in the sampling system before active operations. | subsystem, aetds, batch-sampling, session-328, idempotency:sub-batch-sampling-rep-328
SUB-REQ-027 | The Discharge Monitoring and Control System SHALL isolate the discharge line via a fail-safe valve within 5 seconds of detecting total activity exceeding 80% of the Environment Agency discharge limit, or upon loss of monitoring signal, or upon loss of electrical power. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | SIL 2 safety function per IEC 61511 risk graph — environmental consequence of uncontrolled above-limit discharge. 80% setpoint provides margin before absolute limit. 5-second closure time limits volume discharged above threshold to less than 1 litre at maximum discharge flow rate of 10 L/min. | subsystem, aetds, discharge-monitoring, safety, session-328, idempotency:sub-discharge-isolation-328
SUB-REQ-028 | The Discharge Monitoring and Control System SHALL measure and totalise discharge volume using electromagnetic flowmeters with accuracy of ±1% of reading, recording batch ID, volume, activity concentrations, date/time, and authorising operator identity for each discharge event. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | ±1% flow accuracy required for statutory discharge records submitted to Environment Agency. Operator identity recording required under nuclear site licence arrangements for formal authorisation chain. | subsystem, aetds, discharge-monitoring, session-328, idempotency:sub-discharge-flow-328
SUB-REQ-029 | The Inactive Drain Diversion System SHALL detect contamination above 0.1 Bq/mL gross beta-gamma in inactive drain flows and actuate motorised divert valves to route the affected drain to the active effluent collection tanks within 5 seconds of detection. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 0.1 Bq/mL threshold set at 10% of site discharge limit to provide early warning. 5-second divert time limits volume of contaminated water reaching the trade effluent system to less than 0.5 litres at typical drain flow rates, well within monitoring detection capability. | subsystem, aetds, inactive-drain, session-328, idempotency:sub-inactive-divert-328
SUB-REQ-030 | The Active Drain Collection Network SHALL be constructed from 316L stainless steel or borosilicate glass with double-contained routing through controlled areas, and SHALL withstand liquids in the pH range 1 to 13 with activity concentrations up to 1E6 Bq/L alpha and 1E8 Bq/L beta-gamma without material degradation over a 50-year design life. | subsystem-requirements | Subsystem Requirements |  | Analysis |  |  | pH 1-13 range covers concentrated acid digestion waste (HNO3, HCl) from radiochemical separations and alkaline decontamination solutions. 50-year design life aligns with nuclear facility decommissioning timeline. 316L SS and borosilicate glass are proven materials for nuclear effluent service per GDA assessments. | subsystem, aetds, drain-network, session-328, idempotency:sub-drain-materials-328
SUB-REQ-031 | The Chemical Treatment Plant SHALL collect and package treatment sludge containing co-precipitated actinides as intermediate-level waste in 200-litre drums compatible with the Radioactive Waste Management Facility storage and eventual disposal route. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | Sludge from ferric floc co-precipitation concentrates actinides into a small-volume ILW stream. 200L drum packaging aligns with UK ILW conditioning and storage standards (RWM Letter of Compliance process). Compatibility with downstream waste route is essential to avoid orphan waste. | subsystem, aetds, chemical-treatment, session-328, idempotency:sub-chem-sludge-328
SUB-REQ-032 | The Criticality Warning System SHALL detect a prompt criticality excursion producing a minimum absorbed dose rate of 20 mrad at the detector location within 1 millisecond of onset, using a minimum of two independent neutron-sensitive channels per monitored zone in a 2oo3 coincidence voting arrangement. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 20 mrad detection threshold derived from IAEA NS-R-5 and ANSI/ANS-8.3-1997 requirements for criticality alarm systems in facilities handling fissile materials. 2oo3 voting balances false alarm rate against detection reliability — a single-channel system would produce unacceptable false alarm frequency in a dockyard environment with neutron background from submarine reactor compartments nearby. | subsystem, safety, criticality, session-329, idempotency:sub-cws-detection-329
SUB-REQ-033 | The Criticality Warning System SHALL produce a distinctive warbling alarm tone audible at a minimum of 75 dBA in all occupied areas of the laboratory, distinguishable from fire, contamination, and general evacuation alarms, and not capable of being silenced locally. | subsystem-requirements | Subsystem Requirements |  | Demonstration |  |  | ONR Safety Assessment Principles ECS.3 and IAEA guidance require criticality alarms to be unmistakable and non-defeatable. 75 dBA threshold ensures audibility in laboratory areas with fume cupboard noise (typically 60-65 dBA). Local silencing would allow personnel to ignore the alarm during an event where immediate evacuation is the only safe response. | subsystem, safety, criticality, session-329, idempotency:sub-cws-alarm-329
SUB-REQ-034 | The Fire Detection and Suppression System SHALL achieve detection-to-alarm notification in less than 30 seconds for all fire scenarios within the laboratory, and SHALL initiate automatic suppression within 60 seconds of confirmed detection in zones where automatic suppression is enabled. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 30-second detection-to-alarm derived from BS 5839-1 Category L1 response time requirements. 60-second suppression initiation allows for detector confirmation (prevents single-point false discharge of IG-541, which would require zone evacuation) while limiting fire growth in a facility containing flammable solvents and radioactive materials. | subsystem, fire, session-329, idempotency:sub-fds-response-329
SUB-REQ-035 | The Fire Detection and Suppression System SHALL employ inert gas suppression (IG-541) in zones containing unsealed radioactive materials, alpha glove boxes, and sensitive analytical instruments, and SHALL employ water mist suppression in corridors, offices, and storage areas where water application does not risk contamination spread. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | Water applied to alpha-contaminated surfaces aerosolises plutonium particles and spreads contamination via runoff. IG-541 suppresses fire by oxygen displacement without disturbing contaminated surfaces or damaging sensitive detectors (HPGe, silicon alpha detectors). Zone boundary follows the C3/C4 contamination classification boundary. | subsystem, fire, session-329, idempotency:sub-fds-zones-329
SUB-REQ-036 | The Safety Interlock and Trip System SHALL enforce fissile material mass limits at each workstation using load cells and gamma activity monitors, initiating automatic process isolation when 80 percent of the single-contingency criticality safety limit is reached, with trip logic implemented in hardwired relay circuits achieving SIL 3 per IEC 61511. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Derives from SYS-REQ-009. 80 percent trip threshold provides margin before the actual criticality safety limit. Hardwired relay implementation required for SIL 3 because software-based systems introduce common-cause failure modes that cannot be claimed below 1E-4 PFDavg without extensive software V and V. Load cells plus gamma monitors provide diverse measurement principles. | subsystem, safety, criticality, session-329, idempotency:sub-sit-fissile-329
SUB-REQ-037 | The Safety Interlock and Trip System SHALL use 2oo3 voting logic for all safety-critical trip functions, with each voting channel powered from an independent UPS circuit, and SHALL complete trip initiation within 500 milliseconds of the trip condition being met. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 2oo3 voting allows one channel to fail or be taken offline for maintenance without losing the safety function or causing a spurious trip. 500 ms trip initiation bounds the time between detection and actuation, ensuring criticality and fire trips are fast enough to be effective. Independent UPS circuits prevent single power failure from disabling voting majority. | subsystem, safety, session-329, idempotency:sub-sit-voting-329
SUB-REQ-038 | The Emergency Power System SHALL provide uninterruptible power to all criticality detectors, fire detection panels, safety interlock circuits, and radiation monitors via online double-conversion UPS with a minimum of 30 minutes full-load battery autonomy and 4 hours reduced-load autonomy. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 30-minute full-load autonomy covers diesel generator start and synchronisation including failed-start retry. 4-hour reduced-load provides overnight coverage if diesel fails to start. Online double-conversion ensures zero transfer time — even momentary power interruption to criticality detectors creates an unmonitored gap during which a criticality event could occur undetected. | subsystem, power, safety, session-329, idempotency:sub-eps-ups-329
SUB-REQ-039 | The Emergency Power System SHALL start the standby diesel generator and achieve rated voltage and frequency within 8 seconds of mains power failure, with automatic load transfer to safety-critical ventilation extract fans maintaining containment depression. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 8-second start-to-load derived from the maximum acceptable loss-of-containment duration. Extract fans maintain depression cascade preventing airborne contamination migration between zones. Without extract, C4 zone pressure equalises with C3 within approximately 15 seconds, so 8-second diesel start plus transfer maintains depression before equalisation occurs. | subsystem, power, session-329, idempotency:sub-eps-diesel-329
SUB-REQ-040 | The Spill Containment and Emergency Decontamination System SHALL provide bunded containment under all fume cupboards and glove boxes rated to 110 percent of the largest vessel volume, with stainless steel construction resistant to concentrated nitric acid and sealed joints preventing underfloor seepage to inactive drains. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | 110 percent containment volume is standard for nuclear bunded areas per Environment Agency guidance. Nitric acid resistance required because radiochemical separations use concentrated HNO3 for plutonium dissolution. Sealed joints prevent contaminated liquid reaching inactive (non-monitored) drainage, which would bypass the active effluent treatment system and risk unmonitored discharge. | subsystem, containment, session-329, idempotency:sub-scd-bund-329
SUB-REQ-041 | The Emergency Communications and Alarm System SHALL provide a minimum of four distinct alarm tones: criticality (warbling siren), fire (two-tone), contamination (intermittent), and general evacuation (continuous), each audible at 75 dBA minimum in all occupied zones including hearing protection areas via visual beacon backup. | subsystem-requirements | Subsystem Requirements |  | Demonstration |  |  | Four distinct tones required because each emergency type demands a different response: criticality requires immediate evacuation away from the source, fire requires orderly evacuation via designated routes, contamination requires shelter-in-place or controlled exit, and general evacuation covers all other scenarios. Visual beacons ensure alerting in hearing protection zones where audiometric thresholds may exceed 75 dBA. | subsystem, communications, session-329, idempotency:sub-eca-tones-329
SUB-REQ-042 | The Solid Waste Characterization and Segregation Station SHALL measure the gamma-emitting radionuclide inventory of each waste item using a shielded detector enclosure with a minimum detectable activity of 10 Bq for Cs-137 and Co-60 in a 200L drum geometry with a counting time of not more than 30 minutes. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Waste categorisation into VLLW/LLW/ILW per the Radioactive Substances Regulation 2011 requires quantified activity inventories. The 10 Bq MDA for Cs-137 in a drum geometry ensures reliable discrimination between VLLW (<4 Bq/g threshold) and LLW. The 30-minute limit supports operational throughput of approximately 10 drums per shift. | subsystem, waste-management, characterization, session-330, idempotency:sub-swcss-gamma-mda-330
SUB-REQ-043 | The Solid Waste Characterization and Segregation Station SHALL perform alpha and beta surface contamination screening on each waste package with detection limits of 0.04 Bq/cm2 for alpha and 0.4 Bq/cm2 for beta, recording results against the package unique identifier. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Surface contamination limits for waste packages at UK nuclear sites are derived from IRR 2017 and site-specific dose uptake assessments. These detection limits are one-tenth of the surface contamination clearance values, ensuring reliable compliance verification before packages leave the characterization area. | subsystem, waste-management, characterization, session-330, idempotency:sub-swcss-surface-330
SUB-REQ-044 | The Liquid Waste Conditioning System SHALL achieve a volume reduction factor of at least 10:1 for radioactive liquid concentrates through thin-film evaporation, with a decontamination factor of at least 1000 for non-volatile radionuclides in the distillate returned to the active drain system. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Volume reduction of 10:1 is necessary to minimise the number of cemented wasteform packages generated from liquid waste concentrates. The DF of 1000 ensures distillate activity is below the site discharge authorisation limits for return to the AETDS, preventing re-concentration of radioactivity in the effluent treatment loop. | subsystem, waste-management, liquid-conditioning, session-330, idempotency:sub-lwcs-evaporator-330
SUB-REQ-045 | The Liquid Waste Conditioning System SHALL immobilise concentrated liquid waste into a cement-based wasteform achieving a compressive strength of at least 4 MPa after 28 days curing, and a leach rate of less than 10^-5 g/cm2/day for Cs-137 under standard ANSI/ANS 16.1 test conditions. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Compressive strength of 4 MPa is the RWM Waste Acceptance Criteria minimum for ILW packages destined for geological disposal. The leach rate requirement ensures the wasteform provides adequate long-term containment of mobile fission products during interim storage and post-closure. Without these quantified criteria, packages may be rejected at the disposal facility. | subsystem, waste-management, liquid-conditioning, session-330, idempotency:sub-lwcs-cement-330
SUB-REQ-046 | The Solid Waste Packaging and Compaction System SHALL compact dry active waste in 200L drums using a hydraulic compactor with a minimum force of 50 kN, achieving a volume reduction ratio of at least 3:1 for compactible waste streams. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Volume reduction of 3:1 is the minimum economic justification for mechanical compaction at this facility scale, reducing interim storage requirements by approximately 200 drum positions over the facility lifetime. The 50 kN force is sufficient for laboratory dry active waste (paper, plastics, PPE) without requiring the higher forces needed for metallic waste. | subsystem, waste-management, packaging, session-330, idempotency:sub-swpcs-compaction-330
SUB-REQ-047 | The Solid Waste Packaging and Compaction System SHALL verify that the external surface contamination of each sealed drum does not exceed 4 Bq/cm2 for beta/gamma emitters and 0.4 Bq/cm2 for alpha emitters before release to the Interim Waste Store. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | These are the UK transport regulation surface contamination limits for Type A packages under CDG 2009. Any drum exceeding these limits cannot be transferred to storage or transport without further decontamination. Verification at the packaging stage prevents contaminated drums entering the store where detection and remediation are more difficult. | subsystem, waste-management, packaging, session-330, idempotency:sub-swpcs-surface-verify-330
SUB-REQ-048 | The Interim Waste Store SHALL maintain a minimum spacing of 300 mm between fissile waste packages in all directions, with physical spacer frames preventing closer approach, such that the effective neutron multiplication factor (keff) does not exceed 0.95 for any credible arrangement including seismic displacement and flooding scenarios. | subsystem-requirements | Subsystem Requirements |  | Analysis |  |  | keff limit of 0.95 is the UK standard criticality safety criterion for waste storage (ONR Safety Assessment Principles, para 573). The 300 mm spacing is derived from criticality safety assessment for worst-case fissile loadings in 500L drums containing PuO2-contaminated waste at the maximum permitted mass per drum. Physical spacer frames are required because administrative controls alone are insufficient for multi-decade storage. | subsystem, waste-management, interim-store, safety, session-330, idempotency:sub-iws-criticality-330
SUB-REQ-049 | The Interim Waste Store SHALL continuously monitor the storage environment for temperature (range 5-35 degC, alarm at 40 degC), relative humidity (alarm at 80% RH), and area gamma dose rate (alarm at 25 microSv/h), with all parameters logged at 15-minute intervals and retained for the lifetime of stored packages. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Temperature and humidity limits are derived from RWM guidance on interim storage conditions to prevent corrosion of mild steel drums and degradation of cement wasteforms. The 40 degC alarm prevents thermally-driven degradation. The 25 microSv/h gamma alarm is set at 50% of the 2 mSv/year public dose limit budget allocated to the store, providing early warning of shielding degradation or package failure. | subsystem, waste-management, interim-store, session-330, idempotency:sub-iws-envmon-330
SUB-REQ-050 | The Interim Waste Store SHALL provide storage capacity for a minimum of 200 standard waste packages (200L and 500L drums) with provision for a 20% expansion margin, and SHALL support package placement and retrieval using an overhead crane rated for the maximum credible package weight of 2000 kg. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | 200-package capacity is based on projected waste arisings of approximately 15 ILW drums and 40 LLW drums per year over the 10-year period between scheduled off-site transfers. The 20% margin provides operational flexibility for delayed transfers. The 2000 kg crane rating covers the heaviest credible package: a 500L drum containing cemented alpha-bearing sludge at maximum fill. | subsystem, waste-management, interim-store, session-330, idempotency:sub-iws-capacity-330
SUB-REQ-051 | The Waste Records and Consignment System SHALL create and maintain a waste package data record for each package containing: unique identifier, waste stream origin, radionuclide inventory (activity and uncertainty), waste form description, dose rate at 1 m, surface contamination measurements, package weight, waste category, and storage location, in a format compliant with RWM Data Recording Requirements. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | The RWM Waste Package Data Record is a mandatory regulatory deliverable for any package destined for geological disposal. Without complete and compliant records, packages cannot be consigned. The specified data fields are the minimum set required by the RWM Letter of Compliance process. Records must be created at packaging time and updated throughout the storage period. | subsystem, waste-management, records, session-330, idempotency:sub-wrcs-wpdr-330
SUB-REQ-052 | The Waste Records and Consignment System SHALL maintain a nuclear material accountancy ledger tracking all fissile material (U-235, Pu-239, Pu-241) transfers into, within, and out of the waste management facility, reconciled with the site-level nuclear material accountancy system within 24 hours of any material movement. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | ONR Licence Condition 14 (nuclear material accountancy) requires accurate tracking of all fissile material at licensed nuclear sites. The 24-hour reconciliation window ensures the site accountancy system reflects the current fissile inventory distribution for criticality safety and safeguards purposes. Delays beyond 24 hours risk undetected discrepancies that could compromise either criticality safety arguments or IAEA safeguards compliance. | subsystem, waste-management, records, safety, session-330, idempotency:sub-wrcs-nma-330
SUB-REQ-053 | The Sample Receiving Bay SHALL maintain negative pressure of at least minus 15 Pa relative to the adjacent corridor (C2 zone boundary) during all sample transfer operations, with automatic interlock preventing simultaneous opening of inner and outer airlock doors. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Prevents migration of airborne contamination from the C3 receiving area to uncontrolled C2 corridors during sample delivery. The 15 Pa threshold matches the facility depression cascade setpoints. Interlock prevents containment breach via airlock short-circuit — a credible contamination event given high sample throughput of up to 50 containers per day. | subsystem, sample-receipt, session-331, idempotency:sub-sample-receipt-airlock-331
SUB-REQ-054 | The Dose Rate and Contamination Screening Station SHALL measure ambient dose equivalent rate H*(10) from 0.05 microSv/h to 10 mSv/h with energy response within plus or minus 30 percent across 50 keV to 3 MeV, and SHALL categorise each sample as contact-handled (below 2 mSv/h at surface) or remote-handled (above 2 mSv/h) within 5 minutes of sample presentation. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | The 2 mSv/h contact-handled threshold is the IRR17 boundary for requiring controlled handling procedures. Energy response specification ensures accurate measurement of the mixed gamma fields from reactor-origin samples (Co-60, Cs-137, Mn-54). The 5-minute turnaround prevents bottlenecks at the receiving bay given peak throughput of 50 containers per day. | subsystem, sample-receipt, session-331, idempotency:sub-sample-receipt-doserate-331
SUB-REQ-055 | The Sample Registration and Chain-of-Custody System SHALL assign a unique barcode identifier to each sample within 60 seconds of receipt, recording sample origin, requested analyses, date-time, receiving operator, and external dose rate, with all custody transfer events timestamped and attributable to named individuals in an auditable database. | subsystem-requirements | Subsystem Requirements |  | Demonstration |  |  | Chain-of-custody integrity is a regulatory requirement under ONR licence conditions and is essential for evidential samples used in nuclear material accountancy. The 60-second registration target prevents queuing at the receiving bay. Named-individual attribution supports the ALARP demonstration by tracking who handled what and when. | subsystem, sample-receipt, session-331, idempotency:sub-sample-receipt-registration-331
SUB-REQ-056 | The Sample Preparation Laboratory SHALL perform acid digestion of solid samples using concentrated HNO3, HCl, and HF in PTFE-lined microwave or hotplate dissolution vessels rated to 250 degC and 100 bar, achieving complete dissolution of reactor structural materials (stainless steel, Inconel, Zircaloy) with dissolution completion verified by visual inspection and gravimetric residue measurement below 0.1 percent of original sample mass. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Complete dissolution is essential for representative radiochemical analysis of reactor decommissioning samples. Incomplete dissolution preferentially retains refractory actinides (Pu in particular) leading to underestimation of waste classification. The 0.1% gravimetric threshold is the standard acceptance criterion used by UKAS-accredited nuclear analytical laboratories. | subsystem, sample-receipt, session-331, idempotency:sub-sample-prep-dissolution-331
SUB-REQ-057 | The Sample Preparation Laboratory SHALL add calibrated radiotracer spikes (Am-243, Pu-242, Sr-85, Cs-134) to each sample aliquot before chemical separation, with spike activities calibrated against NPL-traceable standards and dispensed using calibrated micropipettes with accuracy of plus or minus 1 percent of stated volume. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | Radiotracer spiking enables yield correction for the subsequent radiochemical separation steps. Without tracers, chemical recovery losses cannot be quantified, producing unreliable results. NPL traceability is a UKAS accreditation requirement for quantitative radiochemistry. The 1% volumetric accuracy ensures tracer-to-analyte ratio uncertainty does not dominate the overall measurement uncertainty budget. | subsystem, sample-receipt, session-331, idempotency:sub-sample-prep-tracer-331
SUB-REQ-058 | The Sample Storage and Archive Facility SHALL maintain criticality-safe geometry with a minimum spacing of 250 mm between fissile sample containers in all directions using fixed neutron-absorbing cadmium-lined shelving, such that keff does not exceed 0.95 for any credible arrangement including optimum moderation and reflection, with a maximum fissile mass inventory of 200 g Pu-239 equivalent across all stored samples. | subsystem-requirements | Subsystem Requirements |  | Analysis |  |  | Criticality safety is the overriding nuclear safety concern for a sample storage area holding multiple fissile-bearing samples from reactor decommissioning. The 200g Pu-eq limit and 0.95 keff criterion derive from the site criticality safety case assessment, with cadmium-lined shelving providing passive geometric safety independent of administrative controls. This is consistent with the IAEA criticality safety standards for storage arrays. | subsystem, sample-receipt, session-331, idempotency:sub-sample-storage-crit-331
SUB-REQ-059 | The Sample Storage and Archive Facility SHALL maintain a storage environment of 15 to 25 degC and below 60 percent relative humidity, with continuous monitoring at 15-minute intervals and alarm at 30 degC or 70 percent RH, and SHALL provide retention capacity for a minimum of 500 sample containers with a 2-year routine retention period and indefinite retention for evidential or legal-hold samples. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Environmental control prevents sample degradation (acid solution evaporation, biological growth in aqueous samples, corrosion of container closures) during extended storage. The 500-container capacity is derived from the projected 200 samples per week throughput with 2-year retention. Evidential samples supporting nuclear material accountancy or regulatory investigations must be retained indefinitely until released by the regulator. | subsystem, sample-receipt, session-331, idempotency:sub-sample-storage-env-331
SUB-REQ-060 | The Acid Digestion and Dissolution Station SHALL achieve complete dissolution of reactor structural materials (austenitic stainless steel, Inconel 600/625, Zircaloy-2/4) with verified residue below 0.1 percent of original sample mass, using sequential acid additions in PTFE-lined vessels rated to 250 degC and 100 bar, with all acid additions performed within perchloric-acid-rated fume cupboards equipped with washdown systems. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Complete dissolution is prerequisite for quantitative radiochemical analysis. Incomplete dissolution preferentially retains refractory actinides (Pu, Am, Cm are co-precipitated with undissolved metal oxides), leading to systematic underestimation of alpha-emitting inventory — a criticality safety and waste classification concern. Perchloric acid fume cupboards are mandatory when HClO4 is used as a final oxidising acid for organic destruction in Pu-bearing solutions, as dry perchlorate residues are explosive. | subsystem, radchem-sep, session-331, idempotency:sub-radchem-acid-diss-331
SUB-REQ-061 | The Actinide Separation Chemistry Station SHALL achieve chemical recovery of at least 70 percent for Pu-239/240, Am-241, and Cm-244 through the complete separation procedure, with recovery determined by radiotracer yield measurement, and SHALL produce source-ready actinide fractions with a spectral purity sufficient for alpha spectrometry resolution of Pu-239/240 from Pu-238 (minimum peak separation of 40 keV FWHM). | subsystem-requirements | Subsystem Requirements |  | Test |  |  | The 70% recovery threshold is the minimum acceptable for quantitative analysis — below this, measurement uncertainties become unacceptably large for waste categorisation decisions. Spectral purity is essential because Pu-239 (5.16 MeV) and Pu-238 (5.50 MeV) differ by only 340 keV; incomplete separation of interfering isotopes (U, Th, Am) produces spectral overlap that prevents accurate activity ratio determination needed for isotopic characterisation of submarine reactor fuel. | subsystem, radchem-sep, session-331, idempotency:sub-radchem-actinide-recovery-331
SUB-REQ-062 | The Strontium and Caesium Separation Station SHALL isolate Sr-90 from dissolved reactor samples with a decontamination factor of at least 1000 for calcium and at least 10000 for other beta-emitting radionuclides, and SHALL achieve chemical recovery of at least 60 percent for strontium as determined by Sr-85 radiotracer yield. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Sr-90 measurement by Cherenkov counting or LSC of the Y-90 daughter requires high radiochemical purity because other beta-emitters (Cs-137, Ce-144, Ru-106) interfere directly. The calcium decontamination factor is critical because natural calcium in the sample matrix competes with strontium on the crown ether resin, reducing both recovery and selectivity. These DFs are established requirements for UKAS-accredited Sr-90 methods in the nuclear industry. | subsystem, radchem-sep, session-331, idempotency:sub-radchem-sr-sep-331
SUB-REQ-063 | The Tritium and Carbon-14 Preparation Station SHALL recover tritium from aqueous reactor coolant samples by azeotropic distillation with a recovery efficiency of at least 95 percent, and SHALL trap evolved CO2 from acid-digested samples using NaOH bubblers with a trapping efficiency of at least 98 percent for C-14 determination, with all distillation and trapping apparatus operated within a dedicated fume cupboard with molecular sieve extract filtration to prevent tritiated water vapour discharge. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Tritium is the dominant radionuclide by activity in PWR coolant and must be accurately measured for discharge accounting. The 95% recovery threshold ensures measurement uncertainty remains below 10% for typical dockyard coolant samples. The molecular sieve extract filtration is required because tritiated water vapour would otherwise pass through HEPA filters (which capture only particulates) and contribute directly to the site gaseous discharge authorisation. | subsystem, radchem-sep, session-331, idempotency:sub-radchem-tritium-331
SUB-REQ-064 | The Separation Chemistry Fume Cupboard Array SHALL provide a minimum of six standard mineral-acid-rated fume cupboards and two perchloric-acid-rated fume cupboards, each maintaining a face velocity of 0.5 m/s at the 500 mm working sash opening, with continuous airflow monitoring and automatic sash closure within 5 seconds of extract airflow falling below 80 percent of nominal. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Eight fume cupboards is the minimum for concurrent processing of the three analyte group separations plus acid digestion without workflow bottlenecks. Two perchloric-rated cupboards (with stainless steel washdown ductwork) are required specifically for the HClO4 final oxidation step in actinide separations — this cannot be performed in standard cupboards due to explosive perchlorate residue risk. Automatic sash closure on extract failure is a defence-in-depth measure preventing operator exposure if the extract fan trips during active work. | subsystem, radchem-sep, session-331, idempotency:sub-radchem-fume-cupboards-331
SUB-REQ-065 | The HPGe Detector Array SHALL provide energy resolution of 2.0 keV FWHM or better at 1332 keV (Co-60) for each co-axial detector and 1.2 keV FWHM or better at 122 keV (Co-57) for each well-type detector, measured under operational cryocooler vibration conditions. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Energy resolution directly determines the ability to resolve gamma peaks in complex fission product spectra from irradiated submarine reactor components. 2.0 keV at 1332 keV is needed to separate closely spaced peaks such as Eu-152/154 multiplets and to deconvolve Co-60 sum peaks from activation products. Degraded resolution causes systematic positive bias in nuclide activity quantification. | subsystem, gamma-spectrometry, session-333, idempotency:sub-gamma-hpge-resolution-333
SUB-REQ-066 | The HPGe Detector Array SHALL include a minimum of 6 co-axial detectors with relative efficiency of 40 percent or greater and 2 well-type detectors with absolute efficiency of 5 percent or greater at 1332 keV, providing sufficient counting capacity for a throughput of 80 samples per day during submarine refit campaigns. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Detector count and efficiency are sized to the peak demand during submarine refit, when 80+ environmental and structural material samples per day require gamma screening within 48-hour turnaround. Well-type detectors provide the high efficiency needed for low-activity environmental monitoring samples. | subsystem, gamma-spectrometry, session-333, idempotency:sub-gamma-hpge-efficiency-333
SUB-REQ-067 | The Lead Shielding and Counting Chambers SHALL reduce the integral background count rate in the 50 to 2000 keV energy range to less than 1 count per second per detector, using graded-Z shielding of minimum 100 mm lead, 1 mm cadmium, and 1 mm copper. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Background reduction factor of greater than 1000 is required to achieve the 0.5 Bq/L MDA for Cs-137 specified in SYS-REQ-001. Without adequate shielding, environmental background from the nearby dockyard (activation products, cosmic ray interactions) would dominate the spectrum below 500 keV, increasing counting times to impractical durations. | subsystem, gamma-spectrometry, session-333, idempotency:sub-gamma-shielding-background-333
SUB-REQ-068 | The Detector Cryogenic Cooling System SHALL maintain each HPGe detector crystal temperature at 85 K plus or minus 2 K continuously, with automatic shutdown and crystal protection initiated within 30 seconds of temperature exceeding 100 K. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | HPGe detectors require cryogenic temperatures to achieve semiconductor charge carrier mobility needed for spectroscopic-grade resolution. Above 100 K, charge trapping degrades resolution irreversibly within minutes. The 30-second shutdown protects detector investment of approximately 30000 GBP per crystal. | subsystem, gamma-spectrometry, session-333, idempotency:sub-gamma-cryo-temp-333
SUB-REQ-069 | The Digital MCA and Signal Processing System SHALL provide a minimum of 16384 channels per detector with integral nonlinearity of less than 0.05 percent and dead-time correction accurate to within 1 percent at input count rates up to 50000 counts per second. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 16384 channels at 0.05 percent linearity yield approximately 0.12 keV per channel over the 0-2000 keV analysis range, well below the 2 keV detector resolution and sufficient for peak centroid accuracy required by nuclide identification algorithms. 50 kcps dead-time correction handles high-activity submarine reactor component samples without systematic activity underestimation. | subsystem, gamma-spectrometry, session-333, idempotency:sub-gamma-mca-channels-333
SUB-REQ-070 | The Gamma Spectroscopy Analysis and QA Workstation SHALL perform automated nuclide identification against a validated library containing a minimum of 300 nuclides including all fission products, activation products, and actinides relevant to submarine reactor systems, with false positive identification rate of less than 1 percent. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | The nuclide library must cover the full spectrum of isotopes encountered in submarine PWR decommissioning: fission products (Cs-137, Sr-90 via Ba-137m), activation products (Co-60, Fe-55, Ni-63, Mn-54), transuranics (Pu-241 via Am-241 in-growth), and environmental tracers. False positives would trigger unnecessary contamination investigations costing approximately 2000 GBP each. | subsystem, gamma-spectrometry, session-333, idempotency:sub-gamma-analysis-library-333
SUB-REQ-071 | The Gamma Spectroscopy Analysis and QA Workstation SHALL perform automated daily quality control checks on each detector using a Eu-152 reference source, verifying energy calibration within 0.3 keV of reference values and resolution within 10 percent of baseline, and SHALL generate control chart trend data for UKAS accreditation review. | subsystem-requirements | Subsystem Requirements |  | Demonstration |  |  | ISO 17025 and UKAS requirements for accredited nuclear measurements mandate daily QC with documented control charts. Eu-152 is the standard reference source because it provides peaks spanning the full analysis range (122 to 1408 keV). Drift beyond 0.3 keV causes nuclide misidentification; resolution degradation beyond 10 percent indicates detector or electronics failure. | subsystem, gamma-spectrometry, session-333, idempotency:sub-gamma-qc-daily-333
SUB-REQ-072 | The Alpha Detector Chamber Array SHALL achieve energy resolution of 25 keV FWHM or better at 5.486 MeV (Am-241) for each PIPS detector, with minimum detectable activity of 0.5 mBq per sample for Pu-239/240 using a 250000-second counting time. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | 25 keV resolution is required to resolve the Pu-238 (5.499 MeV) and Am-241 (5.486 MeV) alpha peaks which are separated by only 13 keV. The 0.5 mBq MDA derives from SYS-REQ-004 and is needed to demonstrate clearance of submarine structural materials against regulatory limits. | subsystem, alpha-spectrometry, session-333, idempotency:sub-alpha-detector-resolution-333
SUB-REQ-073 | The Electrodeposition and Source Preparation Station SHALL produce alpha sources with areal density of less than 50 micrograms per cm2 on 25 mm diameter stainless steel planchets, with chemical yield recovery for the electrodeposition process of at least 90 percent for Am-243 and Pu-242 tracers. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Source thickness directly determines alpha peak tailing and energy resolution degradation. Above 50 micrograms per cm2, self-absorption broadens peaks beyond 30 keV FWHM, preventing resolution of Pu-238 from Am-241. The 90 percent tracer recovery ensures acceptable measurement uncertainty — below 70 percent recovery, uncertainty on the reported activity exceeds the 20 percent target required by the client. | subsystem, alpha-spectrometry, session-333, idempotency:sub-alpha-electrodep-thickness-333
SUB-REQ-074 | The Alpha Spectrometry Vacuum System SHALL achieve and maintain a chamber pressure of less than 5 Pa within 3 minutes of pump-down initiation, with automatic acquisition inhibit when chamber pressure exceeds 10 Pa during counting. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Alpha particles lose approximately 1 MeV per cm of travel in air at atmospheric pressure. At 5 Pa, energy loss is negligible over the 5-25 mm source-detector distance. Exceeding 10 Pa during counting would degrade energy resolution beyond 30 keV, invalidating the spectrum for peak resolution of closely-spaced actinide lines. | subsystem, alpha-spectrometry, session-333, idempotency:sub-alpha-vacuum-pressure-333
SUB-REQ-075 | The Alpha MCA and Spectral Processing Unit SHALL calculate tracer recovery, isotopic activity ratios, activity per sample, combined measurement uncertainty at k equals 2, and minimum detectable activity for each sample, with results traceable to NIST or NPL-certified tracer solutions. | subsystem-requirements | Subsystem Requirements |  | Demonstration |  |  | Alpha spectrometry quantification requires known-activity tracer isotopes spiked before separation to correct for chemical recovery losses. Without automated tracer recovery calculation, manual errors in activity computation would undermine ISO 17025 accreditation. Traceability to NIST or NPL ensures results are defensible for regulatory submissions. | subsystem, alpha-spectrometry, session-333, idempotency:sub-alpha-mca-tracer-333
SUB-REQ-076 | The Alpha Spectrometry Vacuum System SHALL incorporate a cold trap between counting chambers and the vacuum pump to prevent actinide contamination of pump oil, and SHALL exhaust pump discharge through a HEPA filter to the active extract ductwork. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | Actinide-contaminated pump oil creates a secondary radioactive waste stream and poses an inhalation hazard during pump maintenance. The cold trap captures volatile actinide compounds that may sublime from high-activity sources during long counting periods. HEPA-filtered exhaust prevents airborne release of alpha-emitting particles. | subsystem, alpha-spectrometry, session-333, idempotency:sub-alpha-vacuum-containment-333
SUB-REQ-077 | The Liquid Scintillation Counter Array SHALL achieve a counting efficiency of at least 65% for H-3 and at least 95% for C-14 across the operational quench range (tSIE 200-800), with background count rate not exceeding 2.0 CPM in the H-3 window and 4.0 CPM in the C-14 window. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | These efficiency and background values derive from the SYS-REQ-010 detection limit of 1 Bq/L H-3 at 120 min count time. MDA = 2.71 + 4.65*sqrt(B*t) / (E*V*t) requires E>=0.65 and B<=2.0 CPM to achieve 1 Bq/L. The C-14 figures derive from equivalent calculations for the C-14 window. | subsystem, liquid-scintillation, session-334, idempotency:sub-lsc-counter-efficiency-334
SUB-REQ-078 | The Liquid Scintillation Counter Array SHALL provide pulse shape analysis alpha/beta discrimination with an alpha rejection factor of at least 99.95% and beta spillover into the alpha channel not exceeding 0.1%, verified using mixed Am-241/H-3 reference standards. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Dockyard samples contain alpha-emitting actinides alongside H-3 and C-14. Without effective PSA, alpha events in the beta counting window produce false-positive activity, particularly problematic for low-level H-3 measurements where alpha interference can exceed the H-3 signal. The 99.95% rejection factor ensures alpha contribution to the H-3 window is negligible (<0.1% of MDA). | subsystem, liquid-scintillation, session-334, idempotency:sub-lsc-psa-334
SUB-REQ-079 | The LSC Sample Preparation Station SHALL achieve cocktail dispensing repeatability of 0.5% RSD for volumes between 10 and 15 mL, with sample-to-cocktail ratio configurable between 1:1 and 1:3 to accommodate varying sample salinity and pH. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Cocktail volume consistency directly affects quench level reproducibility and counting geometry. A 1% volume variation translates to approximately 2% variation in quench-corrected activity. The 0.5% RSD target ensures the cocktail dispensing contribution to total measurement uncertainty is below 0.3%, keeping it negligible compared to counting statistics (~3-5% at MDA). | subsystem, liquid-scintillation, session-334, idempotency:sub-lsc-cocktail-334
SUB-REQ-080 | The LSC Sample Preparation Station SHALL provide temperature-controlled dark adaptation storage for a minimum of 200 vials simultaneously, maintaining temperature at 15 +/- 1 degrees C for a minimum dark adaptation period of 2 hours before counting, to reduce chemiluminescence and photoluminescence below 0.5 CPM contribution. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Chemiluminescence from cocktail-sample chemical interaction and photoluminescence from ambient light exposure produce time-dependent false counts that decay over 1-4 hours. At 15C the decay constant is approximately 30 minutes. A 2-hour period at 15C reduces the luminescence contribution to below 0.5 CPM, which is 25% of the H-3 background target. Higher temperatures accelerate chemical luminescence onset. The 200-vial capacity supports batch processing during dockyard campaigns. | subsystem, liquid-scintillation, session-334, idempotency:sub-lsc-dark-adapt-334
SUB-REQ-081 | The LSC Background and QC Reference System SHALL perform automated daily QC checks using sealed H-3 and C-14 reference standards traceable to NPL, with chi-squared test at 95% confidence level and trending analysis detecting systematic drift exceeding 2 sigma over 20 consecutive measurements. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | ISO 17025 and Environment Agency MCERTS requirements mandate regular instrument QC. Daily QC with sealed standards detects detector degradation, electronics drift, and environmental changes before they affect analytical results. The chi-squared test at 95% confidence provides statistically rigorous pass/fail criteria. Two-sigma trending over 20 points detects slow systematic drift that individual chi-squared tests miss. | subsystem, liquid-scintillation, session-334, idempotency:sub-lsc-qc-daily-334
SUB-REQ-082 | The LSC Data Analysis and QA Workstation SHALL calculate total measurement uncertainty for each result per GUM methodology, incorporating counting statistics, quench correction uncertainty, background subtraction uncertainty, volumetric uncertainty, and decay correction uncertainty, reporting combined standard uncertainty and expanded uncertainty at 95% coverage factor. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Regulatory reporting of environmental radioactivity requires uncertainty budgets per GUM (ISO/IEC Guide 98-3). EA discharge authorisation conditions require reported uncertainties. Without propagated uncertainty budgets, results cannot be used for compliance demonstration. Each component contributes differently depending on activity level — at MDA counting statistics dominate, at higher activities volumetric and quench correction dominate. | subsystem, liquid-scintillation, session-334, idempotency:sub-lsc-uncertainty-334
SUB-REQ-083 | The ICP-MS Sample Introduction System SHALL achieve inter-sample carryover of less than 0.01% for uranium and less than 0.005% for plutonium, verified using a 10 ppb rinse-down protocol with 5% HNO3 and 0.05% Triton X-100, to prevent cross-contamination between high-activity dockyard samples and environmental blanks. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Dockyard samples span 6+ orders of magnitude in actinide concentration. Memory effects from high-level samples contaminating subsequent low-level measurements produce false positives. The 0.01% U carryover ensures a 10 ppb sample does not contribute more than 1 fg/L to the next measurement, which is below the instrument detection limit. Plutonium has stricter limits (0.005%) because Pu measurements are used for criticality safety assessments. | subsystem, icp-ms, session-334, idempotency:sub-icpms-memory-334
SUB-REQ-084 | The ICP-MS Instrument SHALL achieve isotope ratio measurement precision of 0.5% RSD for U-235/U-238 and 2% RSD for Pu-239/Pu-240 at concentrations exceeding 10 times the detection limit, with mass bias correction using NIST SRM 3164 (uranium) or CRM 136 (plutonium) certified reference materials. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Uranium enrichment determination (U-235/U-238) for dockyard decommissioning waste characterisation requires 0.5% RSD to distinguish natural uranium (0.0072) from reactor-grade (0.03-0.05) confidently. Pu-239/Pu-240 ratio distinguishes weapons-grade from reactor-grade plutonium, critical for waste categorisation and safeguards. The 2% RSD reflects the greater difficulty of Pu ratio measurements due to lower concentrations and UH+ interference. | subsystem, icp-ms, session-334, idempotency:sub-icpms-isotope-ratio-334
SUB-REQ-085 | The Argon Gas Supply and Plasma System SHALL maintain plasma gas flow stability within 0.5% of setpoint and nebuliser gas flow stability within 1% of setpoint over an 8-hour analytical run, with automatic gas purity monitoring that triggers instrument standby when impurity levels (O2, H2O, N2) exceed 5 ppm total. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Plasma gas flow variations directly affect plasma temperature and ionisation efficiency. A 1% change in nebuliser gas flow causes approximately 5% change in signal intensity and alters oxide formation ratios. The 0.5% stability target keeps signal drift below 3% over 8 hours, within internal standard correction capability. Gas purity monitoring prevents torch damage and spectral interference from N2 and O2 molecular species. | subsystem, icp-ms, session-334, idempotency:sub-icpms-argon-334
SUB-REQ-086 | The LIMS Server and Database SHALL maintain operational availability of at least 99.5% during laboratory working hours (0700-1900 weekdays), with automatic failover to the standby server completing within 60 seconds of primary server failure detection, and zero data loss for committed transactions. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Laboratory operations depend on LIMS for sample tracking and result recording. Extended downtime forces paper-based backup procedures that increase error risk. The 99.5% target allows approximately 2.5 hours unplanned downtime per quarter, which is operationally acceptable. The 60-second failover prevents work interruption. Zero data loss for committed transactions ensures no approved results are lost during failover. | subsystem, lims, session-334, idempotency:sub-lims-availability-334
SUB-REQ-087 | The LIMS Instrument Interface Module SHALL support bidirectional data exchange with all analytical instrument workstations (gamma spectrometry, alpha spectrometry, liquid scintillation, and ICP-MS), pushing work orders and counting/analysis parameters to instruments and receiving validated results back, with automatic data integrity verification using checksums and duplicate detection. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Bidirectional interface eliminates manual transcription at both ends: work order details do not need to be re-entered at the instrument, and results do not need to be manually typed into LIMS. Each manual transcription point has an estimated error rate of 0.1-0.5% for nuclear analytical data. With 50,000 results per year, unidirectional-only interfaces would generate 50-250 transcription errors annually. Checksums prevent silent data corruption during transfer. | subsystem, lims, session-334, idempotency:sub-lims-instrument-ifc-334
SUB-REQ-088 | The LIMS Reporting and Regulatory Compliance Engine SHALL automatically calculate cumulative site discharges against annual authorised limits for all controlled radionuclides, generating alerts when cumulative discharge reaches 75% and 90% of the authorised limit for any radionuclide, and producing quarterly RIFE programme submissions in the format specified by EA/SEPA. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Exceeding authorised discharge limits is a regulatory offence under the Environmental Permitting Regulations. Automated cumulative tracking with early warning at 75% and 90% provides laboratory management and dockyard radiological protection sufficient lead time to implement discharge reduction measures. Manual tracking of cumulative discharges across multiple radionuclides is error-prone and has historically led to reportable events at other nuclear sites. | subsystem, lims, session-334, idempotency:sub-lims-regulatory-334
SUB-REQ-089 | The Electrical Power Distribution System SHALL maintain normal supply availability of at least 99.5% per annum, measured at the main low-voltage switchboard, excluding planned maintenance outages notified 48 hours in advance. | subsystem-requirements | Subsystem Requirements |  | Analysis |  |  | 99.5% equates to approximately 44 hours unplanned downtime per year. Below this, analytical campaigns requiring multi-day continuous instrument operation (e.g., 72-hour LSC counts) would face unacceptable interruption rates. The dockyard HV ring provides dual feeds; 99.5% is achievable with automatic changeover. | subsystem, lab-utilities, session-335, idempotency:sub-epds-availability-335
SUB-REQ-090 | The Electrical Power Distribution System SHALL provide uninterruptible power supply to ICP-MS instruments, gamma spectrometers, LIMS servers, and radiation monitoring data loggers with less than 10ms transfer time and at least 15 minutes autonomy at full rated load. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | ICP-MS plasma extinction occurs within 2ms of power loss, requiring full instrument restart and 30-minute restabilisation. Gamma spectrometry counts spanning hours would lose accumulated data. 15 minutes autonomy covers generator start and load acceptance time with margin. | subsystem, lab-utilities, session-335, idempotency:sub-epds-ups-335
SUB-REQ-091 | The Deionised Water Treatment and Distribution System SHALL produce water meeting ASTM D1193 Type I specifications: resistivity greater than or equal to 18.0 megaohm-centimetres at 25 degrees C, total organic carbon less than 50 parts per billion, and bacterial count below 1 CFU per millilitre. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Trace radiochemical analysis by ICP-MS operates at sub-ppt detection limits. Ionic contaminants in reagent water create spectral interferences and elevated blanks that invalidate results. Type I water is the minimum standard for quantitative trace analysis per ISO 17025 accreditation requirements. | subsystem, lab-utilities, session-335, idempotency:sub-diwater-quality-335
SUB-REQ-092 | When the fire detection system activates a zone alarm, the Laboratory Gas Supply System SHALL automatically isolate all flammable gas supplies (hydrogen, acetylene) to the affected zone within 5 seconds via fail-safe solenoid valves. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Flammable gas accumulation in a radiological area during a fire creates a combined radiological-explosion hazard. 5-second isolation prevents gas accumulation above lower explosive limit in worst-case leak scenarios, derived from zone volume and maximum flow rate calculations. | subsystem, lab-utilities, session-335, idempotency:sub-gas-fire-isolation-335
SUB-REQ-093 | The Building Management System SHALL implement network segmentation isolating operational technology networks from enterprise IT networks via a unidirectional security gateway, with all remote access authenticated via multi-factor authentication and logged with 12-month retention. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | ONR expects nuclear-connected OT networks to be isolated from enterprise networks to prevent cyber-attack vectors reaching safety-related systems. Unidirectional gateway ensures monitoring data flows out but no commands can originate from the IT side. NIS Regulations require audit trail. | subsystem, lab-utilities, session-335, idempotency:sub-bms-cyber-335
SUB-REQ-094 | The Physical Security and Access Control System SHALL enforce three access tiers: Tier 1 (building entry, proximity card), Tier 2 (controlled area, card plus PIN), and Tier 3 (nuclear material accountability zones and high-activity stores, card plus biometric verification), with access rights reviewed quarterly. | subsystem-requirements | Subsystem Requirements |  | Inspection |  |  | NISR 2003 requires graded physical protection proportional to the consequence of malicious act. Three tiers align with the facility consequence categories: low (general labs), medium (active areas), and high (nuclear material stores). Quarterly review catches personnel changes and prevents credential drift. | subsystem, lab-utilities, session-335, idempotency:sub-security-tiers-335
SUB-REQ-095 | The Chilled Water and Heating System SHALL maintain laboratory ambient temperature within 20 plus or minus 2 degrees C in all analytical areas during occupied hours, with chilled water flow temperature controlled to 6 plus or minus 0.5 degrees C. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Analytical instrument calibration and measurement uncertainty budgets assume stable ambient temperature. ICP-MS and gamma spectrometer electronics are temperature-sensitive; ambient excursions beyond 2C from setpoint cause measurable detector gain drift and plasma instability, requiring recalibration. | subsystem, lab-utilities, session-335, idempotency:sub-chw-temp-335
SUB-REQ-096 | The Laboratory Gas Supply System SHALL provide automatic manifold changeover for nitrogen and compressed air supplies, switching from primary to reserve cylinder bank without interruption to laboratory supply pressure, and generating a low-reserve alarm at 20 percent remaining capacity. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | ICP-MS collision cell and glove box atmosphere require uninterrupted nitrogen flow. Manual changeover risks supply interruption during overnight or weekend analytical runs. 20 percent low-reserve threshold provides approximately 8 hours response time for cylinder replacement based on typical consumption rates. | subsystem, lab-utilities, session-335, idempotency:sub-gas-changeover-335
SUB-REQ-097 | The Deionised Water Treatment and Distribution System SHALL recirculate stored water continuously at a minimum flow velocity of 0.5 m/s within the distribution loop, with continuous online monitoring of resistivity and automatic diversion of off-specification water back to the treatment stage. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Stagnant ultra-pure water rapidly degrades through CO2 absorption and biofilm formation. 0.5 m/s minimum velocity prevents dead-leg colonisation. Automatic diversion prevents off-spec water reaching analytical points of use, which would invalidate quality control blanks and potentially contaminate samples. | subsystem, lab-utilities, session-335, idempotency:sub-diwater-recirc-335
SUB-REQ-098 | The Building Management System SHALL present utility alarms in priority order (critical, high, medium, low) on control room workstations within 2 seconds of detection, with critical alarms requiring operator acknowledgement and generating automatic notification to the facility duty manager. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | Nuclear site licence condition requires prompt operator awareness of facility conditions. 2-second alarm presentation ensures utility failures affecting containment systems (e.g., chiller failure causing extract fan motor overheating) are detected before secondary protection trips occur. Prioritisation prevents alarm flooding. | subsystem, lab-utilities, session-335, idempotency:sub-bms-alarms-335
SUB-REQ-099 | The Safety Interlock and Trip System SHALL be subject to proof testing at intervals not exceeding 6 months for all SIL 2 functions and 3 months for all SIL 3 functions, with each proof test exercising the complete trip chain from sensor input through logic solver to final element actuation, and test results recorded in the maintenance management system. | subsystem-requirements | Subsystem Requirements |  | Test |  |  | IEC 61511 requires proof testing at intervals consistent with the target SIL. The 2oo3 voting architecture specified in SUB-REQ-037 achieves SIL 3 only when combined with sufficiently frequent proof testing to detect dangerous undetected failures. Cross-domain analysis of Nuclear Reactor Protection Systems confirms that 3-month proof test intervals are standard practice for SIL 3 nuclear safety functions. Without defined proof test intervals, the claimed SIL cannot be demonstrated in the safety case. | subsystem, safety, proof-testing, validation, session-337
SUB-REQ-100 | The Safety Interlock and Trip System SHALL employ diverse sensing technologies for each safety function such that no single common cause failure mechanism can disable all redundant channels simultaneously, with diversity demonstrated through use of different physical measurement principles or different equipment manufacturers for at least one channel in each 2oo3 voting group. | subsystem-requirements | Subsystem Requirements |  | Analysis |  |  | Cross-domain analysis shows Nuclear Reactor Protection Systems with 0.91 trait similarity require diverse redundancy to defeat common cause failures. IEC 61511 clause 11.4 requires CCF defences proportional to SIL. In a radiochemistry laboratory, common cause failures from chemical attack on sensors or electromagnetic interference from nearby equipment could disable identical redundant channels simultaneously. Diversity in sensing principle provides inherent defence. | subsystem, safety, ccf, validation, session-337
SYS-REQ-001 | The Gamma Spectrometry Suite SHALL achieve a minimum detectable activity (MDA) of 0.5 Bq/L for Cs-137 and Co-60 in a 1-litre Marinelli beaker geometry with a 1-hour counting time, using HPGe detectors with relative efficiency of not less than 30%. | system-requirements | System Requirements |  | Test |  |  | Reactor coolant Cs-137 and Co-60 are primary fuel clad failure indicators. The 0.5 Bq/L MDA at 1-hour count enables detection at levels well below operational action limits (~50 Bq/L for Cs-137) while keeping turnaround within the 4-hour window required by STK-REQ-001. The 30% relative efficiency is the minimum for achieving this MDA in a 1-hour count with typical background. | system, gamma-spec, session-325
SYS-REQ-002 | The Active Ventilation and Containment System SHALL maintain a depression cascade with pressure differentials of at least 15 Pa between successive containment zones (corridor to C1, C1 to C2, C2 to C3, C3 to C4) under all normal operating conditions including door openings. | system-requirements | System Requirements |  | Test |  |  | The depression cascade is the primary engineered control preventing airborne contamination from spreading from higher-activity zones to lower-activity zones and occupied areas. The 15 Pa minimum per zone boundary is derived from HSE/ONR guidance for nuclear facility ventilation and ensures that transient events (door openings creating ~10 Pa perturbation) do not reverse flow direction. This directly supports ALARP for inhalation dose. | system, ventilation, safety, session-325
SYS-REQ-003 | The Active Effluent Treatment and Discharge System SHALL sample and analyse every liquid effluent batch for total alpha (<0.1 Bq/L detection limit), total beta (<1 Bq/L), and tritium (<10 Bq/L) before authorising discharge, with results recorded against the batch in the discharge database. | system-requirements | System Requirements |  | Test |  |  | EA/SEPA discharge authorisations require pre-discharge monitoring of every batch to prevent unauthorised releases. Detection limits are set at approximately 1/100th of typical batch concentration limits to ensure measurement uncertainty does not risk a non-compliant discharge being authorised. Batch-level recording enables retrospective audit of cumulative annual discharges against permit limits. | system, effluent, session-325
SYS-REQ-004 | The Alpha Spectrometry Laboratory SHALL achieve detection limits of 0.5 mBq per sample for Pu-239/240 and Am-241 with energy resolution of not more than 25 keV FWHM at 5.486 MeV (Am-241 peak), using PIPS detectors with a minimum counting time of 72 hours for environmental-level samples. | system-requirements | System Requirements |  | Test |  |  | Nuclear material accountancy for Pu requires quantification at environmental background levels to confirm there is no unaccounted material migration. The 0.5 mBq detection limit is consistent with IAEA requirements for environmental swipe analysis. 25 keV resolution is needed to resolve Pu-238 (5.499 MeV) from Am-241 (5.486 MeV) and Pu-239 (5.157 MeV) from Pu-240 (5.168 MeV) — failure to resolve these peaks invalidates isotopic ratios critical for material identification. | system, alpha-spec, session-325
SYS-REQ-005 | The Laboratory Information Management System SHALL maintain a tamper-evident electronic audit trail recording all sample data entries, modifications, approvals, and deletions with operator identity, timestamp, and reason for change, in compliance with UKAS ISO 17025 and ONR LC25 (operational records) requirements. | system-requirements | System Requirements |  | Inspection |  |  | LC25 requires the licensee to keep adequate records of operations. ISO 17025 clause 8.4 requires control of records including protection against unauthorised changes. The audit trail provides the evidential basis for demonstrating that analytical results are traceable, unaltered, and properly authorised — essential for both regulatory compliance and legal admissibility of results in nuclear safety cases. | system, lims, session-325
SYS-REQ-006 | The Radiation Protection and Health Physics Monitoring System SHALL detect airborne alpha contamination exceeding 1/10th of the derived air concentration (DAC) for Pu-239 within 60 seconds and initiate a local audible and visual alarm within 5 seconds of detection threshold exceedance. | system-requirements | System Requirements |  | Test |  |  | Airborne alpha contamination from Pu compounds is the highest-consequence hazard in the laboratory (committed effective dose per unit intake: 5×10⁻⁵ Sv/Bq for Pu-239 Type S). Detection at 1/10th DAC provides early warning before significant intake occurs. The 60-second detection time is constrained by the filter collection and counting statistics of continuous air monitors operating at 1-2 L/min flow rate. The 5-second alarm latency ensures personnel can respond (evacuate or don respiratory protection) before accumulating significant intake. | system, radpro, safety, session-325
SYS-REQ-007 | The ICP-MS and Elemental Analysis Suite SHALL achieve a detection limit of 0.05 Bq/L for Tc-99 in liquid effluent samples with a measurement uncertainty of not more than 15% (k=2) at 10 times the detection limit. | system-requirements | System Requirements |  | Test |  |  | Tc-99 is a long-lived beta emitter (t½ 2.13×10⁵ years) present in spent fuel and is specifically listed in EA discharge authorisations. LSC measurement of Tc-99 requires prior radiochemical separation, whereas ICP-MS provides faster direct determination at lower detection limits. The 0.05 Bq/L limit is 1/20th of typical batch discharge action levels and the 15% uncertainty at 10× LOD meets ISO 11929 requirements for regulatory reporting. | system, icp-ms, session-325
SYS-REQ-008 | The Active Ventilation and Containment System SHALL employ twin-bank HEPA filtration on all extract pathways from C3 and C4 zones, with each bank achieving a minimum decontamination factor of 1000 (99.9% removal efficiency) for 0.3 micron aerosols, DOP-tested in situ at installation and annually thereafter. | system-requirements | System Requirements |  | Test |  |  | HEPA filtration is the final barrier preventing airborne radioactive particulate discharge to atmosphere. Twin-bank configuration provides redundancy — if one bank fails or is being changed, the second maintains protection. The DF of 1000 per bank (combined DF of 10⁶) reduces stack discharge to negligible levels even during worst-case glove failure or spill events. In-situ DOP testing per BS EN 1822 confirms installed performance, as factory tests do not account for bypass leakage at frame seals. | system, ventilation, safety, session-325
SYS-REQ-009 | The Facility Safety and Emergency Response System SHALL enforce fissile material mass limits per workstation such that the total fissile inventory in any single laboratory room does not exceed 50% of the minimum critical mass for the most reactive credible configuration, with physical controls (container geometry, material form) providing at least two independent barriers to criticality. | system-requirements | System Requirements |  | Analysis |  |  | Nuclear site licence condition LC24 (fissile material) requires the licensee to ensure criticality cannot occur. The 50% margin provides conservative protection against accumulation errors and credible upset conditions. Two independent barriers (e.g. mass limit AND geometry control) ensure that no single failure leads to a critical configuration. This is particularly important during fuel element dissolution where fissile material is in solution form with higher reactivity than solid metal. | system, safety, criticality, session-325
SYS-REQ-010 | The Liquid Scintillation Counting Facility SHALL achieve a detection limit of 1 Bq/L for tritium in reactor coolant water samples with a counting time of not more than 120 minutes, using ultra-low-background counters with figure of merit (E²/B) exceeding 400. | system-requirements | System Requirements |  | Test |  |  | Tritium is produced by neutron activation of boron in PWR coolant and Li in primary circuit materials. It is a key indicator of coolant chemistry status and fuel performance. The 1 Bq/L detection limit is well below the typical coolant concentration (~10⁴-10⁶ Bq/L) but is needed for environmental monitoring of non-active drainage and groundwater. The 120-minute counting time keeps turnaround within the 4-hour operational requirement. FOM > 400 requires Quantulus-class counters with active guard counting. | system, lsc, session-325
SYS-REQ-011 | The Laboratory Information Management System SHALL implement role-based access control with multi-factor authentication for all users, encrypt all data at rest using AES-256 and in transit using TLS 1.2 or later, and maintain network segmentation isolating the LIMS server from general-purpose office networks and from safety-related operational technology networks. | system-requirements | System Requirements |  | Test |  |  | LIMS stores safeguards-relevant nuclear material accountancy data whose integrity directly affects material balance reporting. ONR CNS guidance requires defence-in-depth for computer-based systems on nuclear sites. Network segmentation prevents lateral movement from compromised office networks to safety or safeguards systems. MFA prevents credential theft enabling unauthorised data modification. AES-256 encryption at rest protects against physical media theft from the site. | system, cybersecurity, validation, session-337
SYS-REQ-012 | The Radiochemistry Laboratory SHALL maintain a decommissioning database recording all materials of construction in active zones, all instances of contamination events requiring remediation, and cumulative operational histories for each containment zone, sufficient to support future radiological characterisation during decommissioning planning. | system-requirements | System Requirements |  | Inspection |  |  | LC35 requires decommissioning arrangements to be maintained throughout the facility lifetime. A radiochemistry laboratory accumulates activation and contamination records over decades of operation. Without systematic recording from day one, decommissioning characterisation requires expensive physical sampling of every surface. The database enables waste categorisation estimates and dose predictions essential for decommissioning safety case development. | system, decommissioning, validation, session-337
VER-REQ-001 | Verify IFC-REQ-001: Inspect extract ductwork material certificates (316L stainless steel), pressure test ductwork assembly at minus 3 kPa for 30 minutes with less than 0.5 percent pressure decay, and verify fan-downstream-of-filters arrangement by ductwork isometric review. Pass criteria: material certificates confirm 316L, pressure test holds, fan is downstream of both HEPA banks. | verification-requirements | Verification Requirements |  | Inspection |  |  | Ductwork integrity and material are most efficiently verified by inspection and pressure test during construction, before radioactive commissioning precludes access. | verification, ventilation, session-326
VER-REQ-002 | Verify IFC-REQ-002: Inject step setpoint changes from cascade controller to fan VSD and measure response. Pass criteria: fan speed tracks setpoint within 2 percent, status feedback updates within 500 ms, fault condition correctly reported within 1 second of simulated fault injection. | verification-requirements | Verification Requirements |  | Test |  |  | Functional integration test during commissioning validates the control loop timing that underpins the 2-second cascade response requirement. | verification, ventilation, session-326
VER-REQ-003 | Verify IFC-REQ-003: Commission extract manifold with all fume cupboards at maximum sash opening and all glove boxes at operating pressure. Measure total manifold flow using traverse pitot measurement at manifold entry. Pass criteria: total flow not less than 12000 m3/h, manifold static pressure not less negative than minus 500 Pa. | verification-requirements | Verification Requirements |  | Test |  |  | Full-load commissioning test validates manifold sizing under worst-case simultaneous demand from all containment devices. | verification, ventilation, session-326
VER-REQ-004 | Verify IFC-REQ-004: Inspect as-built ductwork drawings and measure straight-run distances upstream and downstream of sampling point. Pass criteria: not less than 8 diameters downstream of last disturbance, not less than 2 diameters upstream of stack exit, isokinetic nozzle aligned with duct axis within 5 degrees. | verification-requirements | Verification Requirements |  | Inspection |  |  | Physical measurement of straight-run distances confirms ISO 2889 compliance. This is a construction verification that cannot be repeated after active commissioning. | verification, ventilation, session-326
VER-REQ-005 | Verify IFC-REQ-006: With cascade PLC in simulated failure state, activate fire damper closure and smoke extract mode via hardwired safety relays. Pass criteria: fire dampers close within 10 seconds, smoke extract fan starts within 15 seconds, safety relay circuits maintain function with PLC powered off. Perform SIL 2 proof test per IEC 61511. | verification-requirements | Verification Requirements |  | Test |  |  | Independence of emergency override from PLC must be demonstrated by testing with PLC in failed state. This validates the common-cause failure defence required for SIL 2. | verification, ventilation, safety, session-326
VER-REQ-006 | Verify end-to-end ventilation containment: Release tracer aerosol (DOP at 0.3 micrometre) inside a C4 glove box and measure tracer concentration at (a) C3 zone boundary, (b) C2 zone boundary, (c) stack discharge point, and (d) external building envelope. Pass criteria: C3 concentration less than 0.01 percent of source, C2 less than 0.0001 percent, stack less than 0.00001 percent (two HEPA banks), external envelope below detection limit. Depression cascade maintained throughout test. | verification-requirements | Verification Requirements |  | Test |  |  | End-to-end tracer test validates the complete containment chain from primary containment device through extract ductwork, both HEPA banks, and stack discharge. This is the definitive active commissioning test confirming that the ventilation system performs its safety function as an integrated whole, not just at component level. | verification, ventilation, integration, session-326
VER-REQ-007 | Verify IFC-REQ-007: Connect each CAM station to HP Central via RS-485 Modbus RTU bus. Confirm polling interval is 5 seconds or less by monitoring bus traffic with protocol analyser. Inject simulated DAC-fraction readings at 0.5, 1.0, and 3.0 DAC and verify correct display on HP Central mimic within 5 seconds. Confirm all data fields (DAC-fraction, alarm state, instrument status, filter accumulation) are transmitted and recorded. Pass criteria: all data fields present, polling interval within spec, end-to-end latency under 5 seconds. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test verifying the primary data path from distributed CAMs to central display. Tests both routine data and alarm state propagation. | verification, radiation-protection, cam-network, hp-central, session-327
VER-REQ-008 | Verify IFC-REQ-008: For each area gamma monitor, verify 4-20mA analogue output is proportional to log dose rate by injecting known Cs-137 fields at 1, 10, and 100 microSv/h and measuring analogue signal. Simultaneously verify Modbus TCP/IP digital path delivers dose rate, H*(10), cumulative dose, and health status at 10-second intervals. Disable digital path and confirm analogue alarm relay still activates at threshold. Pass criteria: analogue linearity within 5 percent, digital data complete at 10-second intervals, alarm functions on analogue path alone. | verification-requirements | Verification Requirements |  | Test |  |  | Diversity test confirming both data paths function independently. The analogue path independence test is critical because it is the SIL 1 safety function path. | verification, radiation-protection, gamma-array, hp-central, session-327
VER-REQ-009 | Verify IFC-REQ-009: Perform contamination measurement on each HFC monitor and portal monitor. Capture network traffic and confirm measurement results (pass/fail, alpha Bq/cm2, beta-gamma Bq/cm2, timestamp, monitor ID) arrive at HP Central within 2 seconds of measurement completion. Repeat for alarm condition. Pass criteria: all data fields present, latency under 2 seconds for 95th percentile across 20 measurements per monitor. | verification-requirements | Verification Requirements |  | Test |  |  | Verifies real-time contamination data path including latency under operational conditions. Statistical approach (95th percentile) accounts for network variability. | verification, radiation-protection, contamination, hp-central, session-327
VER-REQ-010 | Verify IFC-REQ-010: Issue EPDs to 5 test personnel, enter controlled area, and confirm OPC-UA data (dose rate, integrated dose, constraint percentage, personnel ID) appears on HP Central at 60-second intervals. Trigger EPD dose alarm and verify immediate event notification arrives at HP Central within 5 seconds. Pass criteria: routine data at 60-second intervals with no gaps over 1-hour test, alarm notification within 5 seconds. | verification-requirements | Verification Requirements |  | Test |  |  | End-to-end test of dosimetry-to-central interface under realistic conditions with multiple simultaneous EPDs. Alarm notification timing is critical for RPS situational awareness. | verification, radiation-protection, dosimetry, hp-central, session-327
VER-REQ-011 | Verify IFC-REQ-012: Inject simulated evacuation-level radiation alarm at each CAM and area gamma monitor in turn. Verify hardwired relay contact closure at HP Central within 1 second, and confirm facility evacuation alarm activation within 1 second of relay closure (2 seconds total from radiation alarm). Test with digital network disabled to confirm relay path is independent. Pass criteria: relay closure within 1 second for all monitors, evacuation alarm within 2 seconds total, function maintained with network disabled. | verification-requirements | Verification Requirements |  | Test |  |  | This is the most safety-critical interface in the RP subsystem. The test must demonstrate deterministic hardwired operation independent of digital infrastructure. Network-disabled test proves the safety function does not rely on software. | verification, radiation-protection, hp-central, facility-safety, session-327
VER-REQ-012 | Verify end-to-end radiation protection chain: Release a known-activity Tc-99m aerosol source (1 MBq) inside a fume cupboard in the Radiochemical Separations Laboratory. Verify the sequence: CAM detects airborne contamination and alarms at investigation level within 30 seconds, HP Central displays the alarm and identifies the zone within 5 seconds, area gamma monitors in adjacent zones show no increase (confirming containment), and personnel EPDs in the zone record no significant dose increase (confirming ventilation effectiveness). Pass criteria: complete detection-to-display chain within 35 seconds, no contamination spread beyond source zone, full data chain recorded in historian. | verification-requirements | Verification Requirements |  | Demonstration |  |  | End-to-end system integration test exercising the complete detection-alarm-display chain under realistic conditions. Uses Tc-99m (6-hour half-life) for safe aerosol release testing. Tests both the RP subsystem's detection function and its integration with the ventilation containment system. This is the definitive acceptance test for the radiation protection safety function. | verification, radiation-protection, system-integration, session-327
VER-REQ-013 | Verify IFC-REQ-013: Using inactive simulant at maximum flow rate, confirm gravity and pump transfer through 50mm bore pipework to each delay tank. Verify flow totalisation accuracy within ±2% against calibrated reference. Confirm isolation valve operation at each tank inlet. Pass criteria: flow rate sustains 20 L/min, totaliser agrees with reference within ±2%, each valve closes within 10 seconds. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test at the drain-to-tank boundary using inactive simulant to verify hydraulic performance before active commissioning. | verification, aetds, session-328
VER-REQ-014 | Verify IFC-REQ-014: With tanks filled with traced simulant, execute 5-minute purge cycle then draw triplicate samples from each tank. Analyse samples against known tank concentration. Pass criteria: all samples within ±10% of true concentration, no cross-contamination between tank sample lines (absent tracer in adjacent tank samples). | verification-requirements | Verification Requirements |  | Test |  |  | Validates sample line purge effectiveness and cross-contamination isolation using different tracers per tank. | verification, aetds, session-328
VER-REQ-015 | Verify IFC-REQ-016: Inject calibration sources at 80%, 100%, and 120% of discharge limits into the in-line alpha and beta-gamma monitors during simulated discharge. Verify alarm generation and fail-safe valve closure within 5 seconds at each threshold. Simulate loss of monitoring signal and loss of power — confirm valve closes within 5 seconds in each case. Pass criteria: all alarm setpoints trigger correctly, valve closure time <5 seconds in all failure modes, proportional sampler collects representative composite. | verification-requirements | Verification Requirements |  | Test |  |  | SIL 2 proof test validating the complete safety function chain from detection through logic to final element. Tests all three trip conditions (high activity, signal loss, power loss). | verification, aetds, safety, session-328
VER-REQ-016 | Verify IFC-REQ-015: Transfer a spiked simulant batch through the Chemical Treatment Plant at minimum and maximum flow rates. Confirm treated effluent is returned to receiving tank and post-treatment sample shows decontamination factors meet SUB-REQ-025 targets. Pass criteria: transfer completes without leak at both flow rates, DF for alpha surrogate exceeds 100, DF for Cs/Sr surrogates exceeds 10. | verification-requirements | Verification Requirements |  | Test |  |  | End-to-end treatment loop verification using surrogate radionuclides during inactive commissioning. | verification, aetds, session-328
VER-REQ-017 | Verify IFC-REQ-017: Trigger a batch screening measurement and confirm structured data record appears in LIMS within 30 seconds. Verify OPC UA authentication handshake succeeds and data integrity checksum matches. Corrupt a test message in transit and confirm LIMS rejects it. Pass criteria: latency <30 seconds for 100 consecutive measurements, all corrupted messages rejected. | verification-requirements | Verification Requirements |  | Test |  |  | Validates both the functional data path and the security/integrity mechanisms of the OPC UA interface. | verification, aetds, session-328
VER-REQ-018 | Verify end-to-end AETDS batch cycle: Introduce spiked simulant into the Active Drain Collection Network, confirm collection in delay tank, execute hold period, draw and analyse sample via Batch Sampling Station, transfer through Chemical Treatment Plant if required, obtain discharge authorisation via LIMS, and discharge through the Discharge Monitoring and Control System with continuous monitoring. Pass criteria: full batch cycle completes within 48 hours, all intermediate data records present in LIMS, discharge activity below limits, proportional sample archived. | verification-requirements | Verification Requirements |  | Demonstration |  |  | System-level integration test exercising the complete effluent management chain from waste generation to authorised discharge. 48-hour target includes 24-hour hold period plus processing time. | verification, aetds, integration, session-328
VER-REQ-019 | Verify IFC-REQ-018: Inspect a filled 200-litre sludge drum from the Chemical Treatment Plant. Confirm dose rate at contact is below 2 mSv/h using calibrated survey meter. Verify waste characterisation record includes activity inventory, chemical composition, and package weight. Confirm drum is compatible with RWMF receipt criteria. Pass criteria: contact dose rate <2 mSv/h, characterisation record complete, drum passes RWMF acceptance checklist. | verification-requirements | Verification Requirements |  | Inspection |  |  | Verification of the waste transfer interface by inspection of the physical package and accompanying documentation. | verification, aetds, session-328
VER-REQ-020 | Verify IFC-REQ-019: Simulate criticality trip condition at the Criticality Warning System test input. Measure signal propagation latency from CWS relay output to SIT trip receipt using calibrated timing equipment. Verify latency is less than 10 ms across all three voting channels. Verify galvanic isolation by applying 2.5 kV test voltage between circuits for 60 seconds with no breakdown. Pass criteria: all channels below 10 ms latency, zero isolation failures. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test at system boundaries to verify hardwired interface timing and isolation. | verification, safety, session-329
VER-REQ-021 | Verify IFC-REQ-020: Activate fire detection in a test zone. Measure time from fire confirmation to damper closure command receipt at the Standby Ventilation system. Verify command transmission within 2 seconds. Verify cable integrity by sustained fire test to BS 8519 at 842 degrees C for 120 minutes with circuit continuity maintained. Pass criteria: command latency below 2 seconds, cable circuit integrity maintained for 120 minutes. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test verifying fire-to-ventilation interface timing and fire survivability of cabling. | verification, fire, session-329
VER-REQ-022 | Verify IFC-REQ-021: Remove primary UPS feed to the Criticality Warning System. Verify automatic changeover to secondary feed occurs within 1 ms using oscilloscope monitoring of the 24 VDC supply rail. Verify full criticality detection load is sustained on single feed. Repeat test removing secondary feed. Pass criteria: changeover time below 1 ms on both feeds, no detector dropout or counting interruption during transfer. | verification-requirements | Verification Requirements |  | Test |  |  | Power interface test ensuring zero detection gap during UPS feed changeover. | verification, power, session-329
VER-REQ-023 | Verify IFC-REQ-022: Initiate a simulated fire trip from the Safety Interlock system. Verify isolation commands reach the Depression Cascade Controller. Measure containment depression in non-affected zones during partial isolation and verify minimum 10 Pa is maintained. Pass criteria: trip command received, depression in non-affected zones remains above 10 Pa throughout the isolation transient. | verification-requirements | Verification Requirements |  | Test |  |  | System integration test verifying partial isolation maintains containment where no fire exists. | verification, ventilation, session-329
VER-REQ-024 | Verify IFC-REQ-023: Inspect all floor drain connections between spill containment areas and the active drain network. Verify no cross-connections to inactive drainage by dye tracing from each drain point. Measure water seal depth at each floor trap. Pass criteria: zero cross-connections found, all trap seals at or above 50 mm depth. | verification-requirements | Verification Requirements |  | Inspection |  |  | Physical inspection and dye trace test to verify drainage segregation integrity. | verification, containment, session-329
VER-REQ-025 | Verify end-to-end criticality response: inject simulated neutron burst at CWS detector input exceeding 20 mrad threshold. Verify the complete chain: CWS detection within 1 ms, trip signal to SIT within 10 ms, SIT trip initiation within 500 ms, criticality alarm activation via ECAS, ventilation isolation command to standby ventilation system. Total end-to-end time from neutron burst to completed facility response SHALL be less than 2 seconds. Pass criteria: all subsystems activate in sequence within timing budgets, alarm audible at 75 dBA, no subsystem fails to respond. | verification-requirements | Verification Requirements |  | Test |  |  | System-level integration test exercising the full criticality response chain from sensor to actuator. Tests the interfaces between CWS, SIT, ECAS, and ventilation systems as an integrated safety function. | verification, integration, criticality, session-329
VER-REQ-026 | Verify IFC-REQ-024: Test data transfer from characterization to packaging by performing gamma assay and surface survey on a reference waste item, then confirming the categorisation certificate appears in the packaging system within 5 minutes. Pass criteria: certificate contains correct waste category, activity inventory matching reference values within measurement uncertainty, and approved packaging route. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test to verify the characterization-to-packaging data flow meets the 5-minute latency and data completeness requirements at the system boundary. | verification, waste-management, characterization, session-330
VER-REQ-027 | Verify IFC-REQ-025: Test drum transfer from packaging to store by sealing a test drum, verifying barcode scan retrieves the correct waste package data record, confirming surface contamination below transport limits, and tracking placement in the designated store position. Pass criteria: barcode correctly linked, contamination verified, placement within 4 hours, and store position matches criticality spacing plan. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test at the packaging-to-store boundary ensuring positive identification, contamination verification, and criticality-safe placement are all functioning as a chain. | verification, waste-management, packaging, session-330
VER-REQ-028 | Verify IFC-REQ-026: Test cemented wasteform transfer by producing a test batch, holding for the 7-day curing period, confirming early-age strength test results are recorded, and verifying the package is assigned a criticality-compliant store position. Pass criteria: 7-day hold enforced by system interlock, strength test recorded against package ID, store position respects keff spacing. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test verifying the curing-period hold, strength verification, and criticality-safe placement operate as a controlled sequence. Tests the most safety-significant interface in the waste management chain. | verification, waste-management, liquid-conditioning, session-330
VER-REQ-029 | Verify IFC-REQ-027: Test distillate return by processing a spiked liquid waste batch through evaporation, sampling the distillate at the hold-and-release point, and confirming activity concentrations are below 100 Bq/L alpha and 1000 Bq/L beta/gamma before drain release. Pass criteria: analytical results within limits, flow rate not exceeding 0.5 L/min, hold-and-release interlock prevents release if limits exceeded. | verification-requirements | Verification Requirements |  | Test |  |  | Verifies the critical environmental protection boundary between waste conditioning and the active drain system. Tests both the decontamination performance and the hold-and-release safety function. | verification, waste-management, liquid-conditioning, effluent, session-330
VER-REQ-030 | Verify IFC-REQ-028: Test bidirectional data exchange by creating a waste characterisation record in LIMS and confirming it appears in the Waste Records system within 60 seconds, then creating a waste package record in the Waste Records system and confirming chain-of-custody data is accessible from LIMS. Pass criteria: both synchronisation directions complete within 60 seconds, mutual TLS handshake verified, data integrity confirmed by checksum. | verification-requirements | Verification Requirements |  | Test |  |  | Verifies the information system integration that underpins regulatory record-keeping. Both directions must work because LIMS is the analytical authority and the Waste Records system is the waste package authority — each needs access to the other's data. | verification, waste-management, records, lims, session-330
VER-REQ-031 | Verify IFC-REQ-029: Test concentrate transfer by initiating a batch transfer from the Chemical Treatment Plant to the Liquid Waste Conditioning System with the conditioning system in both ready and not-ready states. Pass criteria: transfer proceeds when conditioning system is ready, physical interlock prevents transfer when not ready, analytical certificate accompanies the batch, and batch volume does not exceed 50 L. | verification-requirements | Verification Requirements |  | Test |  |  | Tests the safety interlock that prevents uncontrolled transfer of radioactive concentrates to an unprepared system. The interlock is a credited safety function in the liquid waste safety case. | verification, waste-management, aetds, liquid-conditioning, session-330
VER-REQ-032 | Verify end-to-end waste lifecycle: Process a solid waste item and a liquid waste batch through the complete RWMF chain from characterization/conditioning through packaging to storage placement, confirming that waste package data records are complete, nuclear material accountancy is reconciled with the site system within 24 hours, and all packages are placed in criticality-compliant store positions with environmental monitoring active. Pass criteria: all intermediate records created automatically, NMA reconciliation confirmed, store environmental parameters within limits, no orphaned packages without complete data records. | verification-requirements | Verification Requirements |  | Demonstration |  |  | End-to-end integration test exercising both solid and liquid waste streams through the entire RWMF. This test verifies that the individual interface tests (IFC-REQ-024 through IFC-REQ-029) function as a coherent system and that no data is lost across the waste lifecycle chain. | verification, waste-management, integration, session-330
VER-REQ-033 | Verify IFC-REQ-030: Commissioning test of pass-through hatch interlock by simultaneously commanding both doors open and confirming interlock prevents second door from opening. Verify container clearance with a 500 mm diameter by 800 mm height test article. Pass criteria: interlock holds under all attempted sequences; test article passes through without obstruction. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test to verify physical interface compliance at the C2/C3 containment boundary. | verification, sample-receipt, session-331
VER-REQ-034 | Verify IFC-REQ-031: Demonstrate end-to-end data transfer from screening station measurement completion to Registration System display within 30 seconds using a test source. Verify dual physical routing by processing one contact-handled and one remote-handled sample through their respective routes. Pass criteria: data arrives within 30 seconds; each sample reaches correct preparation area. | verification-requirements | Verification Requirements |  | Demonstration |  |  | Integration test verifying both the electronic data interface and the physical sample routing paths function as designed. | verification, sample-receipt, session-331
VER-REQ-035 | Verify IFC-REQ-032: Demonstrate message queuing resilience by registering 10 samples with LIMS offline, then restoring LIMS and confirming all 10 records synchronise within 15 minutes of recovery with no data loss. Pass criteria: all 10 records appear in LIMS with correct fields; resync completes within 15 minutes. | verification-requirements | Verification Requirements |  | Demonstration |  |  | Integration test proving the Registration System can operate independently during LIMS outage and resynchronise without data loss — a key architectural decision. | verification, sample-receipt, session-331
VER-REQ-036 | Verify end-to-end Sample Receipt workflow: process a simulated sample from physical arrival at the Receiving Bay through dose rate screening, contamination check, registration, barcode labelling, preparation (acid digestion with tracer spiking), and storage. Pass criteria: sample arrives at storage with complete chain-of-custody record, all screening results recorded in LIMS, preparation records traceable to original registration, and total workflow time under 4 hours for a contact-handled sample. | verification-requirements | Verification Requirements |  | Demonstration |  |  | System-level integration test exercising the complete sample receipt chain from physical arrival to prepared aliquot. The 4-hour target is the operational throughput requirement for routine dockyard samples. | verification, sample-receipt, session-331
VER-REQ-037 | Verify IFC-REQ-033: Analyse five transferred aliquots by ICP-OES to confirm HNO3 molarity is within 2-8 M range. Verify volume does not exceed 100 mL. Verify barcode on PTFE vial matches sample record. Pass criteria: all five aliquots within acid range; volumes at or below 100 mL; barcode-to-record match 100 percent. | verification-requirements | Verification Requirements |  | Test |  |  | Acid matrix compatibility is critical for downstream chromatography performance — feed outside 2-8 M HNO3 causes actinide breakthrough on TEVA resin. | verification, radchem-sep, session-331
VER-REQ-038 | Verify IFC-REQ-034: Weigh ten electrodeposited discs and confirm deposited mass is below 100 micrograms. Measure alpha spectrum on each disc and confirm FWHM resolution is 40 keV or better for the Pu-239/240 peak. Pass criteria: all ten discs below mass limit; all spectra meet resolution threshold. | verification-requirements | Verification Requirements |  | Test |  |  | Source quality directly determines measurement performance — this test verifies the hand-off point between separations and counting. | verification, radchem-sep, session-331
VER-REQ-039 | Verify IFC-REQ-035: Inspect waste segregation during a simulated multi-sample processing campaign. Verify each waste stream is correctly routed to its designated container with barcode linkage to originating work order. Verify organic solvent containers have flash point and activity labels. Pass criteria: zero cross-contamination between streams; 100 percent barcode traceability. | verification-requirements | Verification Requirements |  | Inspection |  |  | Waste segregation failure has regulatory and safety consequences — inspection during operational simulation is the appropriate verification method for procedural compliance. | verification, radchem-sep, session-331
VER-REQ-040 | Verify IFC-REQ-036: Inject a precision pulser signal at the preamplifier test input and measure pulse shape at the MCA input with a digital oscilloscope. Verify rise time is less than 100 ns, amplitude is within 0 to 10 V range, and noise floor is less than 2 mV RMS. Pass criterion: all 8 channels meet specification simultaneously with 15 m cable runs installed. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test confirms cable plant and connector quality after installation. Noise floor measurement validates electromagnetic compatibility in the nuclear facility environment. | verification, gamma-spectrometry, session-333
VER-REQ-041 | Verify IFC-REQ-037: Initiate simultaneous spectrum transfers from all 8 MCA channels to the analysis workstation. Measure transfer completion time for each 16384-channel spectrum. Pass criterion: all transfers complete within 2 seconds, no packet loss on dedicated VLAN, and spectral data integrity verified by checksum comparison. | verification-requirements | Verification Requirements |  | Test |  |  | Concurrent transfer from all channels represents worst-case network load during high-throughput campaigns. | verification, gamma-spectrometry, session-333
VER-REQ-042 | Verify IFC-REQ-038: Process a multi-nuclide reference sample spectrum through the full analysis chain and verify that the workstation automatically posts all required data fields to LIMS upon analyst approval. Pass criterion: nuclide identity, activity, uncertainty, MDA, and QC status fields all populate correctly in LIMS, with timestamp correlation within 5 seconds of approval action. | verification-requirements | Verification Requirements |  | Demonstration |  |  | End-to-end demonstration validates the automated result transfer chain that eliminates transcription errors. | verification, gamma-spectrometry, session-333
VER-REQ-043 | Verify end-to-end gamma spectrometry chain: place a NIST-traceable mixed-nuclide reference source (containing Cs-137, Co-60, Am-241, Eu-152 at certified activities) in a standard counting geometry, acquire spectrum for 3600 seconds, and process through automated analysis to LIMS result. Pass criterion: all reported activities within 10 percent of certified values at 95 percent confidence, MDA for Cs-137 is less than or equal to 0.5 Bq/L equivalent, total turnaround from sample placement to LIMS result less than 90 minutes. | verification-requirements | Verification Requirements |  | Test |  |  | System-level integration test validates that the combined performance of detector, shielding, MCA, and analysis software meets SYS-REQ-001. Exercises the full chain from photon detection to certified result. | verification, gamma-spectrometry, session-333
VER-REQ-044 | Verify IFC-REQ-039: Inspect all 8 vacuum chamber sample trays for correct spring-clip engagement with 25 mm planchets. Verify engraved identification is legible after 10 electrodeposition cycles. Pass criterion: all planchets seat repeatably within 0.5 mm of reference position, identification legible under standard laboratory lighting. | verification-requirements | Verification Requirements |  | Inspection |  |  | Positional repeatability determines counting geometry reproducibility. Identification legibility ensures chain-of-custody integrity throughout sample lifetime. | verification, alpha-spectrometry, session-333
VER-REQ-045 | Verify IFC-REQ-040: For each of the 8 channels, test vacuum interlock by venting one chamber while remaining chambers are under vacuum. Verify that only the vented chamber MCA channel inhibits acquisition while all other channels continue counting uninterrupted. Pass criterion: acquisition inhibits within 5 seconds of vacuum status changing to bad, resumes within 10 seconds of vacuum restoration. | verification-requirements | Verification Requirements |  | Test |  |  | Independent vacuum interlock per channel ensures a single chamber service event does not halt all alpha counting operations. | verification, alpha-spectrometry, session-333
VER-REQ-046 | Verify end-to-end alpha spectrometry: process a NIST-traceable mixed-actinide reference solution (Pu-239, Am-241, Cm-244) through electrodeposition, count for 86400 seconds, and analyse with MCA software. Pass criterion: reported activities within 15 percent of certified values at 95 percent confidence, Pu-239/240 MDA of 0.5 mBq or less, tracer recovery between 70 and 110 percent, and Pu-238 peak resolved from Am-241 with valley-to-peak ratio less than 0.3. | verification-requirements | Verification Requirements |  | Test |  |  | System-level integration test validates the complete alpha analysis chain against SYS-REQ-004. Valley-to-peak ratio confirms adequate source quality and detector resolution for the most demanding peak separation in the analysis. | verification, alpha-spectrometry, session-333
VER-REQ-047 | Verify IFC-REQ-041: Test by processing 10 simulated H-3 distillate and 10 C-14 fraction transfers through the full handover protocol. Verify vial integrity (no leakage after 24h), label readability, documentation completeness, and that receiving QC checks identify any out-of-specification fractions. Pass criteria: 100% documentation compliance, zero vial failures, all barcode scans successful. | verification-requirements | Verification Requirements |  | Test |  |  | Interface testing at the separations-LSC boundary validates the physical handover process that ensures sample integrity and traceability. | verification, liquid-scintillation, session-334
VER-REQ-048 | Verify IFC-REQ-042: Test by loading 50 prepared vials in standard counting cassettes into the sample changer. Verify all barcodes read correctly, vial geometry acceptance rate is 100%, and counting protocol auto-assignment matches LIMS work orders. Pass criteria: zero barcode read failures, zero vial rejection, correct protocol assignment for all 50 vials. | verification-requirements | Verification Requirements |  | Test |  |  | Sample changer interface testing ensures mechanical compatibility and barcode system reliability under realistic batch sizes. | verification, liquid-scintillation, session-334
VER-REQ-049 | Verify IFC-REQ-043: Test by running a full counting batch (20 vials including standards, blanks, and spiked samples) and verifying complete spectral data transfer to the analysis workstation. Verify all specified data fields are present, spectrum channel count matches instrument specification, tSIE values are within expected range for each quench level, and transfer completes within 60 seconds of counting cycle end. | verification-requirements | Verification Requirements |  | Test |  |  | Data transfer integrity testing ensures the analysis workstation receives complete, uncorrupted spectral data for accurate activity calculation and QA review. | verification, liquid-scintillation, session-334
VER-REQ-050 | Verify IFC-REQ-044: Test by processing a validated batch through QA approval and monitoring transfer to LIMS. Verify all required data fields arrive in LIMS within 30 minutes, activity values match workstation values to 6 significant figures, uncertainty values transfer correctly, and authentication/integrity checks pass. Inject deliberate data corruption to verify integrity check rejection. | verification-requirements | Verification Requirements |  | Test |  |  | LIMS interface testing validates the regulated data path from analysis to system of record, including positive and negative testing of integrity controls. | verification, liquid-scintillation, session-334
VER-REQ-051 | Verify end-to-end LSC chain: Test by processing 5 blind H-3 spiked samples and 5 blind C-14 spiked samples through the complete path from separations handover, through sample preparation, counting, data analysis, QA approval, to LIMS result entry. Pass criteria: all 10 results within 10% of the known spike value, all uncertainty budgets correctly calculated, and total turnaround time from vial receipt to LIMS entry not exceeding 48 hours. | verification-requirements | Verification Requirements |  | Test |  |  | End-to-end integration test exercises the full LSC measurement chain under realistic conditions. The 10% accuracy criterion against known spikes validates measurement traceability. 48-hour turnaround validates operational throughput during dockyard campaigns. | verification, liquid-scintillation, session-334
VER-REQ-052 | Verify IFC-REQ-045: Test by processing 20 simulated separated fractions through the autosampler. Verify matrix compatibility (no signal suppression >10%), tube geometry acceptance, label readability, and that autosampler processes all tubes without intervention. Include 3 matrix-mismatched tubes (>5% HNO3) to verify detection of out-of-specification fractions. | verification-requirements | Verification Requirements |  | Test |  |  | Interface testing validates physical and chemical compatibility between separation chemistry output and ICP-MS sample introduction. | verification, icp-ms, session-334
VER-REQ-053 | Verify IFC-REQ-046: Test by running a full calibration and sample batch (30 samples) and verifying complete data transfer. Confirm all mass channels present, CPS values match instrument display, internal standard ratios calculated correctly, and automatic backup completes within 5 minutes of batch end. Simulate network failure to verify data buffering and retry. | verification-requirements | Verification Requirements |  | Test |  |  | Data transfer testing validates the measurement data pipeline integrity and resilience for the complete analytical batch workflow. | verification, icp-ms, session-334
VER-REQ-054 | Verify IFC-REQ-047: Test bidirectional communication with each instrument type by sending 50 work orders and receiving 50 result sets per instrument. Verify work order parameters arrive correctly, result values match source to 6 significant figures, checksums validate on all transfers, and message queuing functions correctly during a simulated 30-minute LIMS outage. Pass criteria: zero data loss, zero corruption, queue drains within 10 minutes of LIMS recovery. | verification-requirements | Verification Requirements |  | Test |  |  | Comprehensive interface testing validates the LIMS integration with all instrument types under normal and degraded conditions. | verification, lims, session-334
VER-REQ-055 | Verify IFC-REQ-048: Integration test of BMS-to-EPDS Modbus TCP interface. Simulate transformer overtemperature condition and verify BMS receives alarm within 1 second and issues load shedding command. Pass criteria: load shedding confirmed on non-essential distribution board within 5 seconds of overtemperature setpoint breach. | verification-requirements | Verification Requirements |  | Test |  |  | Integration test to verify interface compliance and protection response timing at system boundaries. | verification, lab-utilities, session-335
VER-REQ-056 | Verify IFC-REQ-049: Integration test of BMS-to-Chilled Water BACnet/IP interface. Simulate primary chiller fault and verify automatic duty/standby changeover completes with chilled water flow temperature maintained within 6 plus or minus 1C during transition. Pass criteria: changeover complete within 60 seconds with no laboratory temperature excursion. | verification-requirements | Verification Requirements |  | Test |  |  | Duty/standby changeover is the primary resilience mechanism for maintaining laboratory temperature stability during chiller plant faults. | verification, lab-utilities, session-335
VER-REQ-057 | Verify IFC-REQ-050: Integration test of gas isolation on fire alarm. Activate zone fire alarm and measure time from alarm signal to confirmed valve closure on flammable gas supply. Pass criteria: all zone isolation valves closed within 5 seconds and closure status reported to fire panel within 3 seconds of closure. | verification-requirements | Verification Requirements |  | Test |  |  | Fire safety case critical test. Gas isolation timing is a fire safety case claim that must be demonstrated before facility commissioning. | verification, lab-utilities, session-335
VER-REQ-058 | Verify IFC-REQ-054: Integration test of normal-to-emergency power changeover. Simulate normal supply failure and verify generator start, load acceptance, and motorised changeover. Pass criteria: essential loads transferred within 15 seconds of generator reaching rated output, UPS output uninterrupted during transfer. | verification-requirements | Verification Requirements |  | Test |  |  | Power changeover timing is critical to UPS autonomy margin. Test must confirm the entire chain from detection through generator start to load transfer operates within design envelope. | verification, lab-utilities, session-335
VER-REQ-059 | Verify end-to-end utility failure response: Simulate total normal supply loss and verify the complete chain from UPS bridging through generator start, changeover switch operation, BMS alarm presentation, gas supply isolation (if fire alarm concurrent), and laboratory instrument continuity. Pass criteria: ICP-MS plasma maintained throughout, gamma spectrometry counts uninterrupted, LIMS server remains operational, BMS displays correct facility status within 30 seconds of initial failure. | verification-requirements | Verification Requirements |  | Test |  |  | System-level integration test exercising the most critical utility failure scenario. Confirms that the utility subsystem components interact correctly under the highest-consequence failure mode to maintain analytical capability. | verification, lab-utilities, session-335
VER-REQ-069 | Verify SUB-REQ-099: Conduct a proof test of all SIL 3 safety functions by injecting simulated trip conditions through the complete chain from sensor input to final element actuation. Confirm each function trips within its specified response time. Record all proof test results and confirm the test interval does not exceed 3 months since the previous proof test. | verification-requirements | Verification Requirements |  | Test |  |  | Proof testing validation ensures the claimed SIL is maintained throughout the facility operational life. The complete chain test verifies sensor, logic solver, and final element integrity rather than partial testing which could miss degraded components. | verification, safety, validation, session-337
VER-REQ-070 | Verify SUB-REQ-100: Review the safety system design documentation and confirm that each 2oo3 voting group employs at least two different sensing principles or manufacturers. Conduct a common cause failure analysis per IEC 61511 Annex E demonstrating that the beta factor for each voting group does not exceed 0.02 for SIL 3 functions. | verification-requirements | Verification Requirements |  | Analysis |  |  | CCF defence adequacy cannot be demonstrated by testing alone. Design review confirms diversity is implemented as specified. The beta factor analysis per IEC 61511 quantifies residual CCF vulnerability and must meet SIL 3 targets to validate the safety case claims. | verification, safety, ccf, validation, session-337
VER-REQ-071 | Verify SYS-REQ-011: Conduct penetration testing of the LIMS network segment by an independent assessor. Confirm RBAC enforces least-privilege access for each role. Verify MFA is required for all login sessions. Confirm encryption at rest by examining database storage configuration and in transit by capturing network traffic and verifying TLS 1.2 or later. Confirm network segmentation by attempting lateral movement from office and OT network segments. | verification-requirements | Verification Requirements |  | Test |  |  | Cybersecurity controls must be validated by independent testing, not self-assessment alone. Penetration testing is required under NIS Regulations assessment framework. The combination of configuration verification and active testing provides confidence that controls are both configured and effective against realistic attack scenarios. | verification, cybersecurity, validation, session-337

## Trace Links
Source Ref | Target Ref | Link Type | Description | Rationale | Created
SYS-REQ-012 | SUB-REQ-086 | derives | Decommissioning records requirement drives LIMS long-term data retention | The decommissioning database mandated by SYS-REQ-012 requires cumulative contamination history and material records spanning the full facility lifetime. The LIMS server and database is the primary system-of-record for these analytical results, requiring operational availability and data retention capacity sufficient to support eventual decommissioning characterisation | 2026-03-19T14:48:18.482Z
SYS-REQ-009 | SUB-REQ-098 | derives | Safety response drives BMS alarm prioritization for operator awareness | The facility safety case requires operators to distinguish between safety-critical alarms and routine maintenance alerts. Priority-ordered alarm presentation within 2 seconds ensures safety-relevant utility failures (e.g., ventilation loss, power failure) are not masked by lower-priority alarms during emergency conditions | 2026-03-19T14:48:15.034Z
SYS-REQ-007 | SUB-REQ-097 | derives | ICP-MS detection limit drives continuous DI water recirculation | Stagnant deionised water in distribution loops develops ionic and microbial contamination that introduces spectral interferences at sub-ppt levels, directly degrading ICP-MS detection limits specified in SYS-REQ-007. Continuous recirculation at 0.5 m/s prevents biofilm formation and maintains 18.2 MOhm-cm resistivity | 2026-03-19T14:48:11.453Z
SYS-REQ-009 | SUB-REQ-096 | derives | Facility safety drives automatic gas supply changeover for continuity | Nitrogen supply is required for glove box atmospheres and inerting in areas handling fissile materials. Loss of nitrogen supply could compromise containment, so automatic manifold changeover ensures continuous gas availability as part of the safety case | 2026-03-19T14:48:06.962Z
SYS-REQ-011 | SUB-REQ-093 | derives | LIMS network segmentation requirement extends to BMS OT-IT boundary | The system-level mandate for network segmentation isolating safety and safeguards systems in SYS-REQ-011 necessitates the BMS implementing equivalent OT-IT separation to prevent lateral movement from enterprise networks to building control systems | 2026-03-19T14:48:03.908Z
SYS-REQ-009 | SUB-REQ-092 | derives | Facility safety requirement drives automatic gas isolation on fire alarm | The two-independent-barriers safety approach in SYS-REQ-009 requires that fire detection triggers automatic isolation of flammable gas supplies as a defence-in-depth measure, preventing fire escalation in radiological areas | 2026-03-19T14:47:59.878Z
IFC-REQ-055 | VER-REQ-071 | verifies | Lateral movement test during penetration testing verifies unidirectional gateway enforcement | The penetration test explicitly includes lateral movement attempts from BMS to safety system, which directly verifies the data diode enforcement specified in IFC-REQ-055 | 2026-03-19T14:38:15.242Z
SYS-REQ-011 | VER-REQ-071 | verifies | Penetration testing and configuration audit verify LIMS cybersecurity controls | The cybersecurity requirement in SYS-REQ-011 requires independent verification through active testing to confirm controls are both implemented and effective against realistic threats | 2026-03-19T14:38:14.616Z
SYS-REQ-011 | IFC-REQ-055 | derives | LIMS network segmentation requirement extends to BMS-safety system boundary enforcement | The system-level network segmentation mandate in SYS-REQ-011 necessitates a defined interface between general-purpose BMS and safety-rated trip system, implemented as a unidirectional gateway to prevent cross-contamination of safety integrity | 2026-03-19T14:38:14.024Z
SUB-REQ-100 | VER-REQ-070 | verifies | CCF analysis validates diverse redundancy design of safety trip channels | Design review and beta factor analysis together verify that diversity in sensing is implemented and that residual CCF vulnerability meets SIL 3 targets | 2026-03-19T14:37:40.794Z
SUB-REQ-099 | VER-REQ-069 | verifies | Proof test procedure validates ongoing SIL compliance of safety functions | The proof test is the primary mechanism for demonstrating that the claimed SIL is maintained through the facility operational life | 2026-03-19T14:37:40.157Z
SYS-REQ-009 | SUB-REQ-100 | derives | Two-independent-barriers criticality safety requirement drives diverse redundancy in trip sensing | The two-barrier principle in SYS-REQ-009 requires that neither barrier is defeated by a common cause. Diverse sensing within the 2oo3 voting groups ensures the interlock barrier remains effective even under common-mode sensor degradation | 2026-03-19T14:37:12.081Z
SYS-REQ-009 | SUB-REQ-099 | derives | Fissile material safety controls require demonstrated ongoing reliability of trip functions through periodic proof testing | SYS-REQ-009 requires two independent barriers to criticality including the safety interlock system. The claimed SIL for trip functions is only valid when dangerous undetected failure probability is bounded by proof testing at defined intervals per IEC 61511 | 2026-03-19T14:36:47.164Z
STK-REQ-007 | SYS-REQ-011 | derives | Information security stakeholder need drives LIMS access control and encryption requirements | LIMS is the primary repository of safeguards-sensitive NMA data. The cybersecurity requirement directly implements the NISR security plan obligation for electronic systems containing classified nuclear material data | 2026-03-19T14:36:43.014Z
STK-REQ-006 | SYS-REQ-012 | derives | Decommissioning design requirement drives need for operational records to support future characterisation | The decommissioning database is a direct implementation of the LC35 design-for-decommissioning principle: systematic operational records reduce future characterisation effort and decommissioning waste volumes | 2026-03-19T14:36:40.287Z
SUB-REQ-038 | uht-bot:se-radiochem-lab-v2:VER-068 | verifies | Emergency power UPS and battery autonomy test | Mains failure simulation with load measurement proves seamless UPS transition and battery autonomy at 30-min full-load and 4-hour reduced-load. | 2026-03-19T14:25:51.868Z
SUB-REQ-010 | uht-bot:se-radiochem-lab-v2:VER-067 | verifies | CAM alpha sensitivity and reliability test | Check source challenge verifies 0.2 DAC-hours sensitivity; 1000-hour trial quantifies false alarm rate per monitor. | 2026-03-19T14:25:51.249Z
SUB-REQ-037 | uht-bot:se-radiochem-lab-v2:VER-066 | verifies | 2oo3 voting timing and independence test | Signal injection and oscilloscope measurement verifies 500 ms trip time; UPS feed removal proves channel independence. | 2026-03-19T14:25:50.624Z
SUB-REQ-036 | uht-bot:se-radiochem-lab-v2:VER-065 | verifies | Safety interlock SIL 3 proof test | Load cell and gamma trip testing with FMEDA review provides combined evidence of interlock function and SIL 3 integrity. | 2026-03-19T14:25:50.041Z
SUB-REQ-032 | uht-bot:se-radiochem-lab-v2:VER-060 | verifies | Criticality detection proof test | Physical neutron source test directly verifies the 1 ms detection time and 2oo3 coincidence logic of the criticality warning system. | 2026-03-19T14:25:49.456Z
SYS-REQ-009 | IFC-REQ-054 | derives | Facility safety drives power changeover interface between EPDS and emergency supply | Criticality and safety monitoring must not be interrupted during power failure; the automatic changeover interface between normal and emergency power ensures continuous supply to safety-critical systems including criticality warning and safety interlocks. | 2026-03-19T14:23:19.719Z
SYS-REQ-007 | IFC-REQ-053 | derives | ICP-MS detection limit drives DI water interface to acid digestion | Sub-ppt ICP-MS detection limits demand ultra-pure rinse and dilution water at the sample preparation stage; this interface ensures DI water quality is maintained up to the point of use in acid digestion. | 2026-03-19T14:23:19.130Z
SYS-REQ-006 | IFC-REQ-052 | derives | Radiation protection drives security-HP alarm zone data exchange | Radiation protection alarm response requires knowledge of personnel location; the access control to HP alarm interface enables targeted evacuation notifications and dose attribution to individuals present in affected zones. | 2026-03-19T14:23:18.551Z
SYS-REQ-002 | IFC-REQ-051 | derives | Ventilation cascade drives chilled water supply interface to AHU | Maintaining the depression cascade requires conditioned supply air at stable temperature; the chilled water interface to the AHU enables the temperature control that prevents thermal buoyancy effects from disrupting pressure differentials. | 2026-03-19T14:23:17.906Z
SYS-REQ-007 | SUB-REQ-091 | derives | ICP-MS detection limit drives DI water purity requirement | ICP-MS sub-ppt detection limits require ultra-pure water for blank preparation, sample dilution, and rinse cycles; DI water quality directly controls the analytical blank contribution to measurement uncertainty. | 2026-03-19T14:22:54.523Z
SYS-REQ-005 | SUB-REQ-089 | derives | LIMS availability drives electrical power reliability requirement | LIMS 99.5% availability requires reliable power supply; electrical distribution system UPS and changeover capabilities directly underpin the availability target for LIMS and all networked instruments. | 2026-03-19T14:22:53.901Z
SYS-REQ-004 | SUB-REQ-076 | derives | Alpha detection limit drives vacuum system contamination prevention | Cold trapping prevents oil vapour migration from vacuum pump to detector chamber, which would deposit on the silicon surface barrier detector and degrade energy resolution below the threshold needed for 5 mBq MDA. | 2026-03-19T14:22:53.333Z
SYS-REQ-004 | SUB-REQ-075 | derives | Alpha detection limit drives MCA spectral processing requirement | Tracer recovery and activity calculations are fundamental to converting raw alpha spectra into reportable results at the 5 mBq/sample detection limit with proper uncertainty quantification. | 2026-03-19T14:22:52.705Z
SYS-REQ-004 | SUB-REQ-074 | derives | Alpha detection limit drives vacuum chamber pressure requirement | Alpha particle energy resolution required for 5 mBq/sample detection limit demands vacuum better than 1.3 Pa to eliminate air absorption and energy straggling that would broaden peaks and degrade MDA. | 2026-03-19T14:22:52.101Z
SYS-REQ-001 | SUB-REQ-071 | derives | Gamma MDA drives QA verification of analysis results | QA checks on gamma analysis (duplicate blind re-analysis) ensure that reported MDA values are reliable and meet ISO 17025 quality requirements for the 0.1 Bq/sample detection threshold. | 2026-03-19T14:22:31.554Z
SYS-REQ-001 | SUB-REQ-070 | derives | Gamma MDA drives spectrum analysis automation requirement | Automated nuclide identification and activity calculation with library-based analysis is required to convert raw spectra into reportable MDA values meeting the 0.1 Bq/sample threshold with quantified uncertainties. | 2026-03-19T14:22:30.948Z
SYS-REQ-001 | SUB-REQ-068 | derives | Gamma MDA drives cryogenic temperature stability requirement | HPGe detector energy resolution required for MDA at 0.1 Bq/sample is only achievable when detectors are maintained at operating temperature; cryogenic cooling failure directly degrades resolution and MDA. | 2026-03-19T14:22:30.364Z
SYS-REQ-001 | SUB-REQ-066 | derives | Gamma MDA requirement drives minimum detector count | Achieving 0.1 Bq/sample MDA for Cs-137 requires multiple high-efficiency detectors operating in parallel to provide sufficient counting capacity and counting time per sample during high-throughput campaigns. | 2026-03-19T14:22:29.743Z
IFC-REQ-054 | VER-REQ-058 | verifies | Integration test for normal-to-emergency power changeover | Confirms power changeover chain operates within UPS autonomy envelope, preventing analytical instrument loss. | 2026-03-19T14:12:02.906Z
IFC-REQ-050 | VER-REQ-057 | verifies | Integration test for gas isolation on fire alarm | Fire safety case critical test confirming flammable gas isolation timing claim. | 2026-03-19T14:12:02.304Z
IFC-REQ-049 | VER-REQ-056 | verifies | Integration test for BMS-CHW BACnet/IP interface | Confirms duty/standby changeover maintains chilled water temperature during plant fault conditions. | 2026-03-19T14:12:01.694Z
IFC-REQ-048 | VER-REQ-055 | verifies | Integration test for BMS-EPDS Modbus TCP interface | Confirms Modbus TCP polling and load shedding response timing at the system integration boundary. | 2026-03-19T14:12:01.056Z
SYS-REQ-009 | SUB-REQ-094 | derives | Fissile material limits require controlled access to prevent unauthorised accumulation | Criticality safety depends on administrative controls including access restriction to fissile material handling areas. Tier 3 security for nuclear material accountability zones ensures only trained, authorised personnel can access areas where fissile material is present, supporting the mass limit enforcement. | 2026-03-19T14:11:50.669Z
SYS-REQ-002 | SUB-REQ-095 | derives | Depression cascade depends on stable supply air from AHU which needs stable CHW | Supply AHU temperature control affects air density and volume, which directly impacts the depression cascade pressure differentials. CHW temperature stability is a prerequisite for maintaining the supply air conditions that the ventilation system depends on. | 2026-03-19T14:11:50.053Z
SYS-REQ-007 | SUB-REQ-090 | derives | ICP-MS detection limit requires stable uninterrupted power | ICP-MS plasma extinction on power loss requires 30-minute restabilisation, destroying analytical batch continuity. The UPS requirement derives directly from the need to maintain plasma during power transients to achieve 0.05 Bq/L Tc-99 detection. | 2026-03-19T14:11:49.456Z
SYS-REQ-001 | SUB-REQ-090 | derives | Gamma MDA requires uninterrupted power for long counting periods | Achieving 0.5 Bq/L MDA for Cs-137 requires counting times of 12-72 hours. Power interruption during counting loses accumulated spectral data, requiring restart. UPS requirement directly derives from this analytical performance need. | 2026-03-19T14:11:48.826Z
IFC-REQ-047 | VER-REQ-054 | verifies | Integration test for LIMS-to-instrument bidirectional interface | Validates message integrity, queuing resilience, and throughput for the critical LIMS instrument integration under normal and degraded conditions. | 2026-03-19T13:59:20.481Z
SYS-REQ-005 | SUB-REQ-088 | derives | Regulatory reporting derives from system compliance requirement | Cumulative discharge tracking and automated RIFE submissions are the LIMS functions that directly support the system's regulatory compliance obligation under environmental permits. | 2026-03-19T13:59:19.857Z
SYS-REQ-005 | SUB-REQ-087 | derives | Instrument interface derives from system data integrity requirement | Automated bidirectional instrument interfaces are the primary mechanism by which LIMS eliminates transcription errors and maintains data integrity for the system of record. | 2026-03-19T13:59:19.227Z
SYS-REQ-005 | SUB-REQ-086 | derives | Database availability derives from system data management requirement | The system LIMS requirement mandates reliable data management for sample tracking and reporting. The 99.5% availability target and automatic failover are the component-level specifications that ensure the system-level data management commitment is operationally achievable. | 2026-03-19T13:59:18.594Z
SYS-REQ-010 | VER-REQ-051 | verifies | End-to-end integration test verifies system LSC detection limit | The end-to-end test with blind spikes validates the entire measurement chain achieves the 1 Bq/L H-3 detection limit in operational conditions, not just individual component specifications. | 2026-03-19T13:55:45.676Z
SYS-REQ-007 | SUB-REQ-085 | derives | Argon gas stability derives from ICP-MS detection limit | Plasma instability from gas flow variation causes signal noise that directly degrades detection limits. The 0.5% flow stability keeps signal drift within the range correctable by internal standard normalisation, preserving the 0.05 Bq/L Tc-99 MDA. | 2026-03-19T13:55:45.042Z
SYS-REQ-010 | SUB-REQ-082 | derives | Uncertainty calculation derives from detection limit verification | Detection limit calculation per Currie requires a valid uncertainty model. Without GUM-compliant uncertainty propagation, the claimed 1 Bq/L MDA cannot be verified or defended during regulatory audit. | 2026-03-19T13:55:44.438Z
SYS-REQ-010 | SUB-REQ-081 | derives | QC monitoring derives from measurement reliability assurance | Instrument drift or degradation could cause the detection limit to rise above 1 Bq/L without operator awareness. Daily QC with sealed standards provides early warning of performance degradation before it affects reported results. | 2026-03-19T13:55:43.854Z
SYS-REQ-010 | SUB-REQ-080 | derives | Dark adaptation derives from LSC background control needs | Chemiluminescence adds to the apparent background count rate. If luminescence contribution exceeds 0.5 CPM, the effective background rises above 2.5 CPM, degrading the achievable MDA below the system 1 Bq/L target. | 2026-03-19T13:55:43.278Z
SYS-REQ-010 | SUB-REQ-079 | derives | Cocktail preparation quality derives from LSC detection limit | Cocktail dispensing repeatability directly affects quench reproducibility, which propagates into the measurement uncertainty budget. Poor dispensing degrades the achievable MDA below the 1 Bq/L system requirement. | 2026-03-19T13:55:42.669Z
IFC-REQ-046 | VER-REQ-053 | verifies | Data transfer test for ICP-MS instrument to workstation | Validates spectral data pipeline integrity and resilience for the ICP-MS analytical workflow. | 2026-03-19T13:54:45.306Z
IFC-REQ-045 | VER-REQ-052 | verifies | Interface test for separations-to-ICP-MS sample transfer | Validates matrix compatibility and physical handover between radiochemical separations and ICP-MS autosampler. | 2026-03-19T13:54:44.658Z
SYS-REQ-007 | SUB-REQ-084 | derives | Isotope ratio precision derives from waste characterisation requirements | The system Tc-99 detection limit requirement implies the suite must perform high-quality mass spectrometric analysis. Isotope ratio capability for U and Pu is an integral part of the suite's dockyard mission — characterising decommissioning waste by isotopic composition. | 2026-03-19T13:54:44.042Z
SYS-REQ-007 | SUB-REQ-083 | derives | Memory effect control derives from Tc-99 detection limit | Achieving 0.05 Bq/L Tc-99 detection limit requires negligible carryover from preceding high-concentration samples. The 0.01% carryover target ensures prior-sample contribution is below the instrument detection limit. | 2026-03-19T13:54:43.402Z
IFC-REQ-044 | VER-REQ-050 | verifies | LIMS transfer verification for analysis-to-LIMS interface | Validates the regulated data path to the system of record including integrity controls. | 2026-03-19T13:50:47.705Z
IFC-REQ-043 | VER-REQ-049 | verifies | Data transfer integrity test for counter-to-analysis interface | Validates completeness and correctness of spectral data transfer that underpins all result calculations. | 2026-03-19T13:50:47.098Z
IFC-REQ-042 | VER-REQ-048 | verifies | Interface test for sample prep to counter loading | Validates mechanical compatibility and barcode system at the prep-counter interface. | 2026-03-19T13:50:46.486Z
IFC-REQ-041 | VER-REQ-047 | verifies | Interface test for separations-to-LSC handover | Validates physical and documentation integrity of the fraction transfer process at the subsystem boundary. | 2026-03-19T13:50:45.872Z
SYS-REQ-010 | SUB-REQ-078 | derives | Alpha/beta discrimination requirement derives from detection limit in dockyard matrices | Dockyard coolant samples contain alpha emitters. Without 99.95% alpha rejection, the effective H-3 background increases, degrading the achievable detection limit below the 1 Bq/L system requirement. | 2026-03-19T13:50:45.254Z
SYS-REQ-010 | SUB-REQ-077 | derives | Counter efficiency and background requirements derive from system detection limit | The 1 Bq/L H-3 detection limit at 120 min count time requires minimum 65% efficiency and maximum 2 CPM background per the Currie MDA formula. These are the component-level performance figures that enable the system detection limit. | 2026-03-19T13:50:44.622Z
SYS-REQ-004 | VER-REQ-046 | verifies | End-to-end alpha spectrometry system integration test | Validates that the combined electrodeposition, detector, vacuum, and MCA chain meets the 0.5 mBq alpha detection limit system requirement. | 2026-03-19T13:28:38.252Z
IFC-REQ-040 | VER-REQ-045 | verifies | Vacuum interlock independence test per channel | Tests that vacuum interlock operates independently per chamber to maintain throughput during single-chamber servicing. | 2026-03-19T13:28:37.674Z
IFC-REQ-039 | VER-REQ-044 | verifies | Physical inspection of planchet-chamber interface | Confirms mechanical compatibility and identification integrity at the source-detector physical boundary. | 2026-03-19T13:28:37.088Z
SYS-REQ-004 | SUB-REQ-073 | derives | System detection limit cascades to source thickness requirement | Source thickness determines peak shape and resolution. Thick sources degrade resolution below the level needed to achieve quantitative separation of closely-spaced actinide peaks required by SYS-REQ-004. | 2026-03-19T13:27:59.187Z
SYS-REQ-004 | SUB-REQ-072 | derives | System alpha detection limit cascades to detector resolution requirement | The 0.5 mBq detection limit for Pu-239/240 requires adequate energy resolution to separate the peak from Pu-238 and Am-241 interferences, and sufficient counting efficiency over 250000-second acquisitions. | 2026-03-19T13:27:58.530Z
SYS-REQ-001 | VER-REQ-043 | verifies | End-to-end gamma spectrometry system integration test | Validates that combined detector, shielding, MCA, and analysis chain meets the 0.5 Bq/L MDA system requirement. | 2026-03-19T13:25:45.697Z
IFC-REQ-038 | VER-REQ-042 | verifies | End-to-end result transfer demonstration to LIMS | Demonstrates automated result posting eliminates transcription errors. | 2026-03-19T13:25:45.098Z
IFC-REQ-037 | VER-REQ-041 | verifies | Network transfer test for MCA-Analysis data path | Tests worst-case concurrent transfer load on dedicated VLAN. | 2026-03-19T13:25:44.486Z
IFC-REQ-036 | VER-REQ-040 | verifies | Pulser test for HPGe-MCA signal interface | Direct test of signal quality and cable plant performance at the detector-to-MCA boundary. | 2026-03-19T13:25:43.894Z
SYS-REQ-001 | SUB-REQ-069 | derives | System MDA requirement cascades to MCA linearity | ADC nonlinearity directly biases peak area integration. 0.05 percent linearity ensures peak area systematic error contributes less than 0.5 percent to total measurement uncertainty. | 2026-03-19T13:24:55.304Z
SYS-REQ-001 | SUB-REQ-067 | derives | System MDA requirement cascades to shielding background reduction | MDA is directly proportional to the square root of background count rate. The background reduction specified ensures counting times remain practical (under 60000 seconds) for achieving 0.5 Bq/L on 1-litre samples. | 2026-03-19T13:24:54.691Z
SYS-REQ-001 | SUB-REQ-065 | derives | System MDA requirement cascades to detector energy resolution | Achieving 0.5 Bq/L MDA for Cs-137 requires resolving the 661.7 keV peak from Compton continuum of higher-energy nuclides. 2.0 keV resolution is the minimum needed for reliable peak identification and quantification in complex reactor spectra. | 2026-03-19T13:24:54.061Z
SYS-REQ-005 | VER-REQ-036 | verifies | End-to-end verification of sample receipt workflow | This test verifies the complete sample receipt process from physical arrival through screening, registration, and chain-of-custody, validating the LIMS audit trail required by SYS-REQ-005 | 2026-03-19T13:06:59.061Z
SYS-REQ-009 | VER-REQ-032 | verifies | End-to-end verification of waste lifecycle with fissile tracking | This test verifies the complete waste processing chain from characterisation through packaging to storage, including nuclear material accountancy controls required by SYS-REQ-009 | 2026-03-19T13:06:58.404Z
SYS-REQ-009 | VER-REQ-025 | verifies | End-to-end verification of criticality response chain | This test verifies the complete criticality response from detection through interlock trip, alarm, and evacuation initiation, validating the safety function in SYS-REQ-009 | 2026-03-19T13:06:57.812Z
SYS-REQ-003 | VER-REQ-018 | verifies | End-to-end verification of effluent treatment batch cycle | This test verifies the complete batch-before-discharge process from drain collection through treatment, sampling, and controlled release, validating SYS-REQ-003 | 2026-03-19T13:06:57.236Z
SYS-REQ-006 | VER-REQ-012 | verifies | End-to-end verification of radiation protection chain | This end-to-end test verifies the complete radiation protection monitoring system from detection through to alarm and display, validating the overall capability required by SYS-REQ-006 | 2026-03-19T13:06:56.650Z
SYS-REQ-005 | SUB-REQ-051 | derives | Waste records integrate with LIMS for material accountancy | Waste package records must link to LIMS analytical records to maintain nuclear material accountancy — SYS-REQ-005 mandates tamper-evident records and the waste consignment system closes the material balance loop | 2026-03-19T13:05:57.415Z
SYS-REQ-009 | SUB-REQ-050 | derives | Waste store capacity accommodates fissile waste generation | Storage capacity must accommodate the expected waste generation rate while maintaining criticality-safe spacing — insufficient capacity forces unsafe interim storage arrangements | 2026-03-19T13:05:56.802Z
SYS-REQ-009 | SUB-REQ-049 | derives | Waste store environmental monitoring for fissile safety | Continuous monitoring of the interim waste store detects conditions (moisture ingress, temperature excursion) that could compromise the criticality-safe spacing and container integrity required for fissile material storage | 2026-03-19T13:05:56.215Z
SYS-REQ-009 | SUB-REQ-047 | derives | Surface dose rate verification on waste packages | Package dose rate measurement confirms that waste has been correctly characterised and the fissile content is within safe limits — mispackaged fissile material violates the controls in SYS-REQ-009 | 2026-03-19T13:05:55.577Z
SYS-REQ-009 | SUB-REQ-046 | derives | Waste packaging ensures safe geometry for fissile waste | Solid waste compaction must verify package dose rate meets transport limits — fissile-bearing waste packages must maintain criticality-safe geometry per SYS-REQ-009 mass limit controls | 2026-03-19T13:05:54.947Z
SYS-REQ-009 | SUB-REQ-043 | derives | Alpha/beta screening segregates fissile waste | Contact screening determines whether waste contains significant alpha activity (indicating fissile material) and drives waste route assignment — misclassification would defeat the fissile material mass limit controls in SYS-REQ-009 | 2026-03-19T13:05:40.800Z
SYS-REQ-009 | SUB-REQ-042 | derives | Waste characterisation enforces fissile material tracking | Gamma measurement of waste packages determines fissile content — this is required for the nuclear material accountancy component of fissile material control in SYS-REQ-009 | 2026-03-19T13:05:40.208Z
SYS-REQ-005 | SUB-REQ-059 | derives | Sample archive conditions preserve re-analysis capability | LIMS records reference archived samples for verification — the storage environment must preserve sample integrity so that re-analysis against LIMS records is meaningful, supporting the audit trail in SYS-REQ-005 | 2026-03-19T13:05:39.499Z
SYS-REQ-005 | SUB-REQ-057 | derives | Radiotracer spiking supports LIMS quality assurance | Calibrated tracer spikes are the basis for chemical recovery measurement — LIMS must record spike details to maintain the tamper-evident audit trail required by SYS-REQ-005 for each analytical result | 2026-03-19T13:05:38.896Z
SYS-REQ-005 | SUB-REQ-056 | derives | Sample preparation feeds LIMS audit trail | The acid digestion step creates sample aliquots that must be tracked through the LIMS chain of custody — SYS-REQ-005 mandates tamper-evident records, and sample prep is where the original sample is transformed into countable fractions | 2026-03-19T13:05:38.307Z
SYS-REQ-010 | SUB-REQ-063 | derives | Tritium/C-14 preparation for LSC measurement | Tritium requires distillation and C-14 requires combustion/absorption before LSC counting — these preparation steps determine whether the detection limits in SYS-REQ-010 are achievable for these low-energy beta emitters | 2026-03-19T13:05:22.078Z
SYS-REQ-010 | SUB-REQ-062 | derives | Sr/Cs separation provides purified fractions for LSC | Liquid scintillation counting of Sr-90 requires radiochemical separation from interfering beta emitters — the separation quality directly determines whether the 1 Bq/L detection limit in SYS-REQ-010 can be achieved | 2026-03-19T13:05:21.469Z
SYS-REQ-004 | SUB-REQ-060 | derives | Sample dissolution is prerequisite for alpha spectrometry | Solid samples must be completely dissolved before actinide separation and alpha spectrometry — incomplete dissolution causes low recovery, pushing detection limits above the 0.5 mBq specification in SYS-REQ-004 | 2026-03-19T13:05:20.858Z
SYS-REQ-003 | SUB-REQ-045 | derives | Waste immobilisation for safe long-term storage | Concentrated liquid waste must be cemented to prevent leaching and create a passively safe wasteform — this secondary waste arises directly from the chemical treatment process that SYS-REQ-003 mandates | 2026-03-19T13:05:20.238Z
SYS-REQ-003 | SUB-REQ-044 | derives | Liquid waste conditioning reduces volume for storage | Concentrated liquid waste from effluent treatment must be immobilised for disposal — volume reduction is necessary because the waste arises from the batch treatment process mandated by SYS-REQ-003 | 2026-03-19T13:05:19.591Z
SYS-REQ-006 | SUB-REQ-054 | derives | Screening station detects contamination at facility boundary | The dose rate and contamination screening station at sample receipt is the first radiation protection measurement point — detecting elevated samples before they enter the laboratory supports the overall RP monitoring objective in SYS-REQ-006 | 2026-03-19T13:04:59.427Z
SYS-REQ-009 | SUB-REQ-041 | derives | Emergency communications for criticality and safety events | The emergency comms system provides the means to order evacuation during a criticality incident — without reliable, clearly distinguishable alarms the personnel protection objective of SYS-REQ-009 cannot be met | 2026-03-19T13:04:58.805Z
SYS-REQ-009 | SUB-REQ-040 | derives | Spill containment prevents uncontrolled fissile material spread | A radioactive liquid spill containing fissile material must be contained to prevent accumulation in an unsafe geometry — bunded containment directly supports the criticality safety controls in SYS-REQ-009 | 2026-03-19T13:04:58.188Z
SYS-REQ-009 | SUB-REQ-039 | derives | Diesel generator backs up UPS for extended outages | UPS provides ride-through but has limited duration — the diesel generator ensures long-duration power for safety systems including criticality controls required by SYS-REQ-009 | 2026-03-19T13:04:46.469Z
SYS-REQ-009 | SUB-REQ-038 | derives | Emergency power ensures safety systems survive mains loss | Criticality warning, safety interlocks, and radiation monitoring must operate continuously — SYS-REQ-009 mandates fissile material controls that require uninterrupted power to remain functional | 2026-03-19T13:04:45.856Z
SYS-REQ-009 | SUB-REQ-037 | derives | 2oo3 voting ensures reliable safety trip function | The 2oo3 voting architecture prevents single-point failures from defeating the fissile material mass limit enforcement required by SYS-REQ-009 while minimising spurious trips that disrupt operations | 2026-03-19T13:04:45.190Z
SYS-REQ-009 | SUB-REQ-035 | derives | Inert gas suppression protects fissile material areas | Water-based suppression is prohibited in fissile material areas due to neutron moderation risk — inert gas protects criticality safety barriers without introducing a criticality hazard | 2026-03-19T13:04:44.546Z
SYS-REQ-009 | SUB-REQ-034 | derives | Fire detection supports safety system integrity | Fire threatens the integrity of criticality controls (fissile material storage geometry, safety interlocks) — rapid detection and suppression prevents fire-induced loss of criticality safety barriers required by SYS-REQ-009 | 2026-03-19T13:04:43.897Z
SYS-REQ-009 | SUB-REQ-033 | derives | Criticality warning alarm from safety requirement | The audible alarm in SUB-REQ-033 is the personnel protection measure triggered by the criticality detection system — SYS-REQ-009 mandates the overall criticality safety function, the alarm ensures personnel evacuate | 2026-03-19T13:04:43.305Z
SYS-REQ-003 | SUB-REQ-031 | derives | Treatment sludge packaging from effluent treatment | Chemical treatment produces secondary waste (sludge) that must be safely packaged — this is a direct consequence of the treatment process mandated by SYS-REQ-003 to achieve discharge limits | 2026-03-19T13:04:26.255Z
SYS-REQ-003 | SUB-REQ-030 | derives | Active drain material specification for effluent integrity | 316L stainless steel construction prevents corrosion-induced leaks that would cause unmonitored effluent release, bypassing the controlled batch discharge process in SYS-REQ-003 | 2026-03-19T13:04:25.659Z
SYS-REQ-003 | SUB-REQ-029 | derives | Contamination detection prevents uncontrolled discharge | The inactive drain diversion system is the first barrier ensuring contaminated water enters the active treatment system rather than bypassing it — failure defeats the batch-before-discharge control in SYS-REQ-003 | 2026-03-19T13:04:24.967Z
SYS-REQ-003 | SUB-REQ-024 | derives | Effluent containment requirement from discharge standards | Double containment prevents unmonitored release of radioactive effluent — a tank leak bypassing the batch sampling and analysis requirement in SYS-REQ-003 would constitute an uncontrolled discharge | 2026-03-19T13:04:14.216Z
SYS-REQ-003 | SUB-REQ-023 | derives | Batch analysis drives tank sizing and hold time | SYS-REQ-003 mandates sampling every batch before discharge — the 24-hour minimum hold time in SUB-REQ-023 exists to provide adequate time for laboratory analysis turnaround before authorising release | 2026-03-19T13:04:13.592Z
SYS-REQ-008 | SUB-REQ-007 | derives | HEPA twin-bank design requires standby ventilation | The standby ventilation system maintains containment during HEPA bank changeover and fan failure, directly supporting the twin-bank architecture decision in SYS-REQ-008 | 2026-03-19T13:04:03.368Z
SYS-REQ-002 | SUB-REQ-005 | derives | Depression cascade drives stack monitoring requirements | The stack is the final discharge point of the depression cascade — monitoring ensures contaminated air that has been drawn through the cascade and filtered is only released within authorised limits | 2026-03-19T13:04:02.752Z
SYS-REQ-002 | SUB-REQ-001 | derives | Depression cascade requires conditioned supply air | The supply AHU delivers the makeup air that the extract system draws through the zone boundaries to maintain the depression cascade — without controlled supply, pressure differentials cannot be sustained | 2026-03-19T13:04:02.116Z
SYS-REQ-002 | SUB-REQ-064 | derives | Depression cascade drives fume cupboard extract requirements | Fume cupboard face velocity and extract airflow directly contribute to the facility depression cascade — each cupboard extract is a load on the C3 zone extract system | 2026-03-19T12:48:17.874Z
SYS-REQ-004 | SUB-REQ-061 | derives | Alpha spectrometry detection limits drive actinide separation recovery requirements | The system-level 0.5 mBq detection limit for Pu-239/240 can only be achieved if the separation chemistry delivers at least 70% recovery with adequate spectral purity — losses below this threshold push the MDA above specification | 2026-03-19T12:48:17.238Z
IFC-REQ-035 | VER-REQ-039 | verifies | Waste segregation compliance inspection | Verifies waste stream separation and traceability during operational simulation | 2026-03-19T12:48:16.610Z
IFC-REQ-034 | VER-REQ-038 | verifies | Source quality test for separation-to-counting interface | Verifies the electrodeposited source meets mass and resolution requirements for alpha spectrometry | 2026-03-19T12:48:16.028Z
IFC-REQ-033 | VER-REQ-037 | verifies | Acid matrix compatibility test for digestion-to-separation interface | Verifies the feed solution matrix is within the operating envelope for downstream chromatography columns | 2026-03-19T12:48:15.405Z
SYS-REQ-002 | SUB-REQ-053 | derives | Facility depression cascade drives receiving bay negative pressure | The receiving bay at the C2/C3 boundary must maintain the depression cascade specified at system level to prevent contamination migration during sample transfer operations | 2026-03-19T12:43:12.660Z
SYS-REQ-009 | SUB-REQ-058 | derives | Facility criticality limit cascades to sample storage geometry | The system-level fissile mass limit enforcement requirement drives the specific keff and mass limits in the sample storage facility where multiple fissile-bearing samples accumulate | 2026-03-19T12:43:12.055Z
SYS-REQ-005 | SUB-REQ-055 | derives | LIMS audit trail requirement drives chain-of-custody registration | The system-level requirement for tamper-evident audit trails directly cascades to the sample registration system which is the point-of-origin for all sample data entering LIMS | 2026-03-19T12:43:11.453Z
IFC-REQ-032 | VER-REQ-035 | verifies | Message queue resilience test for registration-to-LIMS interface | LIMS outage is a credible operational scenario — the registration system must continue independently without data loss | 2026-03-19T12:42:59.770Z
IFC-REQ-031 | VER-REQ-034 | verifies | Data transfer timing and dual routing test for screening-to-preparation interface | Electronic data latency and physical routing correctness are both safety-significant — a remote-handled sample arriving at a contact-handled bench is a dose uptake incident | 2026-03-19T12:42:59.165Z
IFC-REQ-030 | VER-REQ-033 | verifies | Airlock interlock and clearance test for receiving-to-screening interface | Physical hatch interlock is the primary containment barrier during sample transfer — must be tested before active commissioning | 2026-03-19T12:42:58.520Z
SYS-REQ-005 | IFC-REQ-028 | derives | LIMS audit trail requirement drives waste records integration | The LIMS tamper-evident audit trail (SYS-REQ-005) extends to waste characterisation data that flows between LIMS and the Waste Records system. The bidirectional API with mutual TLS in IFC-REQ-028 ensures that the chain of custody and data integrity required by SYS-REQ-005 is maintained across the system boundary. | 2026-03-19T12:28:16.393Z
SYS-REQ-003 | IFC-REQ-027 | derives | Effluent sampling requirement drives distillate return activity limits | The distillate returned from waste conditioning enters the active drain system that feeds the AETDS. SYS-REQ-003 requires every batch to be sampled before discharge — the activity limits in IFC-REQ-027 ensure that returned distillate does not overload the AETDS chemical treatment capacity or compromise the site's ability to meet discharge authorisation limits. | 2026-03-19T12:28:15.800Z
SYS-REQ-009 | SUB-REQ-052 | derives | System fissile limits require NMA tracking in waste management | Fissile material limits cannot be enforced without accurate accountancy. SUB-REQ-052 provides the tracking mechanism that enables verification of compliance with SYS-REQ-009 as material moves through the waste facility. | 2026-03-19T12:28:15.237Z
SYS-REQ-009 | SUB-REQ-048 | derives | System fissile limits cascade to interim store criticality spacing | The system-level fissile mass limit requirement (SYS-REQ-009) directly drives the need for criticality-safe spacing in the Interim Waste Store. The store is where fissile waste packages accumulate, making it the most criticality-significant location in the waste facility. The keff limit and spacing requirements in SUB-REQ-048 are the engineering implementation of the system fissile control principle. | 2026-03-19T12:28:14.569Z
IFC-REQ-029 | VER-REQ-031 | verifies | Transfer interlock test for CTP to liquid conditioning | Verifies the credited safety interlock preventing uncontrolled transfer of radioactive concentrates | 2026-03-19T12:28:05.995Z
IFC-REQ-028 | VER-REQ-030 | verifies | Bidirectional data exchange test between Waste Records and LIMS | Verifies the information system integration underpinning regulatory record-keeping for waste packages | 2026-03-19T12:28:05.418Z
IFC-REQ-027 | VER-REQ-029 | verifies | Test for distillate return activity limits and hold-and-release interlock | Verifies the environmental protection boundary between waste conditioning and the active drain system | 2026-03-19T12:28:04.821Z
IFC-REQ-026 | VER-REQ-028 | verifies | Integration test for cemented wasteform transfer to store | Verifies the 7-day curing hold, strength test recording, and criticality-compliant placement for the most safety-significant waste transfer | 2026-03-19T12:28:04.189Z
IFC-REQ-025 | VER-REQ-027 | verifies | Integration test for packaging-to-store drum transfer | Verifies barcode identification, contamination verification, and criticality-safe placement as a complete transfer chain | 2026-03-19T12:28:03.588Z
IFC-REQ-024 | VER-REQ-026 | verifies | Integration test for characterization-to-packaging data transfer | Verifies the 5-minute data transfer latency and categorisation certificate completeness at the characterization-packaging boundary | 2026-03-19T12:28:03.005Z
IFC-REQ-023 | VER-REQ-024 | verifies | Drainage segregation inspection | Physical verification of no cross-connection to inactive drains — the fundamental safety claim of IFC-REQ-023 | 2026-03-19T12:12:22.877Z
IFC-REQ-022 | VER-REQ-023 | verifies | Partial isolation containment test | Tests that depression is maintained in non-affected zones during partial isolation — the key safety claim of IFC-REQ-022 | 2026-03-19T12:12:22.264Z
IFC-REQ-021 | VER-REQ-022 | verifies | Power changeover test for CWS dual-redundant UPS | Tests sub-millisecond changeover and load sustainment — verifies zero detection gap claim in IFC-REQ-021 | 2026-03-19T12:12:21.651Z
IFC-REQ-020 | VER-REQ-021 | verifies | Integration test for fire-to-ventilation damper command | Tests command latency and cable fire survivability — the two key acceptance criteria of IFC-REQ-020 | 2026-03-19T12:12:21.037Z
IFC-REQ-019 | VER-REQ-020 | verifies | Integration test for CWS-SIT hardwired interface | Tests relay contact timing and galvanic isolation — the two measurable acceptance criteria of IFC-REQ-019 | 2026-03-19T12:12:20.404Z
SYS-REQ-009 | SUB-REQ-032 | derives | Criticality detection as defence-in-depth behind mass limits | SYS-REQ-009 requires fissile mass limits preventing criticality. SUB-REQ-032 provides the defence-in-depth detection layer — if mass limits are defeated (procedural error, equipment failure), the criticality warning system detects the excursion and triggers evacuation. This is a ALARP measure required by ONR SAPs. | 2026-03-19T12:10:06.082Z
SYS-REQ-009 | SUB-REQ-036 | derives | Fissile mass limit enforcement via safety interlocks | SYS-REQ-009 requires enforcement of fissile mass limits per workstation. SUB-REQ-036 implements this through load cells and gamma monitors with automatic trip at 80 percent of single-contingency limit. This is the primary implementation mechanism for the system-level criticality safety requirement. | 2026-03-19T12:10:05.507Z
IFC-REQ-018 | VER-REQ-019 | verifies | Inspection verification for ILW sludge drum transfer interface | Verifies physical package compliance (dose rate, documentation) at the treatment-to-waste-facility boundary. | 2026-03-19T11:45:21.741Z
SYS-REQ-003 | SUB-REQ-027 | derives | Discharge isolation protects the authorisation requirement | SYS-REQ-003 requires analysis before authorising discharge. SUB-REQ-027 is the safety backstop — if discharge monitoring detects above-limit activity during an authorised discharge, isolation prevents environmental release. This directly protects the integrity of the authorisation process. | 2026-03-19T11:45:10.473Z
IFC-REQ-017 | VER-REQ-017 | verifies | OPC UA data transfer latency and integrity test | Validates both functional data path and security mechanisms for the sampling-to-LIMS interface. | 2026-03-19T11:44:23.213Z
IFC-REQ-015 | VER-REQ-016 | verifies | Treatment loop transfer and DF verification | End-to-end treatment verification confirms both hydraulic interface and treatment efficacy. | 2026-03-19T11:44:22.596Z
IFC-REQ-016 | VER-REQ-015 | verifies | SIL 2 proof test for discharge isolation safety function | Tests all three trip conditions (high activity, signal loss, power loss) and verifies final element response time <5 seconds. | 2026-03-19T11:44:22.001Z
IFC-REQ-014 | VER-REQ-014 | verifies | Sample line purge and cross-contamination test | Validates that sample lines deliver representative aliquots and do not cross-contaminate between tanks. | 2026-03-19T11:44:21.425Z
IFC-REQ-013 | VER-REQ-013 | verifies | Integration test for drain-to-tank flow interface | Hydraulic test verifies flow capacity, totalisation accuracy, and valve operation at the drain collection boundary. | 2026-03-19T11:44:20.810Z
SYS-REQ-003 | SUB-REQ-025 | derives | Treatment capability ensures batches can meet discharge limits | SYS-REQ-003 implies batches must be dischargeable. SUB-REQ-025 provides the treatment capability to bring above-limit batches within discharge authorisation. Without treatment, high-activity batches would be stranded. | 2026-03-19T11:42:33.518Z
SYS-REQ-003 | SUB-REQ-028 | derives | Discharge recording implements the results recording requirement | SYS-REQ-003 explicitly requires results recorded against the batch in the discharge database. SUB-REQ-028 implements this by specifying batch ID, volume, activity, date/time, and operator. | 2026-03-19T11:42:32.902Z
SYS-REQ-003 | SUB-REQ-026 | derives | Sampling accuracy directly implements batch analysis capability | SYS-REQ-003 specifies detection limits for alpha, beta, and tritium. Accurate sampling is prerequisite — if samples are not representative, analysis results are meaningless regardless of instrument sensitivity. | 2026-03-19T11:42:32.275Z
SYS-REQ-003 | SUB-REQ-022 | derives | Hold time enables batch characterisation before discharge | SYS-REQ-003 requires sampling and analysis of every batch before discharge. The 24-hour hold ensures sufficient time for sample collection, laboratory analysis, and result review. | 2026-03-19T11:42:31.653Z
SYS-REQ-006 | SUB-REQ-017 | derives | RP detection function requires dose recording and statutory reporting | SYS-REQ-006 detection of radiation is only meaningful if individual doses are recorded and reported. SUB-REQ-017 CADOR integration implements the legal dose of record submission that makes the detection function operationally complete. | 2026-03-19T11:27:36.352Z
SYS-REQ-006 | SUB-REQ-021 | derives | RP system detection capability requires calibrated instruments and managed sources | SYS-REQ-006 RP detection capability depends on instrument calibration accuracy. SUB-REQ-021 source inventory and calibration management ensures the detection instruments referenced in SYS-REQ-006 remain within calibration and that calibration sources are properly managed under IRR17. | 2026-03-19T11:27:35.647Z
SYS-REQ-005 | IFC-REQ-011 | derives | LIMS integration requirement drives HP Central to LIMS data interface | SYS-REQ-005 audit trail and sample tracking requires radiation monitoring data to be associated with sample records. IFC-REQ-011 specifies the OPC-UA interface enabling this data flow from HP Central to LIMS. | 2026-03-19T11:27:18.238Z
SYS-REQ-006 | SUB-REQ-020 | derives | Continuous RP detection requires UPS-backed central monitoring | SYS-REQ-006 implies continuous detection capability. SUB-REQ-020 UPS requirement ensures the central system maintains this function during mains power loss, which is a credible coincident event with radiation incidents. | 2026-03-19T11:27:17.188Z
SYS-REQ-005 | SUB-REQ-019 | derives | LIMS audit trail requirement cascades to HP Central historian and statutory returns | SYS-REQ-005 tamper-evident audit trail for the laboratory extends to radiation monitoring records. SUB-REQ-019 implements the 10-year historian and statutory returns generation that provides the radiation monitoring component of the laboratory audit trail. | 2026-03-19T11:27:16.595Z
SYS-REQ-006 | SUB-REQ-015 | derives | RP detection extends to controlled area exit contamination screening | SYS-REQ-006 detection scope includes surface contamination at controlled area boundaries. SUB-REQ-015 specifies the portal monitor capability that forms the final contamination barrier at the controlled area exit. | 2026-03-19T11:27:15.977Z
SYS-REQ-006 | SUB-REQ-013 | derives | RP detection requirement cascades to area gamma alarm threshold specification | SYS-REQ-006 detection capability requires defined alarm thresholds. SUB-REQ-013 specifies the three-tier gamma alarm levels that operationalise dose rate detection into graded response actions. | 2026-03-19T11:27:15.500Z
IFC-REQ-012 | VER-REQ-011 | verifies | Hardwired evacuation relay independence test | Safety-critical test proving deterministic evacuation alarm activation independent of digital network | 2026-03-19T11:26:20.420Z
IFC-REQ-010 | VER-REQ-010 | verifies | OPC-UA dosimetry data and alarm notification test | End-to-end test of EPD data path including routine polling and immediate alarm notification | 2026-03-19T11:26:19.784Z
IFC-REQ-009 | VER-REQ-009 | verifies | Latency test for contamination monitoring data path | Confirms 2-second latency requirement met statistically across all contamination monitoring stations | 2026-03-19T11:26:19.185Z
IFC-REQ-008 | VER-REQ-008 | verifies | Dual-path diversity test for area gamma to HP Central | Verifies both analogue 4-20mA and digital Modbus paths independently, with analogue path independence test for SIL 1 function | 2026-03-19T11:26:18.591Z
IFC-REQ-007 | VER-REQ-007 | verifies | Integration test for CAM-to-HP Central Modbus interface | Tests polling interval, data completeness, and alarm state propagation via RS-485 Modbus | 2026-03-19T11:26:17.993Z
SYS-REQ-006 | SUB-REQ-018 | derives | System RP detection drives central monitoring system availability requirement | SYS-REQ-006 detection capability is only effective if the central system aggregating monitoring data maintains high availability. SUB-REQ-018 99.9 percent availability ensures the system-level detection function is continuously operational. | 2026-03-19T11:25:22.724Z
SYS-REQ-006 | SUB-REQ-016 | derives | System RP detection flows to personal dosimetry alarm and dose constraint management | SYS-REQ-006 radiation detection includes personal dose measurement. SUB-REQ-016 implements this through EPD alarm thresholds and dose constraint tracking that directly controls individual exposure. | 2026-03-19T11:25:22.142Z
STK-REQ-004 | SYS-REQ-006 | derives | Stakeholder dose constraint drives RP monitoring system detection requirements | STK-REQ-004 establishes the 6 mSv annual dose constraint. SYS-REQ-006 RP detection capability exists specifically to ensure doses are measured and controlled below this constraint — without detection there is no dose management. | 2026-03-19T11:25:21.477Z
SYS-REQ-006 | SUB-REQ-014 | derives | System detection requirement flows to personnel contamination monitoring sensitivity | SYS-REQ-006 radiation detection must cover both airborne and surface contamination pathways. SUB-REQ-014 implements the surface contamination component through HFC monitor sensitivity requirements at zone boundaries. | 2026-03-19T11:25:15.038Z
SYS-REQ-006 | SUB-REQ-012 | derives | System RP detection requirement flows to area gamma measurement specification | SYS-REQ-006 encompasses the full radiation detection capability. SUB-REQ-012 specifies the area gamma measurement range and energy response needed to complement airborne monitoring with external dose rate surveillance. | 2026-03-19T11:25:14.442Z
SYS-REQ-006 | SUB-REQ-011 | derives | System detection requirement cascades to CAM alarm response levels | SYS-REQ-006 detection capability requires not just measurement but actionable alarm response. SUB-REQ-011 defines the three-tier alarm cascade that operationalises the detection into investigation, action, and evacuation responses. | 2026-03-19T11:25:13.821Z
SYS-REQ-006 | SUB-REQ-010 | derives | System airborne alpha detection requirement flows down to CAM network sensitivity | SYS-REQ-006 specifies the airborne alpha contamination detection capability. SUB-REQ-010 implements this by specifying the CAM network DAC-fraction sensitivity and false alarm rate needed to achieve the system-level detection threshold. | 2026-03-19T11:25:13.186Z
SUB-REQ-003 | VER-REQ-006 | verifies | End-to-end tracer test validates HEPA and containment chain | The DOP tracer test measures actual decontamination factor across both HEPA banks under operational conditions, verifying the per-bank efficiency specified in SUB-REQ-003 as part of the integrated containment chain. | 2026-03-19T11:11:06.930Z
SYS-REQ-002 | IFC-REQ-005 | derives | Depression cascade requires supply-extract coordination interface | SYS-REQ-002 specifies the depression cascade. Coordinated supply-extract modulation (IFC-REQ-005) is essential to prevent over-pressurisation of low-activity zones when extract demand changes. | 2026-03-19T11:11:06.358Z
SYS-REQ-002 | SUB-REQ-009 | derives | Cascade maintenance requires fan condition monitoring | SYS-REQ-002 requires continuous cascade. SUB-REQ-009 provides predictive maintenance capability that prevents unplanned fan failure — a precondition for reliable cascade maintenance. | 2026-03-19T11:10:43.985Z
SYS-REQ-002 | SUB-REQ-002 | derives | Depression cascade maintenance requires fan changeover within time limit | SYS-REQ-002 specifies continuous depression cascade. SUB-REQ-002 specifies the maximum changeover time to standby fan that prevents cascade collapse during single fan failure — directly derived from the need to maintain the cascade continuously. | 2026-03-19T11:10:42.542Z
IFC-REQ-006 | VER-REQ-005 | verifies | Emergency override independence test for IFC-REQ-006 | PLC-failed-state test demonstrates hardwired safety relay independence required by IFC-REQ-006. | 2026-03-19T11:09:20.875Z
IFC-REQ-004 | VER-REQ-004 | verifies | Sampling point inspection for IFC-REQ-004 | Physical measurement confirms ISO 2889 straight-run distances for the isokinetic sampling point. | 2026-03-19T11:09:20.297Z
IFC-REQ-003 | VER-REQ-003 | verifies | Manifold commissioning test for IFC-REQ-003 | Full-load flow measurement confirms manifold capacity and static pressure specified in IFC-REQ-003. | 2026-03-19T11:09:19.660Z
IFC-REQ-002 | VER-REQ-002 | verifies | Integration test for cascade-fan control loop | Step response testing validates the control loop timing specified in IFC-REQ-002. | 2026-03-19T11:09:11.952Z
IFC-REQ-001 | VER-REQ-001 | verifies | Ductwork inspection and pressure test for IFC-REQ-001 | Inspection of material certificates and pressure test confirm ductwork material, integrity, and fan-downstream arrangement specified in IFC-REQ-001. | 2026-03-19T11:09:09.024Z
SYS-REQ-002 | IFC-REQ-002 | derives | Depression cascade derives fan control interface | Maintaining the depression cascade (SYS-REQ-002) requires the cascade controller to modulate extract fan speed; IFC-REQ-002 defines the signal interface that enables this control. | 2026-03-19T11:09:05.439Z
SYS-REQ-008 | SUB-REQ-008 | derives | Twin-bank HEPA requires DP monitoring for loading management | SYS-REQ-008 mandates twin-bank HEPA. SUB-REQ-008 specifies the DP monitoring needed to manage filter loading and trigger replacement before flow degradation compromises the cascade — a necessary operational support requirement for maintaining twin-bank filtration effectiveness. | 2026-03-19T11:08:57.196Z
SYS-REQ-008 | SUB-REQ-003 | derives | Twin-bank HEPA system requirement derives filter efficiency specification | SYS-REQ-008 mandates twin-bank HEPA on C3/C4 extract paths. SUB-REQ-003 specifies the per-bank decontamination factor and safe-change housing requirement that implements this system-level mandate at component level. | 2026-03-19T11:08:53.352Z
SYS-REQ-002 | SUB-REQ-006 | derives | Depression cascade requires adequate face velocities at containment devices | The cascade in SYS-REQ-002 depends on controlled air ingress at each containment device. SUB-REQ-006 specifies the minimum face velocity at fume cupboards and glove box pressure that collectively enforce the cascade direction. | 2026-03-19T11:08:49.871Z
SYS-REQ-002 | SUB-REQ-004 | derives | Depression cascade setpoints derive cascade control response requirement | SYS-REQ-002 specifies the pressure differentials; SUB-REQ-004 specifies how quickly the control system must restore them after perturbation. The response time is directly derived from the need to maintain the cascade under operational disturbances. | 2026-03-19T11:08:46.748Z
STK-REQ-001 | SYS-REQ-010 | derives | Coolant analysis turnaround drives LSC tritium detection requirements | H-3 is listed as a safety-critical coolant parameter in STK-REQ-001. The counting time constraint derives directly from the 4-hour turnaround requirement, given sample preparation takes approximately 1-2 hours. | 2026-03-19T10:47:23.052Z
STK-REQ-002 | SYS-REQ-009 | derives | Nuclear site licence LC24 (fissile material) drives criticality safety requirements | LC24 explicitly requires arrangements to prevent criticality. The laboratory handles fissile material in solution form during dissolution and separation, making criticality safety controls a direct licence condition compliance requirement. | 2026-03-19T10:47:21.828Z
STK-REQ-004 | SYS-REQ-008 | derives | ALARP for inhalation dose drives HEPA filtration requirements | HEPA filtration prevents airborne contamination from reaching the stack and environment, directly supporting ALARP for both on-site workers and the public. | 2026-03-19T10:47:20.609Z
STK-REQ-003 | SYS-REQ-007 | derives | Discharge permit compliance for specific nuclides drives ICP-MS Tc-99 detection capability | EA discharge authorisations list Tc-99 specifically. ICP-MS provides the most practical route to achieving regulatory detection limits for this nuclide without lengthy radiochemical separation. | 2026-03-19T10:46:50.409Z
STK-REQ-004 | SYS-REQ-006 | derives | ALARP dose management drives airborne contamination detection requirements | The dominant inhalation dose pathway from alpha-active materials requires rapid detection and alarming to minimise personnel intake during contamination events, directly supporting the ALARP dose constraint. | 2026-03-19T10:46:49.140Z
STK-REQ-002 | SYS-REQ-005 | derives | Nuclear site licence compliance drives LIMS audit trail requirements | ONR LC25 (operational records) and LC14 (safety documentation) require demonstrable record-keeping. The LIMS audit trail is the primary mechanism for satisfying these conditions for analytical data. | 2026-03-19T10:46:47.845Z
STK-REQ-005 | SYS-REQ-004 | derives | Nuclear material assay accuracy drives alpha spectrometry detection and resolution requirements | IAEA International Target Values for Pu determination require low detection limits and adequate spectral resolution to provide reliable isotopic ratios for material accountancy. | 2026-03-19T10:46:46.597Z
STK-REQ-003 | SYS-REQ-003 | derives | Discharge permit compliance drives batch-level effluent monitoring | The EA/SEPA permit requires demonstrable monitoring before discharge. Batch-by-batch analysis with recorded results provides the verifiable evidence chain required by the permit conditions. | 2026-03-19T10:46:09.576Z
STK-REQ-004 | SYS-REQ-002 | derives | ALARP dose constraint drives depression cascade specification | Maintaining negative pressure prevents airborne contamination spreading to occupied areas, directly reducing inhalation dose which is the dominant exposure pathway for alpha-active material handling. | 2026-03-19T10:46:08.348Z
STK-REQ-001 | SYS-REQ-001 | derives | Coolant analysis turnaround drives gamma spec sensitivity to enable 1-hour counts | The 4-hour turnaround for coolant safety parameters requires that gamma spectrometry achieve adequate sensitivity in short counting times, directly constraining detector efficiency and MDA. | 2026-03-19T10:46:07.089Z

## Linksets
Source Document | Target Document | Default Link Type | Link Count
subsystem-requirements | verification-requirements |  | 8
interface-requirements | verification-requirements |  | 50
system-requirements | verification-requirements |  | 9
system-requirements | subsystem-requirements |  | 102
system-requirements | interface-requirements |  | 10
stakeholder-requirements | system-requirements |  | 13

## Sections
Document | Section | Short Code | Order | Requirements
architecture-decisions | Architecture Decisions | ARC | 0 | 14
interface-requirements | Interface Requirements | IFC | 0 | 55
stakeholder-requirements | Stakeholder Requirements | STK | 0 | 7
subsystem-requirements | Subsystem Requirements | SUB | 0 | 100
system-requirements | System Requirements | SYS | 0 | 12
verification-requirements | Verification Requirements | VER | 0 | 62