## Documents
Slug | Name | Short Code | Sections | Requirements | Created
architecture-decisions | Architecture Decisions | ARC | 1 | 15 | 2026-03-19T15:02:28.870Z
interface-requirements | Interface Requirements | IFC | 1 | 48 | 2026-03-19T15:02:28.395Z
stakeholder-requirements | Stakeholder Requirements | STK | 1 | 12 | 2026-03-19T15:02:26.777Z
subsystem-requirements | Subsystem Requirements | SUB | 1 | 162 | 2026-03-19T15:02:27.888Z
system-requirements | System Requirements | SYS | 1 | 19 | 2026-03-19T15:02:27.316Z
verification-plan | Verification Plan | VER | 1 | 193 | 2026-03-19T15:02:29.379Z

## Requirements
Ref | Text | Document | Section | Pattern | Verification | QA Score | Compliance | Rationale | Tags
ARC-MAIN-001 | ARC: Safety and Interlock Subsystem — dedicated safety processor with hardware watchdog authority. The SIS runs on a processor physically separated from motion control compute with independent power and CAN-bus brake authority. Alternatives considered: software safety monitor on shared CPU (rejected — common-cause fault kills both monitor and controlled system, violates SIL 3 HFT=1 requirement); safety PLC (rejected — latency budget incompatible with 1kHz control loop). Chosen architecture satisfies IEC 61508 SIL 3 Hardware Fault Tolerance = 1 with <1% common-cause exposure. | architecture-decisions | Main |  | Inspection |  |  | Architectural decision: Safety-Integrity Level 3 requires hardware fault tolerance that physically separate execution cannot provide on shared compute. This decision constrains all downstream SIS design. | informational
ARC-MAIN-002 | ARC: Motion Control System — Cartesian-space pipeline with independent tremor filter and scaling stages. Cartesian-space scaling chosen over joint-space to preserve instrument orientation during scaled motion; joint-space scaling produces unexpected end-effector trajectories under redundancy resolution. Tremor filter precedes scaling to avoid amplifying filtered residuals. Six-component pipeline (Filter → Scale → IK → Safety → Servo) chosen for independent testability of each function, enabling SIL 3 V&V decomposition. Redundant compute node mandated by single-point-of-failure analysis: a compute failure in motion control must degrade to safe-stop, not uncontrolled motion. | architecture-decisions | Main |  | Inspection |  |  | Design rationale for the most safety-critical subsystem: captures key architectural choices that constrain all downstream component specifications and must be preserved across future refactors. | informational
ARC-MAIN-003 | ARC: Vision and Imaging System — FPGA-based deterministic image processing pipeline with separated optical acquisition and display stages. The CCU performs per-channel sensor correction (white balance, gamma, chromatic aberration) before the FPGA pipeline adds enhancement and overlay compositing. FPGA was chosen over GPU because IEC 62304 Class C certification requires deterministic worst-case latency guarantees that GPU scheduling cannot provide. The stereo endoscope uses paired CMOS sensors with fixed 6mm baseline rather than a beam-splitter design because paired sensors allow independent channel failure (graceful degradation to 2D) and simpler sterilisation of the rigid optical assembly. Illumination uses closed-loop intensity feedback from the CCU to prevent tissue thermal damage, rather than fixed-intensity operation, because tissue reflectance varies 4x between organ types. | architecture-decisions | Main |  | Analysis |  |  | Architecture driven by three constraints: IEC 62304 Class C certification demands deterministic processing latency; graceful degradation from stereo to mono requires independent channel paths; IEC 60601-2-18 tissue temperature limits require closed-loop illumination control. | informational
ARC-MAIN-004 | ARC: Haptic Feedback Subsystem — Galvanic isolation at Force Signal Conditioner boundary. The Force Signal Conditioner introduces a 4kVrms isolation barrier rather than isolating at the instrument tip. This choice keeps the patient-contact chain entirely analog (passive strain gauge) while confining the isolation requirement to a single, testable PCB boundary. Alternative (digital sensor in handle) was rejected due to sterilisation constraints on active electronics in the wristed instrument channel. | architecture-decisions | Main |  | Inspection |  |  | Architecture decisions are narrative records of design trade-offs; the rationale is embedded in the requirement text. The isolation placement choice (at FSC boundary rather than instrument tip) is driven by sterilisation constraints and regulatory compliance with IEC 60601-1 patient leakage current limits. | informational
ARC-MAIN-005 | ARC: Communication and Data Management System — FPGA-based real-time protocol rather than OS-scheduled networking. A custom FPGA protocol engine rather than a Linux kernel TCP/IP stack is used for kinematic command framing because TCP cannot guarantee sub-millisecond per-frame latency under video traffic coexistence. The FPGA provides hardware-guaranteed timing isolation between priority channels, eliminating head-of-line blocking. A proprietary TDM protocol over 10GbE is preferred over TSN (802.1Qbv) because TSN requires switches to participate in time synchronisation, adding network device dependencies outside the system boundary. | architecture-decisions | Main |  | Inspection |  |  | Architecture decisions are narrative records of design trade-offs; the rationale is embedded in the requirement text. FPGA-based TDM over 10GbE was chosen over Linux TCP/IP and IEEE 802.1Qbv TSN to avoid OS scheduling jitter and external network device dependencies in the deterministic control path. | informational
ARC-MAIN-006 | ARC: Surgical Instrument System uses cable-driven actuation with remote motors in the Instrument Drive Unit rather than direct-drive or gear-driven end-effector actuation. Alternatives considered: direct-drive motors at instrument tip (rejected: tip diameter constraint of 8mm precludes motors with sufficient torque; sterilization incompatible with embedded electronics); gear-driven actuation through instrument shaft (rejected: gear backlash exceeds 0.1mm accuracy requirement; maintenance burden of gear wear in disposable instruments is cost-prohibitive). Cable-driven approach enables all motors to reside in the non-sterile Instrument Drive Unit, separated from the sterile instrument by the Sterile Adapter. The Bouc-Wen hysteresis compensation in the Tool Tip Articulation Controller addresses the primary disadvantage of cable drive — nonlinear cable friction. | architecture-decisions | Main |  | Inspection |  |  | The cable-driven architecture is the only approach that simultaneously satisfies the 8mm instrument diameter constraint, the 0.1mm tip accuracy requirement, the sterile barrier architecture, and the disposable instrument cost target. This decision constrains all downstream instrument mechanical and control design. | informational
ARC-MAIN-007 | ARC: Motion Control System — Dedicated Trajectory Generator between Motion Scaling and Kinematics Engine. The pipeline was extended to include an explicit trajectory generation stage that performs S-curve velocity profiling and acceleration limiting. This was chosen over embedding trajectory generation within the Kinematics Engine because: (1) separation of Cartesian-space trajectory planning from joint-space kinematics enables independent testing and tuning of motion smoothness constraints, (2) the Trajectory Generator can enforce tissue-safe acceleration limits before the kinematic solution, preventing the Kinematics Engine from ever receiving a demand that would produce unsafe tip forces, (3) the lock-free SPSC ring buffer interface provides timing isolation — a transient overrun in trajectory computation does not stall the kinematics cycle. | architecture-decisions | Main |  | Inspection |  |  | Separating trajectory generation from inverse kinematics enables independent testing of motion smoothness and prevents unsafe demands reaching the Kinematics Engine. SPSC ring buffer provides timing isolation between stages. | architecture, motion-control, session-348
ARC-MAIN-008 | ARC: Power Management Subsystem — Dedicated auxiliary 24V supply for safety circuits, electrically isolated from main bus. Considered single-supply architecture with software-controlled priority; rejected because any main bus controller firmware fault could de-energise safety contactors. Hardware isolation ensures safety supervision remains active regardless of software state. | architecture-decisions | Main |  | Inspection |  |  | Hardware isolation of safety circuits from main bus ensures safety supervision cannot be defeated by a firmware fault in the main bus controller. Single-supply with software priority was rejected because software-controlled switching cannot provide the necessary independence for IEC 62304 Class C safety circuitry. | architecture, power-management, session-350
ARC-MAIN-009 | ARC: Energy Delivery System — Dual-modality architecture (RF electrosurgery + 55.5kHz ultrasonic) with centralised Energy Delivery Controller. RF chosen for vessel sealing and haemostasis (monopolar/bipolar); ultrasonic chosen for structures within 1mm of critical vessels due to lower thermal spread. Separate Return Electrode Monitor mandated by IEC 60601-2-2 for monopolar monopolar patient safety; Tissue Effect Monitor added to detect vessel seal endpoints automatically, reducing surgeon reliance on subjective colour assessment. Single controller enforces mutual exclusion — simultaneous RF and ultrasonic activation is prohibited by interlock, not by convention. | architecture-decisions | Main |  | Inspection |  |  | Mixed-modality energy delivery is standard in advanced robotic surgery (cf. da Vinci ESS, Medtronic Thunderbeat), where no single energy type covers all tissue types. Centralised controller enforces safety envelope; distributed generators maintain isolation from each other and from the robot control network. | architecture, energy-delivery, session-352
ARC-MAIN-012 | ARC: Surgeon Input Console — Physical integration of haptic master arms, stereo viewer, foot pedals, and touchscreen into one ergonomic station, with strict separation of safety-critical hardware (foot pedal E-stop, hardwired) from software-mediated controls (touchscreen, voice). Rationale: decoupling safety-path hardware (E-stop pedal → Emergency Stop Chain) from the software configuration interface (touchscreen → Console Computer) ensures a firmware or OS fault on the Console Computer cannot suppress a surgeon-initiated emergency stop. Alternative of routing all pedal inputs through the Console Computer was rejected for this reason. | architecture-decisions | Main |  | Inspection |  |  | Architectural safety requirement: the IEC 62304 and IEC 80601-2-77 framework requires hardware independence between safety-critical inputs (E-stop) and non-safety software systems. This decision is traceable to STK-MAIN-002 (no uncontrolled energy or motion at any time). | architecture, surgeon-console, safety, session-353
ARC-MAIN-014 | ARC: Haptic Feedback Subsystem — galvanic isolation break between patient-side sensing and surgeon-side actuation. The signal chain is split into Force Sensing Module (sterile field) → Force Signal Conditioner (galvanic isolation, ≥4kVrms) → Haptic Controller (console-side, SIL2) → Master Handle Actuator. Alternatives considered: (a) direct digital fibre link eliminating analogue conditioning — rejected due to sensor power budget constraints in the sterile field; (b) single integrated controller spanning both sides — rejected because it creates a conductive path violating IEC 60601-1 patient leakage current limits. The 4kVrms isolation barrier is the single safety-critical boundary in the haptic chain. | architecture-decisions | Main |  | Inspection |  |  | IEC 60601-1 requires patient leakage current < 10µA in CF-type applied parts; a conductive signal path spanning the sterile and console sides would create unacceptable leakage current. The 4kVrms galvanic isolation barrier physically prevents this path. Verification by design review of isolation barrier certification documentation. | architecture, haptic, session-354
ARC-MAIN-015 | ARC: Motion Control System — FPGA-based infrastructure layer (Real-Time Protocol Engine, Inter-Cart Fibre Link, Network Management Controller, Procedure Data Recorder) is separated from the algorithmic layer (Tremor Filter, Motion Scaling, Trajectory Generator, Kinematics Engine, Workspace Safety Enforcer, Joint Servo Controller) because the two layers have fundamentally different failure modes and verification requirements. Infrastructure components are hardware-fixed timing elements verified by oscilloscope measurement; algorithmic components are software-configurable and verified by functional test. This separation also allows the FPGA layer to enforce timing guarantees independently of software faults — if the algorithmic layer hangs, the FPGA continues transmitting watchdog frames to trigger a controlled stop. | architecture-decisions | Main |  | Inspection |  |  | The two-layer decomposition separates hardware-fixed timing from software-configurable algorithms. This allows the FPGA infrastructure layer to enforce timing guarantees independently of software faults: if the algorithmic layer hangs, the FPGA continues transmitting watchdog frames to the patient cart, triggering a controlled stop. Infrastructure components are verified by oscilloscope measurement; algorithmic components are verified by functional test. This separation is consistent with IEC 60601-1 risk management requirements for partitioning safety-critical timing from application software. | architecture, motion-control, infrastructure, session-357
ARC-MAIN-018 | ARC: Motion Control and Scaling Subsystem — linear pipeline topology with safety injection. Components are ordered as: Tremor Rejection Filter → Motion Scaling Module → Trajectory Generator → Kinematics Engine → Joint Servo Controller, with the Workspace Safety Enforcer injecting repulsive constraints directly into the Kinematics Engine rather than acting as a separate gate. The Real-Time Compute Node hosts all pipeline stages on a single RTOS. Alternative considered: distributed pipeline across multiple CPUs for fault isolation. Rejected because cross-CPU IPC adds 0.2-0.4ms per stage, making the 1ms end-to-end budget unachievable. The single-node topology is acceptable because the Real-Time Compute Node is itself triple-redundant (2-of-3 voting) at the hardware level. | architecture-decisions | Main |  | Inspection |  |  | Documents the primary architectural trade-off driving the MC subsystem topology: determinism over distributed fault isolation. The 1ms end-to-end latency in SYS-MAIN-007 is the binding constraint; this decision records why distributed alternatives were rejected and what compensating mechanism (hardware triple-redundancy) provides the required fault tolerance. | architecture, motion-control, session-358
ARC-MAIN-019 | ARC: Verification Plan Coverage — Verification entries cover all safety-critical IFC requirements first (IFC-001 through IFC-013, IFC-020), prioritised by SIL classification under IEC 62061. Interface verification is selected as the minimum viable verification set because every cross-subsystem interface is a potential failure propagation path in a safety-critical medical robot. Performance threshold values (2ms joint force detection, 150ms E-stop de-energisation, 100ms teleoperation latency, 15ms haptic render) are derived from published clinical evidence and applicable standards (IEC 62061, IEC 60601-1), not estimated. | architecture-decisions | Main |  | Inspection |  |  | This architectural decision records the rationale for prioritising safety-critical IFC requirements in the verification plan. The coverage set (IFC-001 to IFC-013, IFC-020) maps directly to cross-subsystem interfaces with SIL 3 classification under IEC 62061; performance thresholds are derived from IEC 60601-1 and published surgical robotics clinical evidence, ensuring the verification plan is traceable to recognised standards rather than arbitrary estimates. | architecture, verification, session-360
ARC-MAIN-020 | ARC: Communication and Data Management System — dual-redundant fibre topology with FPGA-level failover and strict-priority traffic shaping. The architecture separates the real-time kinematic command path (RTPE on compute node) from the link management plane (NMC in FPGA) so that link health monitoring and failover do not share any compute resources with the 1kHz control loop. LVDS hardwired fault notification to SIS ensures the communication system can report its own failure without relying on the link it monitors. PDR uses DMA shared memory to record data without touching the RTPE thread. Alternative considered: single-port fibre with software failover — rejected because software-layer switchover could not meet the 5ms deadline under worst-case OS scheduling latency. | architecture-decisions | Main |  | Inspection |  |  | Architecture decision captures why the RTPE/NMC separation was chosen over integrated approaches, ensuring future maintainers understand the SIL 3 isolation rationale. | architecture, comms, session-369, idempotency:arc-comms-369
IFC-MAIN-001 | The interface between the Joint Force Monitor and the Motion Control and Scaling Subsystem SHALL carry per-axis joint torque readings at 1kHz over a dedicated real-time bus with <200µs latency, using hardware checksums for data integrity. | interface-requirements | Main |  | Test |  |  | 1kHz at <200µs is required for the Joint Force Monitor to execute its 50ms graduated brake response with sufficient measurement cycles (50 samples) to avoid false triggers. Hardware checksum prevents corrupted torque values from masking force limit exceedances. | interface, sis, mcs, joint-force, session-341, idempotency:ifc-jfm-mcs-torque-001-341
IFC-MAIN-002 | The interface between the Emergency Stop Chain and the Power Management Subsystem servo drive contactors SHALL be a hardwired 24V DC control loop; interruption of the loop SHALL cause contactor drop-out within 50ms by capacitor discharge drive. | interface-requirements | Main |  | Test |  |  | Hardwired loop (not digital signal) ensures contactor drop-out survives software fault or bus failure. 50ms is achievable with the capacitor hold-in drive approach without an external power supply. | interface, sis, pms, estop, hardware, session-341, idempotency:ifc-estop-pms-contactor-001-341
IFC-MAIN-003 | The interface between the Communication Monitor and the inter-cart fibre link SHALL expose per-frame CRC pass/fail, round-trip latency in microseconds, and receive buffer occupancy at 1kHz to the Communication Monitor via a sideband status register, without interrupting the data path. | interface-requirements | Main |  | Test |  |  | Sideband status register (not in-band signalling) prevents monitoring from interfering with the real-time control stream. 1kHz status updates are necessary to meet the 3-frame loss detection latency requirement in SUB-MAIN-002. | interface, sis, communication, fibre, session-341, idempotency:ifc-comms-monitor-fibre-001-341
IFC-MAIN-004 | The Safe State Manager SHALL broadcast the current system safety state (OPERATIONAL, DEGRADED, SAFE-HOLD) to all subsystems over a dedicated safety bus within 5ms of any state transition, using a publisher-subscriber protocol with guaranteed delivery. | interface-requirements | Main |  | Test |  |  | 5ms broadcast latency ensures all subsystems receive the new state within the first 2% of the 250ms safe-state window, allowing coordinated response without race conditions between independently-responding subsystems. | interface, sis, safe-state, broadcast, session-341, idempotency:ifc-ssm-broadcast-001-341
IFC-MAIN-005 | The interface between the Surgeon Console and the Motion Scaling Module SHALL transmit Cartesian velocity commands for all 6 DOF at 1kHz using a UDP/multicast protocol over 1GbE fibre with maximum one-way latency of 3ms and packet loss tolerance of 0 (zero packet drop acceptable; retransmit on loss). | interface-requirements | Main |  | Test |  |  | 1kHz command rate is the Nyquist minimum for 500Hz servo bandwidth; 3ms one-way latency is the network share of the 100ms system budget. Zero packet loss is mandatory: a dropped command packet causes a missed servo cycle and potential jerk at the instrument tip. | interface, console-mc, session-340
IFC-MAIN-006 | The interface between the Motion Control System and the Patient-Side Cart SHALL transmit 21 joint-angle setpoints (7 per arm x 3 arms) at 1kHz via CAN-FD at 5Mbps with maximum command-to-actuation latency of 2ms and shall include a CRC on every frame. | interface-requirements | Main |  | Test |  |  | CAN-FD at 5Mbps provides sufficient bandwidth for 21 x 2-byte joint setpoints (42 bytes) at 1kHz = 336kbps, well within 5Mbps limit. 2ms latency includes transmission and joint servo processing. CRC mandated because a corrupted setpoint could command a joint to an incorrect position, creating patient risk. | interface, mc-cart, session-340
IFC-MAIN-007 | The interface between the Motion Control System and the Safety and Watchdog System SHALL provide a dedicated hardwired heartbeat signal at 200Hz; loss of heartbeat for more than 3 consecutive pulses (15ms) SHALL be interpreted by the Safety System as a Motion Control fault requiring immediate brake engagement. | interface-requirements | Main |  | Test |  |  | Hardwired heartbeat is used (not software message) because a software crash could prevent sending a fault notification; hardware signal remains functional even when the compute node OS has hung. 15ms timeout gives 35ms margin within the 50ms emergency stop budget. | interface, mc-safety, session-340
IFC-MAIN-008 | The interface between the Vision and Imaging System and the Surgeon Console SHALL transmit two independent uncompressed 1080p/60Hz video streams over 10GbE, synchronised within 1ms between left and right channels, with end-to-end latency below 50ms. | interface-requirements | Main |  | Test |  |  | Left-right synchronisation within 1ms is required to prevent vergence-accommodation conflict in the stereoscopic display; mismatches above 5ms cause eye strain and nausea in surgeons. 50ms end-to-end is the video share of the total latency budget. | interface, vision-console, session-340
IFC-MAIN-009 | The interface between the Surgical Instrument System and the Motion Control System SHALL transmit instrument type identifier, usage count, and 3-axis tip force measurements at 500Hz via CAN-FD; instrument type SHALL be validated within 500ms of instrument insertion before motion commands are accepted. | interface-requirements | Main |  | Test |  |  | Instrument type is required by the Kinematics Engine to load correct DH parameters and mass properties; wrong kinematics model creates joint-space errors. Force measurements at 500Hz provide the Motion Control pipeline with tissue-contact data for the Workspace Safety Enforcer force limits. | interface, instrument-mc, session-340
IFC-MAIN-010 | The interface between the Stereo Endoscope and the Camera Control Unit SHALL carry two independent HD-SDI video channels at 1.485 Gbps each, transmitting raw Bayer-pattern sensor data at 1920x1080 resolution and 60 frames per second, with BER not exceeding 1e-12. | interface-requirements | Main |  | Test |  |  | Dual independent HD-SDI channels ensure single-channel failure does not affect the other (graceful degradation to 2D). Raw Bayer-pattern data is required because demosaicing and colour correction must be performed in the CCU where calibration data is stored. BER of 1e-12 prevents visible artifacts in the surgical image that could be mistaken for tissue features. | interface, vision, session-341, idempotency:ifc-endoscope-ccu-341
IFC-MAIN-011 | The interface between the Camera Control Unit and the Image Processing Pipeline SHALL transmit two synchronised 3G-SDI video streams at 2.97 Gbps each, carrying colour-corrected 1080p60 10-bit 4:2:2 video with inter-channel temporal skew not exceeding 100 microseconds. | interface-requirements | Main |  | Test |  |  | 3G-SDI at 2.97 Gbps provides sufficient bandwidth for 10-bit 4:2:2 colour depth which preserves tissue colour fidelity for narrow-band imaging modes. The 100us inter-channel skew budget at this interface is tighter than the 500us end-to-end stereo sync requirement because downstream processing adds additional jitter. | interface, vision, session-341, idempotency:ifc-ccu-ipp-341
IFC-MAIN-012 | The interface between the Camera Control Unit and the Surgical Illumination Source SHALL transmit exposure metering data at 60Hz via RS-422 serial link at 115200 baud, with command-to-intensity-change latency not exceeding 16ms (one frame period). | interface-requirements | Main |  | Test |  |  | Closed-loop illumination control requires exposure feedback at frame rate to prevent tissue overheating during endoscope repositioning (tissue reflectance changes abruptly as the tip moves between organs). RS-422 provides differential signalling for noise immunity in the OR electromagnetic environment. 16ms response ensures intensity tracks scene changes within one frame. | interface, vision, session-341, idempotency:ifc-ccu-illumination-341
IFC-MAIN-013 | The interface between the Image Processing Pipeline and the Stereoscopic Display System SHALL transmit two independent DisplayPort 1.2 video streams at 3840x2160 resolution and 60Hz with 10-bit colour depth, with maximum end-to-end latency from IPP output to photon emission not exceeding 8ms. | interface-requirements | Main |  | Test |  |  | 4K resolution per eye at 10-bit depth provides tissue colour fidelity required for narrow-band imaging differential diagnosis. The 8ms photon emission budget includes display controller processing and LCD response time. This budget is derived from the total 50ms hand-to-eye latency budget minus upstream processing and motion control allocations. | interface, vision, session-341, idempotency:ifc-ipp-display-341
IFC-MAIN-014 | The interface between the Image Processing Pipeline and the Procedure Video Recorder SHALL provide a composited 2D 1080p60 video stream via 3G-SDI at 2.97 Gbps, with embedded audio channels for OR ambient audio capture, and a parallel Ethernet link carrying timestamped system event metadata at 1kHz. | interface-requirements | Main |  | Test |  |  | Composited 2D output combines left/right channels with overlay annotations for a single-stream recording suitable for review and teaching. Embedded audio captures surgeon voice notes. The parallel metadata link at 1kHz enables frame-accurate correlation between video and kinematic data per SYS-MAIN-015, which is critical for post-operative complication analysis. | interface, vision, session-341, idempotency:ifc-ipp-recorder-341
IFC-MAIN-015 | The interface between the Force Sensing Module and the Force Signal Conditioner SHALL carry six-axis strain gauge bridge differential signals with a common-mode rejection ratio of at least 80dB, connecting via a 12-conductor shielded cable with maximum length of 0.5m. | interface-requirements | Main |  | Test |  |  | Strain gauge bridge output is a low-level differential signal (typically 1-10mV full scale). 80dB CMRR is required to reject the >10V common-mode noise present in the OR environment from surgical energy generators and motor drive switching. Cable length limit prevents impedance mismatch that degrades CMRR at high frequencies. | rt-vague-interface, red-team-session-502
IFC-MAIN-016 | The interface between the Force Signal Conditioner and the Haptic Controller SHALL transmit six-axis 16-bit digitised force samples at 1kHz via an isolated SPI bus operating at 10MHz, with a maximum bus propagation latency of 100us. | interface-requirements | Main |  | Test |  |  | 16-bit resolution at 10MHz SPI provides sufficient bandwidth for six channels at 1kHz with headroom for protocol overhead. 100us propagation latency budget ensures the full haptic loop (sense-transmit-compute-actuate) stays within the 2ms latency requirement allocated in SUB-MAIN-023. | interface, haptic, digital, session-342, idempotency:ifc-fsc-hc-342
IFC-MAIN-017 | The interface between the Haptic Controller and the Master Handle Actuator motor driver SHALL transmit per-joint torque setpoints at 1kHz via CAN FD at 5Mbit/s, with hardware-enforced torque limiting such that no frame can command a torque greater than 1.2Nm on any joint. | interface-requirements | Main |  | Test |  |  | CAN FD at 5Mbit/s provides sufficient bandwidth for 7-DOF torque commands at 1kHz with message latency under 100us. Hardware torque limiting on the motor driver provides a second layer of force protection independent of the Haptic Controller software, meeting the safety-in-depth requirement for the 1N feedback limit. | interface, haptic, actuator, session-342, idempotency:ifc-hc-mha-342
IFC-MAIN-018 | The interface between the Real-Time Protocol Engine and the Inter-Cart Fibre Link SHALL operate at 10Gbit/s with a maximum per-frame latency of 200us for kinematic command frames, and SHALL support a minimum of 8 logical channels multiplexed onto one physical fibre wavelength. | interface-requirements | Main |  | Test |  |  | 10Gbit/s capacity accommodates 21-joint kinematics (21x64-byte frames at 1kHz = 10.7Mbit/s), two 4K60 video streams (~4Gbit/s combined), and headroom for safety and logging traffic. 200us per-frame latency contributes less than 20% of the 1ms communications budget allocated in SUB-MAIN-027. | interface, comms, fibre, session-342, idempotency:ifc-rpe-fibre-342
IFC-MAIN-019 | The interface between the Network Management Controller and the Safety and Interlock Subsystem SHALL provide a unidirectional status register update at 100Hz, reporting optical power level (dBm), frame loss rate (frames/s), and active link identity (primary or standby), readable by the Communication Monitor component. | interface-requirements | Main |  | Test |  |  | 100Hz polling rate ensures the Communication Monitor (operating at 1kHz internal rate) has a link health update no older than 10ms, consistent with the 5ms fault detection target in SUB-MAIN-028. Unidirectional (read-only) access from Safety subsystem preserves safety isolation — the SIS can observe but not command the network. | interface, comms, safety, session-342, idempotency:ifc-nmc-sis-342
IFC-MAIN-020 | The interface between the Instrument Recognition Module and the Tool Tip Articulation Controller SHALL transfer instrument kinematic model parameters (cable routing geometry, pulley ratios, coupling compliance coefficients, and Bouc-Wen hysteresis model parameters) as a structured data packet of no more than 2KB via the internal CAN-FD bus within 50ms of instrument identity validation. | interface-requirements | Main | event | Test |  |  | The Tool Tip Articulation Controller requires instrument-specific kinematic parameters to compute accurate cable displacement commands. The 2KB limit is derived from the maximum parameter set size for a 6-DoF instrument with per-cable hysteresis model (24 floats per cable x 6 cables + metadata). The 50ms transfer time is a sub-budget of the 200ms total recognition time in SUB-MAIN-032. | surgical-instrument-system, interface, session-346
IFC-MAIN-021 | The interface between the Cable Tensioning System and the Safety and Interlock Subsystem SHALL transmit cable tension anomaly alerts as a priority CAN-FD frame containing affected arm ID, cable channel, measured tension, nominal tension, and timestamp, with a maximum end-to-end latency of 2ms from anomaly detection to Safety and Interlock Subsystem receipt. | interface-requirements | Main | event | Test |  |  | The Safety and Interlock Subsystem must receive cable anomaly data fast enough to enforce the 50ms arm shutdown in SUB-MAIN-038. The 2ms interface latency consumes only 4% of the total budget, leaving adequate margin for safety processing and actuator command. Priority CAN-FD framing ensures the alert is not delayed by regular control traffic. | surgical-instrument-system, interface, safety, session-346
IFC-MAIN-022 | The interface between the Tool Tip Articulation Controller and the Instrument Drive Unit SHALL deliver cable displacement commands for all four instrument DoF as a single CAN-FD frame at 1kHz, with each command expressed as a 16-bit signed integer representing micrometers of cable displacement, achieving command-to-actuation latency of no more than 200 microseconds. | interface-requirements | Main | ubiquitous | Test |  |  | The 1kHz update rate matches the motion control servo loop. 16-bit signed integer resolution provides +/-32mm range at 1um precision, sufficient for the cable displacement range of +/-20mm. The 200us command-to-actuation latency is the Instrument Drive Units share of the 1ms servo cycle, after the Tool Tip Articulation Controllers 500us computation budget. | surgical-instrument-system, interface, session-346
IFC-MAIN-023 | The interface between the Instrument Lifecycle Controller and the Safe State Manager SHALL transmit instrument lockout commands as a CAN-FD frame containing arm ID and lockout reason code, with the Safe State Manager acknowledging receipt within 5ms and inhibiting arm enable until a valid instrument is coupled. | interface-requirements | Main | event | Demonstration |  |  | The Safe State Manager is the single authority for arm enable per the safety architecture (ARC-MAIN-001). Instrument lockout must route through the safety chain rather than being enforced locally by the Instrument Lifecycle Controller, because bypassing the safety processor would violate the SIL 3 architecture. The 5ms acknowledgement ensures the lockout is enforced before the surgeon can begin using the arm. | surgical-instrument-system, interface, safety, session-346
IFC-MAIN-024 | The interface between the Tremor Rejection Filter and the Motion Scaling Module SHALL carry filtered 6-DOF Cartesian velocity vectors (3 translational, 3 rotational) at 1kHz as 64-bit IEEE 754 floating-point values over shared memory, with worst-case read latency below 5 microseconds and a sequence counter for stale-data detection. | interface-requirements | Main |  | Test |  |  | Shared memory is mandated by the sub-millisecond pipeline budget — IPC mechanisms such as sockets or message queues would introduce unacceptable jitter. The 64-bit float format preserves the 15-digit precision needed for sub-100μm tip positioning. The sequence counter enables the Motion Scaling Module to detect and reject stale frames, which is the primary defence against pipeline stall propagation. | interface, motion-control, session-348, idempotency:ifc-trf-msm-velocity-348
IFC-MAIN-025 | The interface between the Motion Scaling Module and the Trajectory Generator SHALL transmit scaled 6-DOF Cartesian velocity commands at 1kHz via shared memory, including the active scaling ratio as a metadata field, with data validity indicated by a monotonically incrementing timestamp synchronized to the Real-Time Compute Node system clock. | interface-requirements | Main |  | Test |  |  | Including the active scaling ratio as metadata enables the Trajectory Generator to adjust acceleration limits proportionally — at 10:1 scaling, the same surgeon hand velocity produces 10× lower instrument velocity, allowing tighter acceleration bounds. The synchronized timestamp is essential for the Trajectory Generator to compute correct velocity integration and detect timing faults. | interface, motion-control, session-348, idempotency:ifc-msm-tg-scaled-velocity-348
IFC-MAIN-026 | The interface between the Trajectory Generator and the Kinematics Engine SHALL deliver interpolated Cartesian pose setpoints (position as 3-element vector in mm, orientation as unit quaternion) at 1kHz via a lock-free SPSC ring buffer, with buffer depth of at least 4 frames to tolerate scheduling jitter up to 3ms without data loss. | interface-requirements | Main |  | Test |  |  | A lock-free single-producer/single-consumer ring buffer eliminates mutex contention in the real-time pipeline. The 4-frame buffer depth provides 3ms of jitter tolerance, which exceeds the measured worst-case PREEMPT_RT scheduling jitter of 1.5ms on the target compute platform. Quaternion representation avoids gimbal lock that would occur with Euler angles during wrist-over manoeuvres common in surgical procedures. | interface, motion-control, session-348, idempotency:ifc-tg-ke-cartesian-pose-348
IFC-MAIN-027 | The interface between the Kinematics Engine and the Joint Servo Controller SHALL transmit per-joint angle setpoints for all 7 DOF of each instrument arm at 1kHz over the EtherCAT fieldbus, with each frame containing position (32-bit), velocity feedforward (32-bit), and torque feedforward (32-bit) per joint, and end-to-end frame delivery latency below 250 microseconds. | interface-requirements | Main |  | Test |  |  | EtherCAT provides deterministic sub-microsecond synchronization across all servo drives, which is mandatory for coordinated multi-joint motion. The 250μs delivery budget consumes one quarter of the 1ms control cycle, leaving 750μs for servo computation and actuation. Velocity and torque feedforward terms are essential for high-bandwidth tracking — position-only control produces unacceptable following error during fast surgical manoeuvres. | interface, motion-control, session-348, idempotency:ifc-ke-jsc-joint-setpoints-348
IFC-MAIN-028 | The interface between the Workspace Safety Enforcer and the Kinematics Engine SHALL provide a real-time workspace boundary constraint set updated at 100Hz, specifying the active Cartesian workspace limits as a convex polytope (up to 24 half-plane constraints) and per-joint angle limits, transmitted via shared memory with atomic compare-and-swap to prevent partial reads. | interface-requirements | Main |  | Test |  |  | The workspace boundary is dynamically reconfigured during surgery (e.g., when the operating table repositions or when the surgeon redefines the working volume). 100Hz update rate ensures the Kinematics Engine operates within constraints that lag the actual boundary by at most 10ms. Convex polytope representation enables efficient real-time collision checking via linear programming, which the Kinematics Engine can evaluate within its 2ms computation window. Atomic CAS prevents the Kinematics Engine from reading a partially-updated polytope, which could create a non-convex feasible region and allow unsafe motion. | interface, motion-control, session-348, idempotency:ifc-wse-ke-boundary-348
IFC-MAIN-029 | The interface between the Auxiliary Power Supply and the Emergency Stop Chain SHALL maintain contactor coil energisation voltage within 22–26V DC continuously, including during mains loss events, with no interruption exceeding 10ms. | interface-requirements | Main |  | Test |  |  | The Emergency Stop Chain contactors must remain closed during normal operation. A voltage interruption exceeding 10ms will cause contactor dropout and trigger an uncontrolled E-stop during surgery. The 10ms limit matches the contactor dropout time specification in IEC 60204-1 Category 3 systems. | rt-missing-failure-mode, red-team-session-502
IFC-MAIN-030 | The interface between Energy Delivery Controller and Electrosurgical Generator SHALL use a dedicated isolated CAN bus at 1Mbit/s, transmitting power setpoint, modality selection, and enable/disable commands from controller to generator at 100Hz, with generator acknowledging each command within 5ms. | interface-requirements | Main |  | Test |  |  | CAN bus provides deterministic messaging with hardware-level error detection and isolation capability. 1Mbit/s supports 100Hz command rate with headroom for status return frames. 5ms acknowledgement timeout allows the controller to detect generator fault within one command cycle and trigger safety shutdown. Isolation prevents RF noise from the generator corrupting control traffic. | interface, energy-delivery, session-352, idempotency:ifc-edc-esg-can-352
IFC-MAIN-031 | The interface between Energy Delivery Controller and Ultrasonic Energy Module SHALL use a dedicated isolated RS-485 link at 115200 baud, carrying power level commands and blade temperature telemetry, with the module transmitting blade temperature to the controller at minimum 50Hz. | interface-requirements | Main |  | Test |  |  | RS-485 is appropriate for point-to-point ultrasonic generator control at the required data rate. 50Hz blade temperature telemetry gives the controller 20ms resolution for the 100°C inhibit threshold response — adequate given the thermal time constant of the blade (estimated 200-500ms for 1°C change at full power). | interface, energy-delivery, session-352, idempotency:ifc-edc-uem-rs485-352
IFC-MAIN-032 | The interface between Return Electrode Monitor and Electrosurgical Generator SHALL include a hardwired safety interlock line that the Return Electrode Monitor holds in a de-energised (fail-safe) state; the Electrosurgical Generator SHALL not enable monopolar output unless this line is actively energised by a Return Electrode Monitor that has confirmed pad impedance below 135 ohms. | interface-requirements | Main |  | Test |  |  | Hardwired fail-safe interlock (de-energise-to-inhibit) ensures monopolar energy is prohibited by hardware in the event of REM failure, communication loss, or power loss — no software path can override it. This is a common safety architecture in electrosurgical units per IEC 60601-2-2 and IEC 61508 SIL 2 guidance for patient-contact safety functions. | interface, energy-delivery, safety, session-352, idempotency:ifc-rem-esg-interlock-352
IFC-MAIN-033 | The interface between Tissue Effect Monitor and Electrosurgical Generator SHALL provide the monitor with access to real-time RF output voltage and current waveform samples at minimum 200kHz sampling rate, enabling impedance calculation at 1kHz; the monitor SHALL write shutoff commands to the generator via the CAN interface within 200ms of detecting seal endpoint. | interface-requirements | Main |  | Test |  |  | 200kHz sampling of RF waveforms (at 300kHz-3MHz carrier) requires adequate Nyquist headroom for the envelope extraction used in impedance calculation. 1kHz impedance update rate is the minimum for the 400ms detection window. 200ms shutoff command latency ensures the system meets the vessel seal endpoint response requirement (SUB-MAIN-052) with margin. | interface, energy-delivery, session-352, idempotency:ifc-tem-esg-impedance-352
IFC-MAIN-034 | The interface between Energy Delivery Controller and Safety and Interlock Subsystem SHALL use the system safety bus to receive system state broadcast at 100Hz; the Energy Delivery Controller SHALL inhibit all energy delivery within 20ms of receiving a SAFE_STATE or E-STOP signal on this bus. | interface-requirements | Main |  | Test |  |  | Energy delivery must respond to system-level emergency stop faster than the 50ms deactivation requirement (SUB-MAIN-048). 20ms allows the EDC to process the safety signal and command generator deactivation with 30ms margin for the generator hardware response. Bus-based integration avoids a dedicated hardwire for every energy modality, but the safety bus itself is hardware-isolated and deterministic. | interface, energy-delivery, safety, session-352, idempotency:ifc-edc-safety-bus-352
IFC-MAIN-035 | The interface between the Foot Pedal Array and the Energy Delivery Controller SHALL transmit energy modality (RF monopolar, RF bipolar, ultrasonic) and activation state (active/inactive) over a dedicated isolated CAN bus at 1Mbit/s, with a maximum message latency of 10ms. | interface-requirements | Main |  | Test |  |  | Energy activation commands are safety-critical: a delayed or missed deactivation command could deliver unintended energy to tissue. 10ms message latency budget is compatible with the 50ms end-to-end pedal latency requirement (SUB-MAIN-055) after accounting for pedal debounce (10ms) and controller processing (5ms). | interface, surgeon-console, energy-delivery, session-353, idempotency:ifc-pedal-edc-can-353
IFC-MAIN-036 | The interface between the Foot Pedal Array and the Motion Control and Scaling Subsystem SHALL transmit instrument clutch state (engaged/disengaged) and camera control commands over the console CAN bus within 10ms, using message priority higher than configuration traffic but lower than safety bus messages. | interface-requirements | Main |  | Test |  |  | Clutch engage/disengage decouples master arm motion from instrument motion, allowing the surgeon to reposition hands. A 10ms latency ensures sub-50ms end-to-end response (SUB-MAIN-055). Priority scheme prevents clutch messages from being starved by high-bandwidth configuration traffic. | interface, surgeon-console, motion-control, session-353, idempotency:ifc-pedal-mcs-clutch-353
IFC-MAIN-037 | The interface between the Console Computer and the Real-Time Protocol Engine SHALL exchange session management messages (case start, case end, configuration sync, surgeon identity) over gigabit Ethernet using a defined session protocol, with message delivery confirmation within 100ms. | interface-requirements | Main |  | Test |  |  | Session management messages configure the patient-side cart with the correct instrument parameters, scaling, and surgeon profile before motion is enabled. The 100ms confirmation window allows the Console Computer to detect a failed session start and display an error before the surgical team proceeds. | interface, surgeon-console, comms, session-353, idempotency:ifc-cc-rtpe-session-353
IFC-MAIN-038 | The interface between the Voice Command Module and the Console Computer SHALL deliver recognised command identifiers (from a defined vocabulary enumeration) with a confidence score and timestamp over USB 3.0, and SHALL NOT transmit raw audio data outside the module to protect patient privacy. | interface-requirements | Main |  | Test |  |  | Privacy constraint: raw audio captured in the OR may contain patient-identifiable information and must not be stored or transmitted beyond the recognition module. The command ID plus confidence score gives the Console Computer sufficient information to confirm or reject the command without retaining audio. | interface, surgeon-console, voice, privacy, session-353, idempotency:ifc-vcm-cc-commands-353
IFC-MAIN-039 | The interface between the Surgeon Interface Panel and the Console Computer SHALL carry master arm pose data at 1kHz in the upstream direction and haptic force commands at 1kHz in the downstream direction over a dedicated real-time EtherCAT bus, with end-to-end latency no greater than 1ms in each direction. | interface-requirements | Main |  | Test |  |  | The haptic control loop is closed across this interface; a bidirectional 1kHz rate and 1ms one-way latency are the minimum specification consistent with transparent teleoperation as defined in SYS-MAIN-001 (<1ms end-to-end motion scaling). EtherCAT is selected over CAN or EtherNet/IP because it provides deterministic cycle times below 100us supporting the 1kHz control rate with headroom for jitter tolerance. | interface, surgeon-console, haptic, session-356, idempotency:ifc-sip-cc-ethercat-356
IFC-MAIN-040 | The interface between the Console Computer and the Inter-Cart Fibre Link SHALL carry a continuous 6-DOF Cartesian velocity command stream encoded as 48-byte little-endian frames at 1kHz, with a per-frame sequence number and 16-bit CRC-CCITT checksum. | interface-requirements | Main |  | Test |  |  | The Console Computer is the upstream source of surgeon hand motion data. Explicit frame numbering at the interface allows the Real-Time Protocol Engine at the patient end to detect dropped frames immediately rather than processing stale data. CRC-CCITT provides the standard error detection for deterministic serial links, and is already used in the EtherCAT frame format for consistency. | interface, motion-control, infrastructure, inter-cart-fibre, session-357, idempotency:ifc-cc-icfl-command-357
IFC-MAIN-041 | The interface between the Inter-Cart Fibre Link and the Real-Time Protocol Engine SHALL deliver decoded 6-DOF velocity frames with a worst-case inter-frame delivery jitter of no more than 5 microseconds, providing the sequence number and CRC validity flag alongside each frame. | interface-requirements | Main |  | Test |  |  | The Real-Time Protocol Engine's TDM scheduler has a 1-microsecond jitter budget (SUB-MAIN-071). Optical transceiver deserialization and clock recovery introduce additional timing uncertainty; 5 microseconds is the worst-case SERDES CDR lock window for the selected 10GBase-SR transceiver at 10m. Validity flags allow the RTPE to apply the fault-halt policy (SUB-MAIN-072) without a separate CRC computation on the RT task. | interface, motion-control, infrastructure, inter-cart-fibre, real-time-protocol-engine, session-357, idempotency:ifc-icfl-rtpe-frames-357
IFC-MAIN-042 | The interface between the Network Management Controller and each Joint Servo Controller SHALL use EtherCAT with a bus cycle of 1ms, delivering a 16-byte process data object containing target joint angle (float32), feed-forward torque (float32), control mode (uint8), and node fault mask (uint8) per servo node per cycle. | interface-requirements | Main |  | Test |  |  | 1ms EtherCAT cycle matches the 1kHz kinematic pipeline rate end-to-end. The feed-forward torque field enables velocity-mode operation to complement the position PID in Joint Servo Controller, reducing steady-state tracking error at high speed. The fault mask byte allows per-node disablement without a separate out-of-band channel, consistent with IEC 61800-3 functional safety profile requirements for SIL 2 motion systems. | interface, motion-control, infrastructure, network-management, joint-servo, session-357, idempotency:ifc-nmc-jsc-ethercat-357
IFC-MAIN-043 | The interface between the Real-Time Compute Node and the Procedure Data Recorder SHALL transfer kinematic sample frames via PCIe DMA at 1kHz, where each frame contains a 64-bit UTC timestamp, 7-element float32 joint angle array, 7-element float32 joint torque array, and 6-element float32 Cartesian velocity array for each active arm, with a DMA transfer latency not exceeding 10 microseconds per frame. | interface-requirements | Main |  | Test |  |  | PCIe DMA is the only interconnect capable of sustaining continuous 1kHz structured data transfer from a PREEMPT_RT kernel with sub-100-microsecond latency without consuming CPU cycles on a memory-copy path. A 10-microsecond DMA ceiling leaves the real-time CPU interrupt budget intact (SUB-MAIN-011 specifies 50-microsecond ISR worst case). The timestamp must be captured in hardware at the RTPE frame boundary, not in software, to avoid jitter from OS scheduling. | interface, motion-control, infrastructure, procedure-data-recorder, session-357, idempotency:ifc-rtcn-pdr-dma-357
IFC-MAIN-044 | The interface between the Main Power Distribution Unit and the Power Sequencing Controller SHALL use a galvanically isolated CAN FD bus operating at 1 Mbit/s, transmitting branch current, voltage, and fault status at 10 Hz with a maximum latency of 5 ms. | interface-requirements | Main |  | Test |  |  | CAN FD at 1 Mbit/s provides sufficient bandwidth for 12-branch telemetry frames at 10 Hz while meeting automotive EMC standards tested for hospital OR environments. Galvanic isolation prevents ground loop noise from the high-current main bus from corrupting the low-voltage control bus, a known failure mode in earlier surgical robot power architectures. | interface, power-management, can-bus, session-361, idempotency:ifc-pdu-psc-can-361
IFC-MAIN-045 | The interface between the UPS Battery Module and the Main Power Distribution Unit SHALL provide a 48 VDC bulk link with a maximum impedance of 50 mOhm at the PDU input terminals, supporting peak discharge currents of up to 200 A for a maximum of 500 ms during startup inrush, and sustain 20 A continuous during surgical operation. | interface-requirements | Main |  | Test |  |  | The 48 VDC bus voltage is the minimum that allows the DC-DC converters within the PDU to maintain regulated outputs during rapid transient loads from arm actuator energisation (up to 180 A peak measured at worst-case six-axis simultaneous joint engagement). The 50 mOhm impedance limit ensures less than 1 V bus droop during 20 A continuous load, preventing brownout resets in the downstream servo controllers. | interface, power-management, ups, power-bus, session-361, idempotency:ifc-ups-pdu-dc-link-361
IFC-MAIN-046 | The interface between the Power Sequencing Controller and the Auxiliary Power Supply SHALL use an isolated discrete signal to command charge/standby/discharge modes, with mode transitions acknowledged within 50 ms via a return status signal. | interface-requirements | Main |  | Test |  |  | Discrete hardware signalling is used rather than bus-based control to ensure the Auxiliary Power Supply can receive commands even when the CAN bus has faulted, maintaining safety circuit availability in the worst-case scenario of bus communication failure during a power event. | interface, power-management, auxiliary-power, session-361, idempotency:ifc-psc-auxpsu-control-361
IFC-MAIN-047 | The interface between the Network Management Controller and the Safety and Interlock Subsystem SHALL transmit COMM_FAULT notifications within 2ms of link failure detection, using a dedicated LVDS hardwired signal independent of the fibre link under test. | interface-requirements | Main |  | Test |  |  | LVDS hardwired signal ensures the fault notification channel is not dependent on the fibre link it monitors — the fibre link cannot report its own failure. 2ms budget matches the Real-Time Protocol Engine fault reporting window and keeps total fault propagation under 5ms required by SUB-MAIN-028. | interface, comms, network-management, safety, session-369, idempotency:ifc-nmc-sis-fault-notify-369
IFC-MAIN-048 | The interface between the Real-Time Protocol Engine and the Procedure Data Recorder SHALL deliver a time-stamped stream of decoded kinematic frames at 1kHz using shared memory DMA transfer with less than 100us delivery latency, without interrupting the Real-Time Protocol Engine deterministic cycle. | interface-requirements | Main |  | Test |  |  | Shared memory DMA avoids adding any kernel scheduling latency to the RTPE deterministic cycle. 100us delivery budget allows the Procedure Data Recorder to buffer a full cycle before the next frame arrives. This decouples recording latency from the hard real-time control loop. | interface, comms, real-time-protocol-engine, procedure-data-recorder, session-369, idempotency:ifc-rtpe-pdr-dma-369
STK-MAIN-001 | The system SHALL enable surgeons to perform minimally invasive procedures with sub-millimetre tip precision, 3D visualisation, and hand-tremor elimination, reducing incision trauma compared to open surgery. | stakeholder-requirements | Main |  | Test |  |  | Core value proposition of the surgical robot — patients require reduced trauma and recovery time; surgeons require precision beyond unaided manual capability for complex anatomical regions. | stakeholder, core-capability, session-341, idempotency:stk-precision-001-341
STK-MAIN-002 | The system SHALL not deliver uncontrolled force, energy, or motion to the patient at any time, including during power failure, communication loss, instrument fault, or operator error. | stakeholder-requirements | Main |  | Test |  |  | Fundamental patient safety obligation. Uncontrolled actuator motion or energy delivery during loss of control is a Class III hazard (irreversible patient harm). Regulatory bodies (FDA 510(k), MDR) mandate fault-containment as a baseline approval criterion. | stakeholder, safety, session-341, idempotency:stk-safety-001-341
STK-MAIN-003 | The system SHALL integrate with existing operating room infrastructure including hospital networks, anaesthesia monitoring systems, and electrosurgery generators without requiring room modification. | stakeholder-requirements | Main |  | Inspection |  |  | Hospitals will not accept systems requiring structural OR modifications; retrofit-free integration is a purchasing prerequisite for the majority of target customers and is mandated by facilities management policies. | stakeholder, integration, session-341, idempotency:stk-or-integration-001-341
STK-MAIN-004 | All patient-side components entering or approaching the sterile field SHALL be designed for draping, sterilisation, or single-use, maintaining sterility throughout the procedure duration. | stakeholder-requirements | Main |  | Inspection |  |  | Surgical site infection (SSI) is a primary post-operative complication; regulatory sterility requirements (ISO 13485, EU MDR Annex I) impose design obligations. Breach of sterile field is a never-event in regulated healthcare environments. | stakeholder, sterility, session-341, idempotency:stk-sterility-001-341
STK-MAIN-005 | The surgeon console SHALL be ergonomically operable for procedures of up to 4 hours duration without inducing physical fatigue, and SHALL support a structured simulation-based training pathway to proficiency. | stakeholder-requirements | Main |  | Test |  |  | Surgeon fatigue during long procedures increases error rate and is a patient safety risk; ergonomic design is required by IEC 62366 (usability). Simulation training pathway is required by hospital credentialing committees before granting robotic surgery privileges. | stakeholder, ergonomics, training, session-341, idempotency:stk-ergonomics-001-341
STK-MAIN-009 | Surgeons shall be able to perform minimally invasive procedures with instrument precision exceeding freehand capability, including motion scaling and tremor elimination. | stakeholder-requirements | Main |  | Demonstration |  |  | Fundamental stakeholder need: surgical robot adoption is justified only if it enables operations beyond freehand limits. | stakeholder, precision, session-340
STK-MAIN-010 | The system shall prevent inadvertent patient tissue damage caused by control errors, communication failures, or unintended instrument motion. | stakeholder-requirements | Main |  | Test |  |  | Regulatory requirement per ISO 13485 and IEC 60601-1: active surgical devices cannot create new patient hazards. Uncontrolled instrument motion in a body cavity is immediately life-threatening. | stakeholder, safety, session-340
STK-MAIN-011 | All components contacting the patient or operating within the sterile field shall be sterilisable or provided as sterile-draped without compromising system function. | stakeholder-requirements | Main |  | Inspection |  |  | Infection control and regulatory requirement: surgical site infection from non-sterile instrumentation can be fatal; EN ISO 11135 sterilisation standards must be met for EU and US market approval. | stakeholder, sterility, session-340
STK-MAIN-012 | The surgeon shall maintain full situational awareness through high-definition 3D visualisation of the surgical field at all times during an active procedure. | stakeholder-requirements | Main |  | Demonstration |  |  | Stakeholder need from surgical user research: loss of visual feedback during robotic surgery is the leading cause of conversion to open surgery; 3D HD stereo has been shown to reduce depth-perception errors by 40% vs. 2D in laparoscopic studies. | stakeholder, vision, session-340
STK-MAIN-013 | The system shall enable rapid instrument exchange during a procedure without breaking the sterile field or requiring scrub nurse re-draping. | stakeholder-requirements | Main |  | Demonstration |  |  | Operational efficiency requirement: laparoscopic procedures require 8-15 instrument changes on average; each break in sterility adds risk and OR time, increasing cost and infection probability. | stakeholder, workflow, session-340
STK-MAIN-014 | The system shall support clinical data recording for post-operative review, audit, and regulatory reporting of each procedure. | stakeholder-requirements | Main |  | Inspection |  |  | Hospital governance and regulatory mandate: EU MDR Article 83 and FDA 21 CFR Part 820 require post-market surveillance data; malpractice liability creates institutional demand for full procedure recording. | stakeholder, data, session-340
STK-MAIN-015 | The system shall operate continuously for at least 8 hours without requiring maintenance intervention during scheduled surgical lists. | stakeholder-requirements | Main |  | Test |  |  | Operational requirement from theatre schedulers: a typical surgical list runs 7-9 hours with up to 4 procedures; unplanned system downtime forces cancellation, increasing patient waiting lists and OR costs. | stakeholder, availability, session-340
SUB-MAIN-001 | The Safety and Interlock Subsystem SHALL detect any single axis joint torque exceedance above 110% rated limit and initiate a controlled brake sequence within 50ms, and SHALL trigger an emergency stop within 20ms at 150% rated limit. | subsystem-requirements | Main |  | Test |  |  | Graduated response: 50ms at 110% allows controlled deceleration reducing tissue trauma risk; immediate stop at 150% prevents structural damage to arm or patient. Thresholds derived from actuator datasheet maximum continuous and peak ratings. | subsystem, sis, safety, session-341, idempotency:sub-sis-force-limits-001-341
SUB-MAIN-002 | When communication latency between surgeon console and patient-side cart exceeds 20ms or 3 consecutive frames are lost, the Safety and Interlock Subsystem SHALL freeze all arm motion and hold joints at current position within 10ms of threshold breach. | subsystem-requirements | Main |  | Test |  |  | 20ms / 3-frame threshold was selected as the point at which bilateral haptic control becomes unstable (Ryu et al. 2004 teleoperation stability criterion). At this threshold, uncontrolled arm motion during communication fault is safer than continued operation. | subsystem, sis, safety, communication, session-341, idempotency:sub-sis-comms-loss-001-341
SUB-MAIN-003 | The Safety and Interlock Subsystem SHALL achieve a complete safe state (all joints braked, all surgical energy de-energised) within 250ms of any safety event trigger, across all single-point failure modes. | subsystem-requirements | Main |  | Test |  |  | 250ms system-level budget is derived from SYS-MAIN-002. This SUB requirement allocates the full system budget to SIS as the responsible subsystem; downstream subsystems (servo drives, energy generator) must respond within their share of this window. | rt-sil-gap, red-team-session-502
SUB-MAIN-004 | The Watchdog Timer Controller SHALL operate on a processor physically isolated from the motion control CPU and SHALL maintain braking authority independently of motion control software state. | subsystem-requirements | Main |  | Inspection |  |  | Common-cause failure between safety monitor and controlled system is the primary SIL 3 architectural hazard. Physical isolation (separate processor, separate power rail, hardware brake authority) eliminates this common cause. Required by IEC 61508 SIL 3 architectural constraints (HFT=1). | rt-missing-failure-mode, red-team-session-502
SUB-MAIN-005 | The Emergency Stop Chain SHALL be a hardwired series loop independent of software, completing through all E-stop actuators (surgeon console, patient-side x3, facility), and SHALL de-energise servo drive contactors within 50ms of any break in the loop. | subsystem-requirements | Main |  | Test |  |  | Software-controlled E-stop cannot be relied upon as a safety function because software faults (deadlock, exception) are the fault mode it is intended to protect against. Hardwired series-loop is the IEC 60204 standard for Category 0 stop in machinery. | rt-sil-gap, red-team-session-502
SUB-MAIN-006 | The Motion Control System SHALL execute the complete kinematic computation pipeline (tremor filter, motion scaling, inverse kinematics, safety enforcement, servo command) within 10ms per 1kHz cycle. | subsystem-requirements | Main |  | Test |  |  | 10ms computation budget is the largest single allocation within the 100ms end-to-end system latency budget; remaining 90ms covers network (3ms), sensor acquisition (2ms), actuation settling (5ms), and display pipeline (50ms for video). If this budget is exceeded, the overall 100ms SYS-MAIN-007 cannot be met. | subsystem, motion-control, performance, session-340
SUB-MAIN-007 | The Motion Control System SHALL reject all Cartesian velocity command components above 6Hz by at least 40dB using the Tremor Rejection Filter before motion scaling is applied. | subsystem-requirements | Main |  | Test |  |  | Derived from SYS-MAIN-009: tremor filtration must occur before motion scaling in the pipeline to prevent scaling up residual filter artefacts. 40dB attenuation at 6Hz cutoff reduces 0.3mm/s tremor amplitude to 0.03mm/s, within acceptable tissue contact tolerance at 5:1 scaling. | subsystem, motion-control, tremor, session-340
SUB-MAIN-008 | The Kinematics Engine SHALL compute joint-angle setpoints for all 7 DOF of a single instrument arm within 2ms of receiving a Cartesian end-effector command. | subsystem-requirements | Main |  | Test |  |  | 2ms IK computation is the largest individual stage within the 10ms pipeline budget. Damped least-squares Jacobian pseudo-inverse complexity for 7-DOF is O(n^3) = O(343 FLOP), achievable in under 1ms on target hardware; 2ms includes singularity handling and redundancy resolution. | subsystem, motion-control, kinematics, session-340
SUB-MAIN-009 | The Joint Servo Controller SHALL achieve position tracking error below 0.1 degrees RMS during continuous trajectory following at maximum instrument velocity of 200mm/s tip speed. | subsystem-requirements | Main |  | Test |  |  | 0.1-degree joint error at 570mm arm reach produces approximately 1mm tip error, which at 10:1 motion scaling maps to 0.1mm at the surgeon console — below the surgeon's proprioceptive discrimination threshold of 0.3mm. Derived from system-level tip accuracy requirement. | rt-missing-failure-mode, red-team-session-502
SUB-MAIN-010 | The Workspace Safety Enforcer SHALL prevent any joint-angle command that would exceed hardware end-stop minus 5-degree software margin, and SHALL enforce the trocar-pivoting constraint to within 2mm of the insertion point at the abdominal wall. | subsystem-requirements | Main |  | Test |  |  | 5-degree software margin prevents mechanical joint binding under servo control; trocar-pivoting constraint to 2mm limits lateral force on the abdominal wall to under 2N based on tissue stiffness model, preventing port-site herniation or inadvertent viscus contact. | subsystem, motion-control, safety, session-340
SUB-MAIN-011 | The Real-Time Compute Node SHALL guarantee a worst-case interrupt latency of 50 microseconds on all Motion Control threads, and SHALL assert the hardware safety output within 5ms of detecting a motion-control thread heartbeat timeout. | subsystem-requirements | Main |  | Test |  |  | 50 microsecond PREEMPT_RT latency is achievable on target hardware (Intel Xeon with kernel 6.x PREEMPT_RT, measured P99.99 latency <45us in qualification testing). 5ms watchdog assertion leaves 45ms for the Safety and Watchdog System to complete brake engagement within the 50ms emergency stop budget. | rt-implausible-value, red-team-session-502
SUB-MAIN-012 | The Safe State Manager SHALL initiate all safety state transitions automatically, without requiring any operator action. Recovery from SAFE-HOLD to OPERATIONAL SHALL require an explicit, deliberate surgeon re-engagement sequence. | subsystem-requirements | Main |  | Test |  |  | Analog: Reactor Trip Subsystem (nuclear) requires automatic trip initiation — waiting for operator confirmation during a fault introduces unacceptable delay and human-error exposure. For surgical safety, the same principle applies: automatic entry to safe state, deliberate manual recovery. This design pattern is codified in IEC 61508 for safety instrumented systems. | subsystem, sis, safety, auto-initiation, session-341, idempotency:sub-sis-auto-initiation-001-341
SUB-MAIN-013 | The Stereo Endoscope SHALL provide a minimum optical resolution of 20 line pairs per millimetre across both channels, with less than 2% geometric distortion at the image periphery, to enable the surgeon to distinguish tissue structures at 0.5mm scale. | subsystem-requirements | Main |  | Test |  |  | Sub-millimetre instrument precision (STK-MAIN-001) requires the surgeon to resolve tissue structures at 0.5mm, which demands 20 lp/mm optical resolution. 2% distortion limit prevents spatial misjudgement during instrument manipulation near field edges. | subsystem, vision, endoscope, session-341, idempotency:sub-endoscope-resolution-341
SUB-MAIN-014 | The Camera Control Unit SHALL maintain inter-channel synchronisation between left and right stereo video streams with temporal skew not exceeding 500 microseconds, to prevent stereoscopic fusion artifacts that could cause surgeon depth misjudgement. | subsystem-requirements | Main |  | Test |  |  | Temporal skew above 500us between stereo channels causes perceptible depth shimmer during instrument motion at typical surgical velocities (5-20mm/s), degrading the surgeon's depth accuracy below the 1mm threshold required by SYS-MAIN-003. | subsystem, vision, ccu, session-341, idempotency:sub-ccu-sync-341
SUB-MAIN-015 | The Surgical Illumination Source SHALL regulate light intensity such that tissue surface temperature at the endoscope distal tip does not exceed 41 degrees Celsius under any operating mode, as measured at 10mm distance from the tip with tissue-equivalent thermal phantom. | subsystem-requirements | Main |  | Test |  |  | IEC 60601-2-18 mandates maximum tissue temperature limits for endoscopic illumination. The 41C limit applies at the distal tip where energy density is highest. Closed-loop regulation is required because tissue reflectance varies 4x between organ types, making fixed intensity unsafe. | subsystem, vision, illumination, session-341, idempotency:sub-illumination-thermal-341
SUB-MAIN-016 | The Image Processing Pipeline SHALL add no more than 2ms total processing latency from input frame reception to output frame availability, measured end-to-end across all processing stages including edge enhancement, noise reduction, and overlay compositing. | subsystem-requirements | Main |  | Test |  |  | Total system visual feedback latency budget is 50ms (surgeon hand motion to display update). Motion control consumes 10ms, display scan-out consumes 16ms. The image processing pipeline's 2ms allocation ensures the vision chain does not exceed its share of the latency budget, preventing surgeon-perceptible lag that degrades hand-eye coordination. | subsystem, vision, ipp, session-341, idempotency:sub-ipp-latency-341
SUB-MAIN-017 | The Stereoscopic Display System SHALL achieve less than 1% inter-channel crosstalk (ghosting) across the full luminance range at viewing distances between 500mm and 700mm, to prevent false depth cues that could cause instrument positioning errors. | subsystem-requirements | Main |  | Test |  |  | Crosstalk above 1% introduces ghost images that create false depth cues. During precise dissection near critical structures (nerves, vessels), even 2mm of apparent depth error from display crosstalk could result in inadvertent tissue damage. The 500-700mm viewing distance range covers ergonomic surgeon positioning. | subsystem, vision, display, session-341, idempotency:sub-display-ghosting-341
SUB-MAIN-017 | The Tool Tip Articulation Controller compute board SHALL operate from the 5V ±5% rail supplied by the Power Management Subsystem, drawing a maximum continuous power of 8W during full-bandwidth kinematic computation. | subsystem-requirements | Main | ubiquitous | Test |  |  | TTAC is a signal-processing compute board driving cable displacement commands to the IDU. The 8W budget is derived from FPGA plus motor-driver gate-drive logic at full computational load; the 5V rail is the standard embedded compute supply in the instrument drive chain.
SUB-MAIN-018 | The Procedure Video Recorder SHALL record composited 2D video continuously for at least 8 hours at 1080p60 resolution with H.265 encoding at 50Mbps CBR, with frame-accurate timestamps synchronised to the system event log within 1ms accuracy. | subsystem-requirements | Main |  | Test |  |  | SYS-MAIN-015 requires recording of all video streams at sufficient quality for post-operative review and audit. 8-hour continuous recording matches the system operational endurance requirement (SYS-MAIN-013). Frame-accurate synchronisation with the event log enables post-operative correlation of instrument movements with video for complication analysis. | subsystem, vision, recorder, session-341, idempotency:sub-recorder-duration-341
SUB-MAIN-018 | The Image Processing Pipeline SHALL operate from the 12V ±5% rail supplied by the Power Management Subsystem, drawing a maximum continuous power of 35W during dual-channel 1080p60 stereo processing. | subsystem-requirements | Main | ubiquitous | Test |  |  | The IPP is an FPGA-based image processing board processing two independent 1080p60 HD-SDI streams. The 35W budget is derived from FPGA device power at full utilisation plus DDR memory. 12V allows on-board DC-DC regulation for FPGA core and I/O rails.
SUB-MAIN-019 | The Surgeon Console SHALL operate from a 24V ±5% medical-grade isolated power supply, drawing a maximum continuous power of 120W including master manipulator motors, stereoscopic display panels, and embedded compute. | subsystem-requirements | Main | ubiquitous | Test |  |  | The Surgeon Console is the largest non-arm power consumer. The 120W budget covers dual 1080p display panels (~50W), master manipulator force-feedback motors (~40W), and embedded compute/comms (~30W). 24V medical-grade isolated supply is specified per IEC 60601-1 clause 8 for patient-proximate equipment.
SUB-MAIN-019 | When one stereo channel of the Stereo Endoscope fails, the Vision and Imaging System SHALL continue to provide the surgeon with 2D monocular video from the remaining channel at 1080p60 with no interruption exceeding 500ms, and SHALL display a persistent visual alert on the remaining channel indicating loss of stereoscopic depth perception. | subsystem-requirements | Main |  | Test |  |  | Complete loss of visual feedback during surgery is catastrophic. Single-channel failure must degrade gracefully to 2D rather than blacking out. The 500ms switchover limit ensures the surgeon does not lose visual contact with instruments in tissue. The persistent alert is required because operating without depth perception changes the surgical technique required. | subsystem, vision, degraded-mode, safety, session-341, idempotency:sub-vision-degraded-mono-341
SUB-MAIN-020 | The Camera Control Unit SHALL operate from the 12V ±5% rail supplied by the Power Management Subsystem, drawing a maximum continuous power of 18W during dual-channel 1080p60 HD-SDI capture and format conversion. | subsystem-requirements | Main | ubiquitous | Test |  |  | The CCU processes dual HD-SDI streams from the stereo endoscope. The 18W budget covers two SDI receivers, genlock circuitry, and format conversion ASICs. 12V supply shared with IPP to simplify rail distribution in the vision rack.
SUB-MAIN-020 | The Surgical Illumination Source SHALL provide fluorescence excitation at 805nm with irradiance of ≥5 mW/cm² at tissue surface to enable ICG fluorescence imaging at tissue depths up to 10mm, with mode switching between visible and NIR completing within 200ms. | subsystem-requirements | Main |  | Test |  |  | 5 mW/cm² is the minimum surface irradiance required for clinically detectable ICG fluorescence at 10mm tissue depth based on Beer-Lambert attenuation at 805nm. Below this threshold, signal-to-noise ratio falls below clinical utility (>3:1 contrast ratio) for sentinel node mapping. Value derived from published photon transport models for ICG in human tissue. | subsystem, vision, illumination, session-341, idempotency:sub-illumination-nir-341
SUB-MAIN-021 | The Haptic Feedback Subsystem SHALL operate from the 24V ±5% rail supplied by the Power Management Subsystem, drawing a maximum continuous power of 40W per instrument arm during simultaneous force rendering on both haptic actuator channels. | subsystem-requirements | Main | ubiquitous | Test |  |  | The HFS drives force-feedback actuators at the master manipulator with peak forces up to 5N per axis. The 40W per-arm budget is derived from actuator stall current headroom plus conditioning electronics. 24V rail is shared with master manipulator motor drive to minimise converter stages.
SUB-MAIN-021 | The Camera Control Unit SHALL provide at least three surgeon-selectable image enhancement modes (standard white light, narrow-band imaging for vascular contrast, and ICG fluorescence overlay) with mode switching completing within 100ms and no frame drops during transition. | subsystem-requirements | Main |  | Demonstration |  |  | Different surgical phases require different visualisation modalities. Narrow-band imaging enhances mucosal vascular patterns for tumour margin identification. ICG overlay shows perfusion. Mode switching must be seamless because the surgeon may need to toggle rapidly between views during active dissection near critical structures. | subsystem, vision, ccu, session-341, idempotency:sub-ccu-modes-341
SUB-MAIN-022 | The Force Signal Conditioner SHALL operate from the galvanically isolated 5V ±2% supply within the Haptic Feedback Subsystem, drawing a maximum continuous power of 3W during six-axis force/torque signal conditioning. | subsystem-requirements | Main | ubiquitous | Test |  |  | The FSC requires galvanic isolation from the main power bus per ARC-MAIN-004 to maintain patient electrical safety. The 5V isolated supply is generated by a dedicated isolated DC-DC converter in the HFS. The 3W budget covers six analogue conditioning channels plus ADC.
SUB-MAIN-022 | The Haptic Feedback Subsystem SHALL measure instrument-tissue interaction forces with a resolution of 0.05N and a range of ±30N across all six force/torque axes, sampled at 1kHz. | subsystem-requirements | Main |  | Test |  |  | Minimum force resolution of 0.05N is derived from the haptic discrimination threshold: human fingers detect force differences of approximately 0.1N; the sensor must resolve half this value to prevent perceptible quantisation steps during delicate tissue manipulation. Inadequate resolution produces a coarse, jerky haptic feedback that degrades surgical precision and may mask tissue damage. | subsystem, haptic, performance, session-342, idempotency:sub-haptic-force-res-342
SUB-MAIN-023 | When the Image Processing Pipeline watchdog timer expires without a valid frame completion token within 40ms, the Vision and Imaging System SHALL suppress the display output and assert an IMAGE_PIPELINE_FAULT signal to the Safety and Interlock Subsystem within 5ms. | subsystem-requirements | Main | event | Test |  |  | The IPP operates autonomously on a real-time FPGA pipeline; a hung or corrupted pipeline could deliver frozen or misleading frames to the surgeon. The 40ms watchdog (two missed frames at 50fps) triggers display suppression to prevent the surgeon acting on stale imagery, and SIS notification enables the safe-state decision chain.
SUB-MAIN-023 | The Haptic Feedback Subsystem SHALL deliver rendered force feedback to the master handle actuators with an end-to-end latency no greater than 2ms from force measurement at the instrument tip. | subsystem-requirements | Main |  | Test |  |  | The haptic control loop must close at >500Hz to maintain passivity and prevent energy accumulation that leads to instability (oscillation or divergence). A 2ms bound provides margin above the 1kHz sampling cycle while keeping the phase lag below 1 degree at the highest feedback bandwidth (20Hz), ensuring stable contact rendering on stiff tissue. | subsystem, haptic, performance, session-342, idempotency:sub-haptic-latency-342
SUB-MAIN-024 | The Console Computer SHALL implement network traffic isolation between the surgical control network and the hospital information network, with no direct data path between safety-critical control functions and external network interfaces. | subsystem-requirements | Main |  | Test |  |  | Console Computer is the only subsystem with external network access for hospital PACS/HIS integration. A bridge to the CAN-FD surgical control bus would create a remote code execution vector on the motion control chain. IEC 81001-5-1 mandates network boundary controls for Class IIb medical devices.
SUB-MAIN-024 | The Haptic Feedback Subsystem SHALL limit the maximum feedback force applied to the surgeon's master handles to 1N in any single axis under all operating conditions, including sensor fault conditions. | subsystem-requirements | Main |  | Test |  |  | Exceeding 1N feedback force risks startle response causing unintended master handle motion that is teleoperated to the instrument, potentially causing patient injury. The 1N limit is derived from IEC 80601-2-77 guidance on master device force limits in surgical robot systems. The limit must hold under sensor fault to prevent uncontrolled force buildup. | subsystem, haptic, safety, session-342, idempotency:sub-haptic-force-limit-342
SUB-MAIN-025 | The Force Signal Conditioner SHALL provide galvanic isolation of not less than 4kVrms (50Hz, 1 minute) between the strain gauge bridge circuit and the digital signal processing stage. | subsystem-requirements | Main |  | Test |  |  | IEC 60601-1 patient leakage current limits for Type CF applied parts require isolation of the patient-contact instrument circuit from the mains-connected digital electronics. A 4kVrms isolation barrier provides the required 1500V patient-to-earth working voltage margin with appropriate derating. Failure of this isolation constitutes a Class I critical hazard. | subsystem, haptic, safety, isolation, session-342, idempotency:sub-haptic-isolation-342
SUB-MAIN-026 | When the Force Sensing Module on one instrument arm reports a sensor fault, the Haptic Feedback Subsystem SHALL disable force feedback on that arm only, continue providing force feedback on all remaining arms at full specification, and notify the surgeon via a visual alert within 200ms. | subsystem-requirements | Main |  | Test |  |  | Disabling only the faulted arm preserves surgical utility on remaining arms, which is safer than total feedback loss since the surgeon retains haptic sense for other instruments. The 200ms alert requirement matches the minimum safe reaction window from IEC 62443 guidance for real-time operator warnings in safety-relevant systems. | subsystem, haptic, degraded-mode, safety, session-342, idempotency:sub-haptic-degraded-342
SUB-MAIN-027 | The Communication and Data Management System SHALL maintain end-to-end kinematic command transmission latency below 1ms from Surgeon Console to Patient-Side Cart under peak load (all 21 joint channels active, full video traffic). | subsystem-requirements | Main |  | Test |  |  | 1ms is the latency budget allocated to the communications layer in the 1ms end-to-end control loop specified by SYS-MAIN-001. Exceeding this budget causes the kinematics pipeline to run on stale commands, introducing effective dead time that degrades motion tracking and may cause trajectory overshoot at high motion speeds. | subsystem, comms, performance, session-342, idempotency:sub-comms-latency-342
SUB-MAIN-028 | The Communication and Data Management System SHALL detect fibre link failure within 5ms and complete switchover to the standby fibre path within 10ms, with no loss of kinematic command frames during the switchover. | subsystem-requirements | Main |  | Test |  |  | 10ms switchover time is derived from the safety requirement in SYS-MAIN-002 and SYS-MAIN-005: a 10ms gap in commands at 1kHz represents 10 missed frames, within the 100ms safe coast period defined for the motion controller. Faster detection (5ms) provides margin before the safety monitor declares a communication fault and initiates safe-hold. | subsystem, comms, reliability, safety, session-342, idempotency:sub-comms-failover-342
SUB-MAIN-029 | The Kinematics Engine SHALL authenticate all joint-space command inputs via HMAC-SHA256 signed frames, rejecting any command with an invalid signature within one 1ms control cycle and logging authentication failures to the Procedure Data Recorder. | subsystem-requirements | Main |  | Test |  |  | The Kinematics Engine is a purely digital component — a compromised kinematics computation could generate arm trajectories exceeding workspace limits, causing patient injury. HMAC-SHA256 per-frame authentication ensures only authorised motion commands reach the inverse kinematics solver. The 1ms rejection window matches the 1kHz servo rate.
SUB-MAIN-029 | The Procedure Data Recorder SHALL record all kinematic data at 1kHz, both stereo video streams, and all system events without data loss for a minimum continuous operating period of 8 hours, with post-procedure data protected in WORM mode. | subsystem-requirements | Main |  | Test |  |  | 8-hour capacity matches SYS-MAIN-013 (operational capability for 8 consecutive hours). WORM protection is required by FDA 21 CFR Part 820 (QSR) device history records and by IEC 62304 for surgical device audit trails. Data loss during recording constitutes a regulatory non-conformance and may compromise post-incident investigation. | subsystem, comms, recording, compliance, session-342, idempotency:sub-comms-recorder-342
SUB-MAIN-030 | The Trajectory Generator SHALL validate all motion waypoints against a cryptographically signed workspace envelope before generating trajectory segments, rejecting any waypoint outside the signed envelope and initiating a controlled stop within 50ms. The envelope definition SHALL be loaded from write-protected memory at startup and verified by RSA-2048 signature. | subsystem-requirements | Main |  | Test |  |  | The Trajectory Generator is a purely digital component computing motion paths for all robot arms. An attacker with access to waypoint inputs could inject trajectories moving instruments outside the sterile field or into anatomical structures. Signing the workspace envelope and validating each waypoint prevents malicious waypoints from being executed. RSA-2048 signature protection is consistent with IEC 62443 requirements for safety-critical motion systems.
SUB-MAIN-030 | When the Real-Time Protocol Engine detects a frame sequence error or CRC failure, the Communication and Data Management System SHALL discard the corrupted frame, log the error with timestamp and channel identifier, and continue processing the next valid frame without resetting the communication channel. | subsystem-requirements | Main |  | Test |  |  | Frame-level error recovery is preferred over channel reset because a reset introduces a 10ms+ dead period that would trigger the comms-loss safety path. Discarding a single corrupted frame produces at most one missed command cycle (1ms), which the motion controller can interpolate safely. Logging enables post-procedure analysis of communication reliability. | subsystem, comms, fault-handling, session-342, idempotency:sub-comms-framing-342
SUB-MAIN-031 | The Real-Time Protocol Engine SHALL authenticate all synchronisation messages using IEEE 1588v2 PTP with HMAC-SHA256 message authentication codes, discarding timing frames with invalid MACs and logging each rejection. When authenticated timing frames are unavailable for more than 10ms, the system SHALL enter a safe hold state and alert the surgeon console. | subsystem-requirements | Main |  | Test |  |  | The Real-Time Protocol Engine distributes the master clock used by every servo controller and sensor sampler. A time-injection attack shifting the distributed clock could cause phase misalignment between motion command generation and joint servo execution, producing uncontrolled arm movements. IEEE 1588v2 with HMAC authentication prevents clock spoofing. The 10ms safe-hold threshold is the maximum permissible phase error before joint servo controllers saturate error integrals and produce runaway torque commands.
SUB-MAIN-032 | The Instrument Recognition Module SHALL read and validate the instrument identity chip within 200ms of mechanical coupling detection, providing instrument type code, calibration offsets, remaining use count, and sterilization history to the Tool Tip Articulation Controller and Instrument Lifecycle Controller. | subsystem-requirements | Main | event | Test |  |  | STK-MAIN-013 requires rapid instrument exchange. The 200ms budget derives from 15-second total swap time: 10s manual handling, 3s coupling, 2s recognition and reconfiguration. 200ms chip read ensures recognition is not the bottleneck. | surgical-instrument-system, instrument-recognition, session-346
SUB-MAIN-033 | The Instrument Drive Unit SHALL actuate all four instrument degrees of freedom (wrist pitch, yaw, roll, and grip) with a position accuracy of +/-0.1mm at the instrument tip across the full 10-procedure instrument lifetime, at a servo update rate of 1kHz. | subsystem-requirements | Main | ubiquitous | Test |  |  | SYS-MAIN-001 specifies master-to-slave motion scaling. The 0.1mm tip accuracy is the instrument subsystems share of the overall 1mm system accuracy budget. The 1kHz rate matches the motion control servo loop. Accuracy must hold across instrument lifetime because cable stretch degrades positioning. | surgical-instrument-system, instrument-drive-unit, session-346
SUB-MAIN-034 | The Cable Tensioning System SHALL maintain cable tension on all four instrument DoF cables within +/-5% of the instrument-specific nominal set-point, and SHALL detect cable tension deviation exceeding 15% within 10ms, reporting a tension anomaly to the Safety and Interlock Subsystem. | subsystem-requirements | Main | ubiquitous | Test |  |  | Cable tension directly governs instrument tip accuracy. The 5% tolerance derives from the 0.1mm tip accuracy requirement and the cable-to-tip displacement ratio of approximately 4:1. A 15% deviation indicates fraying, disconnection, or mechanical failure requiring safety intervention. The 10ms detection window ensures motion arrest before tip displacement exceeds the safe envelope. | surgical-instrument-system, cable-tensioning, session-346
SUB-MAIN-035 | The Sterile Adapter SHALL maintain sterile barrier integrity per ISO 11607-1 under continuous operating loads of 50N axial force and 2Nm torque per rotary feedthrough channel for a single surgical procedure up to 8 hours, and SHALL transmit torque through all six sealed rotary feedthroughs with no more than 5% torque loss. | subsystem-requirements | Main | ubiquitous | Test |  |  | SYS-MAIN-006 and STK-MAIN-004/008/011 require sterile field compliance. The 50N and 2Nm loads represent worst-case instrument insertion force and wrist actuation torque from cadaver studies of complex procedures. The 5% torque loss limit ensures the Cable Tensioning System can compensate without saturating. 8-hour duration covers extended procedures. | surgical-instrument-system, sterile-adapter, session-346
SUB-MAIN-036 | The Tool Tip Articulation Controller SHALL compute cable displacement commands for all four instrument DoF from a desired end-effector pose within 500 microseconds worst-case latency, loading instrument-specific kinematic models from the Instrument Recognition Module at instrument coupling time. | subsystem-requirements | Main | ubiquitous | Test |  |  | The 1kHz motion control loop allocates 1ms per cycle. The Tool Tip Articulation Controller shares this cycle with the Kinematics Engine and Joint Servo Controller. The 500us budget is the instrument subsystems allocation after 300us for forward kinematics and 200us for servo command dispatch. Exceeding this budget causes jitter visible as instrument tip tremor. | surgical-instrument-system, articulation-controller, session-346
SUB-MAIN-037 | The Instrument Lifecycle Controller SHALL prevent coupling of any instrument that has exceeded its manufacturer-defined use limit (actuation cycles, sterilization count, or calendar age), inhibiting arm enable via the Safe State Manager until a valid instrument is detected, and SHALL log all lifecycle events to the Procedure Data Recorder per FDA 21 CFR Part 820. | subsystem-requirements | Main | ubiquitous | Demonstration |  |  | STK-MAIN-007/010 require prevention of inadvertent patient tissue damage. An instrument past its rated lifecycle has degraded cable integrity, worn joints, and reduced force accuracy. Regulatory traceability (21 CFR 820) mandates that every instrument use event is recorded with disposition. Arm lockout is the enforcement mechanism because the Safe State Manager is the single authority for arm enable. | surgical-instrument-system, lifecycle-controller, session-346
SUB-MAIN-038 | When the Cable Tensioning System detects a tension anomaly on any single cable, the Surgical Instrument System SHALL disable the affected instrument arm within 50ms while maintaining full motion control on all remaining instrument arms, and SHALL display the affected arm identity and failure type on the surgeon console. | subsystem-requirements | Main | event | Test |  |  | SYS-MAIN-016 requires graceful degradation when one arm fails. A cable tension anomaly means the affected instruments tip position is no longer trustworthy. The 50ms shutdown window is derived from the maximum safe tip displacement at full operating speed (100mm/s): 5mm of uncontrolled travel is the safety limit. Other arms must remain operational because mid-procedure instrument loss is recoverable but total system shutdown may endanger the patient. | surgical-instrument-system, degraded-mode, session-346
SUB-MAIN-039 | The Trajectory Generator SHALL compute interpolated Cartesian pose setpoints at 1kHz with S-curve velocity profiling, limiting instrument tip acceleration to 2g and jerk to 50g/s, ensuring smooth instrument motion that does not induce tissue tearing or excessive contact forces. | subsystem-requirements | Main |  | Analysis |  |  | Trajectory smoothness directly governs tissue interaction safety. The 2g acceleration limit derives from biomechanical studies of safe tissue manipulation forces during laparoscopic procedures. S-curve profiling eliminates jerk discontinuities that cause vibration in the cable-driven transmission, which degrades instrument tip position accuracy. The 1kHz rate matches the servo loop frequency to avoid interpolation artefacts. | subsystem, motion-control, session-348, idempotency:sub-trajectory-generator-motion-profile-348
SUB-MAIN-040 | When the primary haptic force-rendering processor fails, the haptic feedback subsystem SHALL switch to a secondary rendering path within 50 ms, maintaining contact-force reproduction accuracy within 20 percent of nominal and holding maximum perceivable force above 5 N until the surgeon withdraws the instrument. | subsystem-requirements | Main |  | Test |  |  | Haptic feedback is system-essential: loss of force cues mid-dissection prevents tissue plane discrimination and risks inadvertent perforation. 50 ms switchover is within the 200 ms perceptual threshold for force discontinuity per IEC 60601-1 Clause 14. | redundancy, safety, haptics, session-367
SUB-MAIN-040 | The Motion Scaling Module SHALL apply the surgeon-selected scaling ratio (3:1, 5:1, or 10:1) to filtered Cartesian velocity commands with gain accuracy of ±0.5%, and SHALL complete the scaling computation within 100 microseconds per cycle to maintain pipeline timing margin. | subsystem-requirements | Main |  | Test |  |  | Scaling accuracy of ±0.5% ensures the surgeon perceives consistent motion amplification across the workspace. At 10:1 scaling, a 0.5% error corresponds to 50μm at 10mm instrument travel — within the 100μm instrument tip repeatability budget. The 100μs execution budget allocates half the 200μs total pipeline margin to downstream stages (Trajectory Generator, Kinematics Engine). | subsystem, motion-control, session-348, idempotency:sub-motion-scaling-accuracy-348
SUB-MAIN-041 | When the primary IEEE 1588 grandmaster clock source fails or exceeds 1 microsecond offset from UTC, the time protocol engine SHALL switch to the hot-standby grandmaster within 200 ms without disrupting subsystem synchronisation by more than 5 microseconds, and SHALL log the switchover event with timestamp and root cause code. | subsystem-requirements | Main |  | Test |  |  | Time protocol engine is system-essential: inter-subsystem synchronisation loss causes motion command latency spikes that manifest as jerky or uncontrolled arm motion. The 5 microsecond continuity window is derived from the motion control loop rate of 4 kHz; offsets beyond this cause missed motion cycles. 200 ms switchover is within the 500 ms maximum tolerable disruption defined in IEC 62304 architectural risk analysis. | redundancy, timing, session-367
SUB-MAIN-041 | When the Trajectory Generator detects that a computed trajectory segment would exceed the workspace boundary or violate acceleration limits, the Trajectory Generator SHALL clamp the output to the last safe pose and assert a trajectory-violation flag to the Workspace Safety Enforcer within 1ms, halting all further interpolation until the Workspace Safety Enforcer acknowledges the clamp. | subsystem-requirements | Main |  | Test |  |  | The Trajectory Generator is the first stage that can detect demand violations before they propagate to the Kinematics Engine. Without this clamp, an out-of-bounds demand would force the Workspace Safety Enforcer to reject the entire joint-angle command downstream, causing a harder motion discontinuity. Clamping at the Cartesian level preserves motion smoothness during limit events. The 1ms assertion budget allows the violation to be caught within the same control cycle. | subsystem, motion-control, safety, session-348, idempotency:sub-tg-safety-clamp-348
SUB-MAIN-042 | When primary AC mains supply fails or drops below 85 VAC, the power management subsystem SHALL transfer all patient-safety loads to the internal UPS within 10 ms, sustaining surgical operations at full rated power for a minimum of 15 minutes to allow orderly procedure completion and instrument retraction. | subsystem-requirements | Main |  | Test |  |  | Power management is system-essential and Regulated: an uncontrolled power loss mid-procedure immobilises robotic arms in situ, creating a patient entrapment hazard and preventing safe instrument withdrawal. 10 ms transfer time is derived from the motion controller's minimum command cycle; slower transfer causes the servo watchdog to trigger an uncontrolled halt. 15 minutes is the clinical consensus minimum for orderly procedure close cited in IEC 60601-1-1 Annex J. | redundancy, power, safety, session-367
SUB-MAIN-042 | When the Motion Scaling Module receives a velocity command magnitude exceeding 200mm/s (corresponding to maximum safe instrument tip velocity at 1:1 scaling), the Motion Scaling Module SHALL reject the command, hold the last valid output, and report an over-velocity fault to the Safety and Interlock Subsystem within 500 microseconds. | subsystem-requirements | Main |  | Test |  |  | The Motion Scaling Module is the earliest point in the pipeline where absolute velocity limits can be enforced independent of the selected scaling ratio. The 200mm/s threshold corresponds to the maximum safe instrument tip velocity derived from tissue damage studies — exceeding this at any scaling ratio indicates either a sensor fault or an uncontrolled input. The 500μs detection budget ensures the fault is flagged within the same control cycle, preventing propagation to the Trajectory Generator. | subsystem, motion-control, safety, session-348, idempotency:sub-msm-safety-overvel-348
SUB-MAIN-043 | When the primary procedure data recorder storage medium fails or write latency exceeds 200 ms, the procedure data recorder SHALL simultaneously stream all kinematic, video, and event data to a secondary hot-standby recorder such that no more than 2 seconds of procedure data is lost and recording resumes on the backup without operator intervention. | subsystem-requirements | Main |  | Test |  |  | Procedure data recorder is system-essential: complete surgical records are required by IEC 62304 and regulatory bodies (FDA 21 CFR Part 820) for post-incident reconstruction, device liability, and surgical training. A 2-second data gap is the maximum acceptable loss established by clinical risk analysis; longer gaps may obscure the causal chain in adverse event investigations. | redundancy, data-recording, regulatory, session-367
SUB-MAIN-043 | The Power Management Subsystem SHALL maintain all surgical robot system functions from the UPS Battery Module for a minimum of 30 minutes following loss of mains power, permitting controlled procedure completion and safe shutdown. | subsystem-requirements | Main |  | Test |  |  | IEC 60601-1 clause 11.8.4 requires medical electrical equipment to maintain operation from internal energy source for defined periods; 30 minutes is the minimum required to complete a laparoscopic procedure segment and park all arms safely. Loss of power mid-procedure constitutes a patient safety event. | subsystem, power-management, session-350, idempotency:sub-power-ups-duration-350
SUB-MAIN-044 | The interlock subsystem SHALL be designed, verified, and validated to IEC 61508 SIL 3, achieving a PFH (probability of dangerous failure per hour) of less than 1e-7 per hour, with documented FMEA, fault-injection testing, and independent third-party assessment prior to regulatory submission. | subsystem-requirements | Main |  | Analysis |  |  | Interlock subsystem is Regulated and classified as the last line of defence against uncontrolled robotic motion near the patient. SIL 3 is required because the hazardous event (uncontrolled arm motion causing patient injury) is classified as Catastrophic + Frequent in the system-level HAZOP. IEC 61508 is the applicable functional safety standard for programmable electronic safety systems in medical robotics per IEC 80601-2-77. | compliance, safety, SIL3, regulatory, session-367
SUB-MAIN-044 | The Power Management Subsystem SHALL energise subsystems in the following order during startup: Safety and Interlock Subsystem, Auxiliary Power Supply, Communication and Data Management System, Motion Control System, Surgical Instrument System. Shutdown SHALL reverse this sequence. | subsystem-requirements | Main |  | Inspection |  |  | Energising the safety subsystem first ensures that the watchdog and E-stop chain are active before any motion-capable subsystem receives power. Reverse-sequence shutdown ensures that motion-capable drives are de-energised before safety supervision is withdrawn. Any other order creates a window where actuators are powered without protection. | subsystem, power-management, session-350, idempotency:sub-power-sequencing-order-350
SUB-MAIN-045 | The motion control system software SHALL be developed to IEC 62304 Class C (safety class), with complete requirements traceability, code review, unit and integration test evidence, and a software hazard analysis prior to release. All Class C software modules SHALL achieve modified condition/decision coverage (MC/DC) of 100 percent. | subsystem-requirements | Main |  | Inspection |  |  | Motion control system is Regulated: its software directly commands robotic arm position; a defect causing unintended motion constitutes a Catastrophic hazard under ISO 14971. IEC 62304 Class C is mandatory for software whose failure can cause death or serious injury; MC/DC 100 percent is required by DO-178C and adopted by IEC 62304 supplementary guidance for safety-critical medical motion control. | compliance, software, IEC62304, session-367
SUB-MAIN-045 | The Auxiliary Power Supply SHALL remain energised and supply the Safety and Interlock Subsystem, Watchdog Timer Controller, and Emergency Stop Chain contactor coils during any main bus fault, brownout below 85% nominal, or deliberate mains disconnection. | subsystem-requirements | Main |  | Test |  |  | The Safety and Interlock Subsystem must remain operational to initiate controlled safe-state during power faults. A brownout that de-energises the watchdog before safe-state is reached would leave joint motors in an undefined state — creating a patient harm risk equivalent to a software crash. | subsystem, power-management, session-350, idempotency:sub-power-aux-isolation-350
SUB-MAIN-046 | The workspace safety enforcer SHALL comply with IEC 80601-2-77:2021 Clause 201.11 (Accuracy of controls and instruments) and ISO 10218-1:2011 Clause 5.4 (Safety-rated monitored stop), achieving a safety-rated monitored stop reaction time of less than 10 ms from workspace boundary violation detection to servo torque cutoff. | subsystem-requirements | Main |  | Test |  |  | Workspace safety enforcer is Regulated: it prevents robotic arms from entering anatomical exclusion zones during surgery. IEC 80601-2-77 is the applicable collateral standard for surgical robot systems; ISO 10218-1 Clause 5.4 provides the safety-rated stop performance benchmark. 10 ms is derived from the worst-case arm velocity (0.5 m/s) and maximum tolerable penetration depth (5 mm) before tissue contact. | compliance, safety, workspace, regulatory, session-367
SUB-MAIN-046 | While the Image Processing Pipeline is processing stereoscopic video, it SHALL detect any frame that contains artefacts exceeding 5% pixel corruption or latency exceeding 33ms (2-frame drop at 60fps) and substitute a frozen clean frame, generating a latency-exceeded alert to the Surgeon Console within 10ms of detection. | subsystem-requirements | Main |  | Test |  |  | Lint finding: Image Processing Pipeline is Functionally Autonomous (FPGA-based, minimal external oversight). A corrupted or latency-violated frame delivered to the surgeon without notification constitutes a patient safety risk — the surgeon may act on an inaccurate view of the surgical field. Frozen frame substitution is the standard safety response in medical video chains (IEC 62304 guidance); the 5% threshold is derived from psychophysical research showing artefacts above this level impair depth perception in stereoscopic displays. | subsystem, vision, safety, session-350, idempotency:sub-image-safety-constraint-350
SUB-MAIN-047 | The console computer SHALL be qualified as a medical device accessory per EU MDR 2017/745 Annex I and FDA 21 CFR Part 820, with a Quality Management System certified to ISO 13485:2016, and SHALL display a current CE mark and 510(k) clearance number on the labelling prior to clinical deployment. | subsystem-requirements | Main |  | Inspection |  |  | Console computer is Regulated: as the primary surgeon-facing interface that initiates and controls all robotic motion, it is a medical device accessory that requires regulatory approval before clinical use. EU MDR and FDA 510(k) clearance are the applicable market authorisations; ISO 13485 QMS certification is a prerequisite for both submission pathways. | compliance, regulatory, MDR, session-367
SUB-MAIN-047 | The Electrosurgical Generator SHALL produce monopolar RF output power of 10-400W and bipolar RF output power of 10-80W at frequencies between 300kHz and 3MHz, with actual output power within 10% of the selected setting across the load impedance range of 100Ω to 2kΩ. | subsystem-requirements | Main |  | Test |  |  | Power range and impedance bounds are derived from IEC 60601-2-2 and clinical electrosurgical unit specifications. Monopolar 400W maximum covers major vessel haemostasis; bipolar 80W maximum covers delicate tissue coagulation. 10% output accuracy prevents surgeon from applying unintended power levels that could cause deep thermal injury. | subsystem, energy-delivery, performance, session-352, idempotency:sub-eds-esg-rf-power-352
SUB-MAIN-048 | The Electrosurgical Generator SHALL achieve full output power within 100ms of receiving an activation command and SHALL reduce output to below 1W within 50ms of receiving a deactivation command. | subsystem-requirements | Main |  | Test |  |  | Activation 100ms bound allows surgeon intent recognition by the Energy Delivery Controller before energy reaches tissue. Deactivation 50ms bound is the safety-critical parameter: at 400W monopolar, 50ms corresponds to approximately 20J of unintended energy delivery — the upper threshold for acceptable unintended thermal damage per IEC 60601-2-2 risk analysis. | subsystem, energy-delivery, safety, performance, session-352, idempotency:sub-eds-esg-latency-352
SUB-MAIN-049 | The Ultrasonic Energy Module SHALL drive the ultrasonic transducer at 55.5kHz ± 500Hz with selectable power levels from 10% to 100% in 10% increments, and SHALL detect blade temperature via thermocouple and inhibit activation when blade temperature exceeds 100°C. | subsystem-requirements | Main |  | Test |  |  | 55.5kHz is the standard resonant frequency for Harmonic-class ultrasonic surgical devices; ± 500Hz tolerance maintains resonance efficiency above 95%. 100°C blade temperature inhibit prevents retained heat burns from a previously activated blade contacting unintended tissue — a documented adverse event class in ultrasonic surgery (FDA MAUDE database, 2015-2022). | subsystem, energy-delivery, performance, safety, session-352, idempotency:sub-eds-uem-freq-temp-352
SUB-MAIN-050 | The Energy Delivery Controller SHALL enforce mutual exclusion between RF and ultrasonic modalities such that activation of one modality SHALL automatically inhibit the other, with the inhibition taking effect within 10ms of detecting concurrent activation requests. | subsystem-requirements | Main |  | Test |  |  | Simultaneous RF and ultrasonic energy delivery through the same instrument port could generate resonance artefacts in tissue or overload instrument drive electronics. 10ms inhibition response is derived from the 100ms activation latency requirement — inhibition must be faster than the energy-reaching-tissue window. | subsystem, energy-delivery, safety, session-352, idempotency:sub-eds-edc-mutual-exclusion-352
SUB-MAIN-051 | The Return Electrode Monitor SHALL continuously sample patient return electrode impedance at a minimum of 100Hz and SHALL inhibit monopolar energy delivery within 500ms when sampled impedance exceeds 135 ohms, with inhibition persisting until the pad impedance is restored and the surgeon actively resets the alarm. | subsystem-requirements | Main |  | Test |  |  | 135Ω threshold and 500ms response are per IEC 60601-2-2 Annex J requirements for REM circuits. Partial pad lift is the primary cause of alternate site burns in monopolar electrosurgery; active surgeon reset prevents automatic re-enable after transient impedance excursion, ensuring the surgeon consciously acknowledges the alarm before energy resumes. | rt-missing-failure-mode, red-team-session-502
SUB-MAIN-052 | The Tissue Effect Monitor SHALL detect vessel seal completion by impedance rise signature and SHALL command the Electrosurgical Generator to cease energy delivery within 200ms of endpoint detection, with endpoint defined as an impedance increase of at least 1.5kΩ occurring within any 400ms window during active vessel sealing. | subsystem-requirements | Main |  | Test |  |  | 1.5kΩ rise in 400ms is derived from validated impedance signatures for collagen denaturation in bipolar vessel sealing (cf. LigaSure algorithm patent family, published clinical studies on tissue impedance monitoring). Automatic cutoff prevents over-application beyond seal completion, which is the primary cause of vessel charring and reduced seal burst strength. | subsystem, energy-delivery, performance, session-352, idempotency:sub-eds-tem-endpoint-352
SUB-MAIN-053 | The Energy Delivery Controller SHALL automatically terminate energy delivery and generate a surgeon console alarm if continuous activation exceeds 15 seconds for RF modality or 5 seconds for ultrasonic modality, with resumption requiring explicit surgeon re-activation. | subsystem-requirements | Main |  | Test |  |  | 15s RF timeout prevents inadvertent extended electrosurgery from a held footswitch (e.g., surgeon distraction or equipment malfunction). 5s ultrasonic timeout is shorter because blade thermal accumulation above the 100°C inhibit threshold occurs within 7-10s of continuous activation at full power — the 5s limit provides 2-5s safety margin. Both timeouts are surgeon-acknowledged to prevent silent energy restart. | subsystem, energy-delivery, safety, session-352, idempotency:sub-eds-edc-timeout-352
SUB-MAIN-054 | The Energy Delivery System SHALL provide electrical isolation between the energy delivery circuit and the robotic control network such that patient leakage current through any energy path does not exceed 10μA in normal condition and 50μA in single-fault condition, per IEC 60601-1 Type CF applied part classification. | subsystem-requirements | Main |  | Test |  |  | Type CF classification is mandatory for equipment making direct cardiac contact or used near the heart — surgical robot instruments may be used in cardiac procedures. 10μA/50μA limits are per IEC 60601-1 Table 1 for Type CF applied parts. Isolation is safety-critical: conductive coupling between the RF generator and servo control bus could introduce high-frequency noise causing servo instability or patient microshock. | subsystem, energy-delivery, safety, session-352, idempotency:sub-eds-isolation-leakage-352
SUB-MAIN-055 | The Foot Pedal Array SHALL transmit energy activation, clutch, and camera control pedal events to the corresponding subsystem controllers within 50ms of pedal mechanical actuation. | subsystem-requirements | Main |  | Test |  |  | 50ms is the maximum latency consistent with transparent surgeon control: beyond 50ms, pedal-to-action delays become perceptible and can cause unintended energy delivery or motion. The E-stop pedal is hardwired and has a separate sub-millisecond hardware path; this requirement covers only the software-mediated pedals. | subsystem, surgeon-console, foot-pedal, session-353, idempotency:sub-sic-pedal-latency-353
SUB-MAIN-056 | The Voice Command Module SHALL recognise predefined surgical commands from the active vocabulary with a word error rate below 5% in an operating room acoustic environment with background noise up to 65dB SPL. | subsystem-requirements | Main |  | Test |  |  | A 5% WER ceiling ensures that no more than 1 in 20 commands is misrecognised in the worst-case OR noise environment. Validated against reference vocabulary of 200 surgical commands. Higher error rates create workflow interruption and require repeated commands, increasing procedure duration. | subsystem, surgeon-console, voice, session-353, idempotency:sub-sic-voice-wer-353
SUB-MAIN-057 | The Voice Command Module SHALL dispatch a recognised command to the Console Computer within 200ms of speech onset. | subsystem-requirements | Main |  | Test |  |  | 200ms is the perceptibility threshold for voice-to-action feedback in human factors studies for surgical environments. Commands exceeding 200ms latency break cognitive flow and may lead the surgeon to repeat the command, creating duplicate-command hazard. | subsystem, surgeon-console, voice, session-353, idempotency:sub-sic-voice-latency-353
SUB-MAIN-058 | The Console Computer SHALL require surgeon authentication before enabling robotic motion, and SHALL log the authenticated user identity, authentication time, and case start time to the Procedure Data Recorder. | subsystem-requirements | Main |  | Test |  |  | Regulatory requirement under MDR 2017/745 and FDA 21 CFR Part 820 for medical device traceability. Authentication prevents unauthorised use and creates an auditable record linking each surgical case to a credentialled operator. | subsystem, surgeon-console, authentication, compliance, session-353, idempotency:sub-sic-auth-353
SUB-MAIN-059 | The Arm Positioning System SHALL prevent motorised adjustment of any axis while the system is in OPERATIONAL state, and SHALL complete position lock-out within 500ms of the system entering OPERATIONAL state. | subsystem-requirements | Main |  | Test |  |  | Arm movement during robotic operation would change the master-to-slave kinematic calibration mid-procedure, causing instrument tip position error. 500ms lock-out allows the safety state broadcast (IFC-MAIN-004) to propagate and the positioning motor controllers to confirm brake engagement before motion commands are accepted. | subsystem, surgeon-console, arm-positioning, safety, session-353, idempotency:sub-sic-armpos-lockout-353
SUB-MAIN-060 | The Console Computer SHALL complete a startup self-test within 90 seconds of power-on, verifying connectivity to the Surgeon Interface Panel, Voice Command Module, Arm Positioning System, and Real-Time Protocol Engine, and SHALL display the test result on the Surgeon Interface Panel before enabling case start. | subsystem-requirements | Main |  | Test |  |  | 90-second startup target matches pre-operative room setup workflow; longer startup delays operating theatre utilisation. Self-test covers all console-side interfaces so fault detection occurs before the surgical team proceeds to patient positioning. | subsystem, surgeon-console, startup, session-353, idempotency:sub-sic-startup-selftest-353
SUB-MAIN-061 | When the Voice Command Module fails or is disabled, the Surgeon Input Console SHALL continue to provide full surgical system control through the Surgeon Interface Panel and Foot Pedal Array with no reduction in motion control, energy delivery, or camera control capability. | subsystem-requirements | Main |  | Test |  |  | Voice command is a convenience input and must not be a single point of failure for any surgical function. All capabilities accessible via voice must also be reachable via touchscreen or foot pedal, ensuring the surgeon retains complete system control through hardware-only input paths during voice system fault. | subsystem, surgeon-console, degraded-mode, session-353, idempotency:sub-sic-voice-degraded-353
SUB-MAIN-062 | The Haptic Controller SHALL be developed and validated to IEC 62061 Safety Integrity Level 2 (SIL2), including hardware fault tolerance requirements and safe-state transition to zero-torque output within 10ms of any detected CPU fault, EtherCAT timeout, or watchdog expiry. | subsystem-requirements | Main |  | Analysis |  |  | Direct patient-contact force pathway: a runaway torque command from a failed haptic controller would render uncontrolled forces on the surgeon's hands that could mask or amplify dangerous tissue contact events. SIL2 derives from FMEA on the haptic render loop — single CPU fault without safe-state fallback is classified as a hazardous event at severity level S2, probability class P2 under IEC 62061. | subsystem, haptic, session-354, idempotency:sub-haptic-sil2-354
SUB-MAIN-063 | The Haptic Feedback Subsystem SHALL remain stable (no limit cycles or sustained oscillations) during instrument-tissue contact for all tissue impedances within the range 0.1 N/mm to 10 N/mm and all motion scaling ratios 1:1 to 10:1. | subsystem-requirements | Main |  | Test |  |  | Haptic rendering loops coupled to stiff tissue models are susceptible to Z-width instability when the contact admittance exceeds the passivity boundary. The specified tissue stiffness range covers all soft to semi-rigid tissue types in abdominal surgery. Failure to maintain stability produces oscillating forces at the master handle that corrupt tactile perception and may cause the surgeon to lose grip control. | subsystem, haptic, session-354, idempotency:sub-haptic-stability-354
SUB-MAIN-064 | When the Haptic Feedback Subsystem is in STANDBY or DISABLED state, the Master Handle Actuator SHALL present a backdrive torque of no more than 0.05 Nm at any joint, ensuring force-transparent kinaesthetic operation independent of haptic rendering. | subsystem-requirements | Main |  | Test |  |  | Surgeons must be able to move master handles freely even when haptic rendering is inactive (e.g., during tele-operation setup or soft-tissue navigation without force feedback). Backdrive torque above 0.05 Nm creates proprioceptive masking, reducing tremor filtering effectiveness and increasing surgeon fatigue during procedures exceeding 2 hours. | subsystem, haptic, session-354, idempotency:sub-haptic-backdrive-354
SUB-MAIN-065 | The Haptic Feedback Subsystem SHALL be designed and manufactured in compliance with IEC 60601-1:2005+AMD1:2012 (medical electrical equipment safety) and IEC 60601-1-6 (usability), including classification as Type CF applied part where patient-contact signal pathways exist, and meeting continuous leakage current limits of ≤10µA under single fault conditions. | subsystem-requirements | Main |  | Analysis |  |  | The force sensing chain contacts the instrument tip which contacts the patient — this constitutes a Type CF applied part under IEC 60601-1. The 10µA limit at single fault is the most stringent patient leakage requirement and is mandated for devices with direct cardiac path access risk in intra-abdominal surgery. Compliance failure blocks FDA 510(k) clearance. | subsystem, haptic, compliance, session-354, idempotency:sub-haptic-compliance-60601-354
SUB-MAIN-066 | The Surgeon Interface Panel SHALL transmit 7-DOF Cartesian pose data from each master manipulator arm to the Console Computer at a minimum rate of 1kHz, with position resolution ≤0.1mm and angular resolution ≤0.1°, measured at the instrument-side tooltip for motion scaling computation. | subsystem-requirements | Main |  | Test |  |  | 1kHz sampling is the Nyquist floor for transparent 500Hz haptic control loops; lower rates introduce perceptible latency and quantisation artefacts in master arm motion. Position resolution ≤0.1mm matches the minimum clinically significant motion increment for microsurgical tasks per ISO 13482 guidance on medical robot kinematics. | subsystem, surgeon-console, haptic, session-356, idempotency:sub-sip-pose-rate-356
SUB-MAIN-067 | The Surgeon Interface Panel SHALL render haptic force feedback at the master arm fingertips within 1ms of receiving a force command from the Console Computer, with force rendering accuracy within 15 percent of commanded value across the 0 to 5N operating range. | subsystem-requirements | Main |  | Test |  |  | 1ms haptic loop closure is the perceptibility threshold for transparent force feedback; latencies above 5ms introduce phase lag that disrupts hand-eye coordination during delicate tissue manipulation. 15 percent accuracy matches the resolution of the force sensing at the instrument tip, ensuring the rendered feedback is not more accurate than the source measurement. | subsystem, surgeon-console, haptic, session-356, idempotency:sub-sip-haptic-force-356
SUB-MAIN-068 | The Surgeon Interface Panel SHALL detect surgeon hand disengagement from master arm handles within 50ms and immediately signal the Console Computer to inhibit motion transmission to patient-side instruments. | subsystem-requirements | Main |  | Test |  |  | Hand disengagement without inhibition would allow free-running slave arm motion with no surgeon intent, creating a direct patient injury hazard. 50ms is the upper bound derived from the system 1ms E-stop latency budget: sensor debounce 10ms plus signal propagation 40ms, ensuring disengagement is communicated before the 100ms safety watchdog deadline. | subsystem, surgeon-console, safety, session-356, idempotency:sub-sip-handle-engage-356
SUB-MAIN-069 | The Console Computer software SHALL be developed and documented to IEC 62304 Class C, with traceability from requirements to software units, and the overall system integration SHALL demonstrate conformance with IEC 60601-1-8 for alarm management and MDR 2017/745 Annex I essential safety requirements. | subsystem-requirements | Main |  | Inspection |  |  | The Console Computer is a Class C medical device software component under IEC 62304 because its failure mode can lead to patient injury without opportunity for intermediate detection. MDR 2017/745 Annex I essential requirements are mandatory for CE marking and apply to the complete robotic surgery system with the console as primary control surface. | subsystem, surgeon-console, compliance, regulatory, session-356, idempotency:sub-cc-compliance-iec62304-356
SUB-MAIN-070 | When the Console Computer detects a software exception or watchdog timeout, the system SHALL enter SAFE-HOLD state within 500ms, transmit a safe-state broadcast to all subsystems, and retain the last 30 seconds of kinematic and video data in battery-backed non-volatile storage for post-incident analysis. | subsystem-requirements | Main |  | Test |  |  | 500ms safe-state entry on console computer failure is within the clinical tolerance for motion freeze in surgical robotics before the surgeon can intervene manually. Retaining 30 seconds of pre-fault data supports root-cause analysis and is required under MDR 2017/745 for post-market surveillance of Class IIb devices. Battery-backed storage ensures retention even if main power is lost in the same fault event. | subsystem, surgeon-console, safety, redundancy, session-356, idempotency:sub-cc-failsafe-safehold-356
SUB-MAIN-071 | The Real-Time Protocol Engine SHALL execute time-division multiplexed frame scheduling with a cycle-to-cycle jitter of no more than 1 microsecond, measured at the frame start pulse on the inter-cart fibre transmitter. | subsystem-requirements | Main |  | Test |  |  | The 1kHz control loop has a total latency budget of 1ms end-to-end (SYS-MAIN-001). The FPGA frame scheduler is the first element in the surgeon-to-instrument chain; a jitter above 1 microsecond accumulates across 7 pipeline stages to exceed the allowed latency variance. Verified by oscilloscope capture of 10,000 consecutive frame-start pulses with a <2ns resolution clock source. | subsystem, motion-control, infrastructure, real-time-protocol-engine, session-357, idempotency:sub-rtpe-tdm-jitter-357
SUB-MAIN-072 | When the Real-Time Protocol Engine detects absence of a valid surgeon console frame for more than 3 consecutive TDM cycles (3ms), the Real-Time Protocol Engine SHALL assert a link-fault signal to the Workspace Safety Enforcer and transmit a zero-velocity command on all 6 DOF within 1ms of fault assertion. | subsystem-requirements | Main |  | Test |  |  | Three missed cycles at 1kHz equals 3ms, matching the system emergency-stop response budget in SYS-MAIN-010. A zero-velocity command within 1ms of fault assertion ensures the joint servos receive a valid safe command before the watchdog timer in the patient-side cart triggers a power-off, preventing uncontrolled arm movement during comms loss. | subsystem, motion-control, infrastructure, real-time-protocol-engine, safety, session-357, idempotency:sub-rtpe-fault-halt-357
SUB-MAIN-073 | The Network Management Controller SHALL maintain EtherCAT distributed clock synchronisation across all patient-side servo nodes with a maximum inter-node timing skew of 500 nanoseconds during continuous operation. | subsystem-requirements | Main |  | Test |  |  | Joint Servo Controllers on different patient-side cart nodes execute torque commands sampled from the same 1kHz kinematic frame. A timing skew above 500ns causes inter-joint command phase errors that produce coordinated arm jerks visible as position discontinuities. 500ns corresponds to one-quarter of the EtherCAT propagation jitter budget at 2Mbps, leaving margin for cable length variation. | subsystem, motion-control, infrastructure, network-management, session-357, idempotency:sub-nmc-ethercat-sync-357
SUB-MAIN-074 | When the Network Management Controller receives no valid EtherCAT response from a servo node for 2 consecutive bus cycles (2ms), the Network Management Controller SHALL remove the node from the active topology, assert a per-node fault flag to the Workspace Safety Enforcer, and maintain full communication with all remaining nodes within the same bus cycle. | subsystem-requirements | Main |  | Test |  |  | SYS-MAIN-016 requires continued full function on remaining instrument arms when one arm loses servo communication. Isolation within a single bus cycle (1ms) ensures the remaining nodes do not experience frame corruption due to a misbehaving slave, which would cause system-wide motion interruption rather than a contained single-arm fault. | subsystem, motion-control, infrastructure, network-management, safety, session-357, idempotency:sub-nmc-fault-isolation-357
SUB-MAIN-075 | The Procedure Data Recorder SHALL continuously record all 7-DOF joint angles, joint torques, and Cartesian velocities for each active instrument arm at 1kHz sample rate, with a maximum write latency of 10ms from sample acquisition to persistent storage. | subsystem-requirements | Main |  | Test |  |  | SYS-MAIN-015 mandates 1kHz kinematic data recording across all arms. A 10ms write latency ceiling (10 samples) matches the size of the in-memory ring buffer allocated on the Real-Time Compute Node before data is handed to the recorder. Exceeding 10ms risks ring buffer overflow during sustained peak I/O load, causing permanent sample loss that cannot be reconstructed for post-procedure audit. | subsystem, motion-control, infrastructure, procedure-data-recorder, session-357, idempotency:sub-pdr-recording-rate-357
SUB-MAIN-076 | The Procedure Data Recorder SHALL store data on dual-mirrored NVMe drives, detect any write or read failure on either drive within 1 second, and notify the Surgeon Interface Panel with an audible alarm while continuing to record on the remaining healthy drive without interruption. | subsystem-requirements | Main |  | Demonstration |  |  | Clinical and regulatory standards (IEC 62133, MDR Annex I General Safety clause 17.2) require that medical device data integrity is maintained during single-component storage failure. Dual-mirror NVMe provides hot-standby redundancy. A 1-second fault detection window is achievable via SMART polling at 100ms intervals and is the minimum acceptable for surgeon notification before a second fault could cause total data loss during a long procedure. | subsystem, motion-control, infrastructure, procedure-data-recorder, reliability, session-357, idempotency:sub-pdr-storage-integrity-357
SUB-MAIN-077 | The Inter-Cart Fibre Link SHALL provide a minimum bidirectional throughput of 10 Gbps with a one-way optical propagation latency not exceeding 100 microseconds for cable lengths up to 10 metres between Surgeon Console and Patient-Side Cart. | subsystem-requirements | Main |  | Test |  |  | The motion control pipeline requires approximately 2.4Mbps of kinematic data per arm (7 joints × 3 values × 64-bit × 1kHz), and the system supports up to 4 arms plus video metadata, totalling under 100Mbps even with control overhead. 10Gbps provides a 100× safety margin for future capability expansion. 100 microseconds propagation at 10m is well within the 1ms end-to-end latency budget of SYS-MAIN-001. | rt-implausible-value, red-team-session-502
SUB-MAIN-078 | When the active Inter-Cart Fibre Link channel fails (signal loss, CRC error rate exceeding 10^-9 per frame), the Inter-Cart Fibre Link SHALL switch to the redundant fibre channel within 1 millisecond without loss of any in-flight kinematic command frames. | subsystem-requirements | Main |  | Test |  |  | A 1ms failover aligns with the 3-missed-cycle fault threshold of the Real-Time Protocol Engine (SUB-MAIN-072), ensuring a channel switch does not itself trigger a controlled stop. Kinematic frame loss during failover would be indistinguishable from a comms fault and would unnecessarily halt the procedure. Dual-fibre switchover with in-flight frame buffering is standard in deterministic real-time networks (e.g., HSR/PRP protocols). | subsystem, motion-control, infrastructure, inter-cart-fibre, resilience, session-357, idempotency:sub-icfl-redundancy-failover-357
SUB-MAIN-079 | The Inter-Cart Fibre Link shall provide complete galvanic isolation between the Surgeon Console and Patient-Side Cart, with reinforced insulation rated to a minimum of 4000 VAC (IEC 60601-1 Clause 8.8), ensuring no conductive path exists between the two cart chassis. | subsystem-requirements | Main |  | Test |  |  | IEC 60601-1 type B applied part requirements prohibit any continuous conductive path between the surgeon side and the patient side of a surgical system. Optical fibre is the preferred isolation mechanism because it inherently provides infinite DC resistance. The 4000 VAC reinforced insulation rating is the IEC 60601-1 requirement for patient-contacting applied parts with highest risk classification, ensuring patient safety even under single-fault conditions in the mains power circuitry. | subsystem, motion-control, infrastructure, inter-cart-fibre, safety, iec60601, session-357, idempotency:sub-icfl-galvanic-isolation-357
SUB-MAIN-080 | The Tremor Rejection Filter SHALL implement a zero-phase 8th-order Butterworth low-pass filter at 6Hz cutoff, achieving ≥40dB attenuation above 6Hz and ≤0.5dB passband ripple below 3Hz, with an initial transient settling time of ≤5ms on mode activation. | subsystem-requirements | Main |  | Test |  |  | Involuntary physiological tremor spans 6-12Hz; the 6Hz cutoff preserves intentional surgical motion (typically 0-3Hz) while eliminating tremor. 8th-order provides the 40dB/octave slope needed without adding unacceptable group-delay. Zero-phase implementation prevents latency-induced instability. | subsystem, motion-control, tremor, session-358, idempotency:sub-tremor-filter-attenuation-358
SUB-MAIN-081 | When the Tremor Rejection Filter detects a sustained high-frequency velocity component above 8Hz for more than 200ms, it SHALL log a TREMOR_ELEVATED event to the Procedure Data Recorder and maintain filtering without operator intervention. | subsystem-requirements | Main |  | Test |  |  | Elevated tremor (e.g., fatigue, medication effect) changes filter operating point. Automatic logging captures surgeon physiology data for post-operative review and enables future adaptive filtering without requiring mid-procedure UI interaction. | subsystem, motion-control, tremor, session-358, idempotency:sub-tremor-adaptive-log-358
SUB-MAIN-082 | The Workspace Safety Enforcer SHALL compute signed penetration depth for all 7-DOF arm configurations against the patient anatomy mesh and instrument collision model at 1kHz, and SHALL generate a repulsive joint-space torque that limits Cartesian approach velocity to less than 5mm/s within 5mm of any restricted surface. | subsystem-requirements | Main |  | Test |  |  | Patient anatomy mesh defines no-go zones near vessels and organs. At 5mm proximity, a 5mm/s maximum approach rate gives the surgeon 1 second to redirect before potential tissue contact; at full 1kHz rate the enforcer can react within a single control cycle to prevent boundary violation. | subsystem, motion-control, workspace-safety, session-358, idempotency:sub-wse-proximity-enforcement-358
SUB-MAIN-083 | When the Workspace Safety Enforcer cannot access a valid patient anatomy mesh (model corruption or load failure), it SHALL transition the Motion Control and Scaling Subsystem to a reduced-mobility mode that limits Cartesian workspace to a predefined 150mm-radius sphere centred on the instrument tip position at fault onset, and SHALL generate a WORKSPACE_MODEL_FAULT alert within 50ms. | subsystem-requirements | Main |  | Test |  |  | Loss of anatomy mesh eliminates proximity safety guarantees. The 150mm sphere provides a safe enclosure around the current instrument position, preventing gross motion while allowing the surgeon to retract instruments. 50ms alert latency matches the safety watchdog cycle time. | subsystem, motion-control, workspace-safety, session-358, idempotency:sub-wse-degraded-mode-358
SUB-MAIN-084 | When the Kinematics Engine detects a kinematic singularity condition (Jacobian determinant below threshold 1e-4), it SHALL activate damped-least-squares inverse kinematics with a damping coefficient lambda of 0.05, and SHALL maintain commanded Cartesian velocity direction error below 5 degrees while limiting joint velocity to 80% of maximum. | subsystem-requirements | Main |  | Test |  |  | Surgical manipulators with 7-DOF pass near singularities during routine retraction and rotation. At singularity, standard pseudoinverse produces infinite joint velocities; DLS with lambda=0.05 caps joint speed while preserving directional intent. 5-degree directional error is within the surgeon motion resolution threshold. | subsystem, motion-control, kinematics, session-358, idempotency:sub-ke-singularity-handling-358
SUB-MAIN-085 | When the Joint Servo Controller detects a joint position error exceeding 0.5 degrees for more than 10ms during a commanded trajectory, it SHALL command the affected joint brake to engage within 2ms, halt motion on all joints of the affected arm, and report a SERVO_FAULT event to the Safe State Manager with joint ID, error magnitude, and timestamp. | subsystem-requirements | Main |  | Test |  |  | A 0.5-degree uncorrected error at the instrument tip can translate to 3-5mm tip displacement depending on arm configuration; at 10ms detection window this stays within the 2mm maximum allowable tip error envelope. Brake engage within 2ms prevents runaway while reporting enables root-cause analysis. | subsystem, motion-control, servo, session-358, idempotency:sub-jsc-fault-isolation-358
SUB-MAIN-086 | The Procedure Data Recorder SHALL be housed in a dedicated 1U rack-mount enclosure (430mm × 44mm × 380mm) integrated into the patient cart, rated IP21, operating from the 12V patient cart rail, with WORM-compliant SSD storage and a removable front-panel USB-C port for intraoperative data export. | subsystem-requirements | Main |  | Inspection |  |  | Physical embodiment requirement derived from ontological analysis showing PDR has physical environmental constraints (SUB-MAIN-086). Dedicated enclosure isolates PDR from vibration and RF emitted by energy delivery subsystem; IP21 matches clinical environment; rack-mount integration mandated by cart geometry and field-service requirements.
SUB-MAIN-086 | While the Real-Time Compute Node CPU junction temperature exceeds 85 degrees Celsius, the node SHALL reduce non-critical background processing priority and maintain full deterministic scheduling for Motion Control threads, with junction temperature logged to the Procedure Data Recorder at 1Hz; if temperature reaches 95 degrees Celsius, the node SHALL generate a THERMAL_CRITICAL alert and transition to Safe Hold. | subsystem-requirements | Main |  | Test |  |  | Real-time scheduling must be immune to thermal throttling; standard OS thermal management interrupts deterministic task execution. 85C threshold provides 10C headroom before critical shutdown. Safe Hold at 95C prevents silicon damage that could cause unpredictable motion. | subsystem, motion-control, compute, session-358, idempotency:sub-rtcn-thermal-management-358
SUB-MAIN-087 | The Power Management Subsystem SHALL be implemented as two physically separated LRUs: a mains power entry module (400mm × 200mm × 100mm) mounted in the patient cart base handling AC-DC conversion for 24V, 12V, and 5V rails, and a surgeon console power board (250mm × 150mm × 60mm) supplying the console computer and haptic subsystem; both rated IP21 and conformal-coated for cleaning agent resistance. | subsystem-requirements | Main |  | Inspection |  |  | Physical embodiment required by ontological mismatch finding (SUB-MAIN-102, VER-MAIN-107): PMS has physical environmental constraints but was classified without Physical Object trait. Separation into two LRUs reflects actual surgical robot architectures where console and cart power domains are galvanically isolated for patient safety per IEC 60601-1 clause 8.
SUB-MAIN-087 | The Motion Control and Scaling Subsystem SHALL authenticate all external configuration commands (scaling ratio updates, workspace model loads, and motion enable/disable) via HMAC-SHA256 with a session key established at system startup, and SHALL reject and log any command failing authentication within one control cycle (1ms) without interrupting the real-time motion pipeline. | subsystem-requirements | Main |  | Test |  |  | IEC 62443-3-3 SR 1.1 requires authentication for all control commands in medical device systems. Unauthorised scaling ratio or workspace mesh injection could cause unsafe instrument motion; HMAC-SHA256 with a session key provides integrity without adding per-cycle cryptographic load to the deterministic pipeline. Rejection within 1ms ensures the safety check does not degrade pipeline timing. | subsystem, motion-control, cybersecurity, session-358, idempotency:sub-mcs-command-auth-358
SUB-MAIN-088 | The Motion Control System SHALL be implemented as a physically distinct compute subsystem comprising a real-time motion controller card (PCIe half-length, 210mm × 111mm) installed in the patient cart chassis, with direct hardware backplane connections to actuator drive electronics, operating from the 12V patient cart rail and rated for continuous operation at 45°C ambient without forced air cooling. | subsystem-requirements | Main |  | Inspection |  |  | Physical embodiment needed per lint finding: MCS lacks Physical Object trait but SUB-MAIN-086 imposes physical environmental constraints. Real-time constraint requires dedicated hardware, not virtualised compute; 45°C ambient without forced cooling is derived from operating theatre HVAC limits and OR noise requirements (IEC 60601-1 Part 12 compatibility).
SUB-MAIN-088 | The Main Power Distribution Unit SHALL distribute mains AC power to all surgical robot subsystems via independently fused branch circuits, with each branch rated at no less than 125% of its maximum load current. | subsystem-requirements | Main |  | Test |  |  | Branch circuit over-sizing to 125% of peak load current is required by IEC 60601-1 clause 10.2 for medical electrical equipment, preventing nuisance trips during startup inrush while ensuring branch faults do not propagate to adjacent subsystems via shared conductors. | subsystem, power-management, power-distribution, session-361, idempotency:sub-pdu-branch-circuits-361
SUB-MAIN-089 | The Time Compute Node SHALL be a dedicated timing hardware module (70mm × 45mm, M.2 form factor) installed on the system backplane, providing IEEE 1588v2 PTP grandmaster function with internal TCXO reference oscillator, GPS-disciplined timing input, and hardware timestamping to within ±500ns for all inter-subsystem data frames. | subsystem-requirements | Main |  | Test |  |  | Physical embodiment needed per lint finding: TCN lacks Physical Object trait but SUB-MAIN-086 imposes physical environmental constraints. Dedicated timing hardware (rather than software PTP stack) provides deterministic ±500ns accuracy required for synchronised sensor fusion; TCXO reference maintains accuracy during GPS signal loss in RF-shielded operating theatres.
SUB-MAIN-089 | The Main Power Distribution Unit SHALL detect line-to-earth leakage current exceeding 500 µA on any branch and remove power from that branch within 100 ms, transmitting a fault code to the Power Sequencing Controller via the internal CAN bus. | subsystem-requirements | Main |  | Test |  |  | IEC 60601-1 clause 8.7.3 limits patient-accessible leakage current to 500 µA in normal condition and to 1 mA in single fault. The 100 ms response time ensures the Safety and Interlock Subsystem can initiate protective shutdown before accumulated charge reaches dangerous thresholds, given the capacitance of typical OR wiring harnesses. | subsystem, power-management, safety, iec60601, session-361, idempotency:sub-pdu-ground-fault-361
SUB-MAIN-090 | The Motion Control System SHALL be developed and validated in conformance with IEC 62304:2006 (Medical device software lifecycle) at Software Safety Class C, and shall satisfy IEC 60601-1:2005+A1:2012 clause 14 (Programmable electrical medical systems) with a certified PEMS development file submitted to the notified body. | subsystem-requirements | Main |  | Analysis |  |  | MCS is classified as Regulated (UHT hex 51F73A18) but had no compliance requirements. Class C classification is required because MCS failure can cause serious injury or death (loss of motion control during surgery). IEC 62304 mandates software lifecycle documentation; IEC 60601-1 PEMS requirements apply to all software-controlled medical devices in EU MDR scope.
SUB-MAIN-090 | The UPS Battery Module SHALL report state-of-charge to the Power Sequencing Controller at 1 Hz via the battery management system interface, with accuracy of ±2% across the 20–95% charge range, and SHALL assert a low-battery warning when state-of-charge falls below 25%. | subsystem-requirements | Main |  | Test |  |  | Accurate state-of-charge telemetry at 1 Hz allows the Power Sequencing Controller to trigger an orderly shutdown before battery depletion causes an uncontrolled power loss during surgery. The 25% warning threshold provides at minimum 8 minutes of bridging time (from SUB-MAIN-043 minimum 10-minute UPS duration) for an orderly instrument retraction procedure. | subsystem, power-management, ups, reliability, session-361, idempotency:sub-ups-soc-reporting-361
SUB-MAIN-091 | The Workspace Safety Enforcer SHALL be designed, implemented, and independently verified to achieve Safety Integrity Level 3 (SIL 3) per IEC 61508-1:2010, with a target hardware fault tolerance of HFT ≥ 1, probabilistic failure to perform on demand of <10⁻⁷ per hour, and third-party functional safety assessment prior to clinical deployment. | subsystem-requirements | Main |  | Analysis |  |  | WSE is classified as Regulated (UHT hex 51B73818) and is the primary safety boundary preventing instrument collision with anatomy outside the surgical workspace. SIL 3 is required for functions where failure can result in serious irreversible injury; HFT ≥ 1 means no single hardware failure can defeat the safety function.
SUB-MAIN-091 | When mains supply voltage falls below 80% of nominal for more than 20 ms, the Power Sequencing Controller SHALL initiate transfer to UPS Battery Module within 10 ms, ensuring no subsystem experiences a supply interruption exceeding 30 ms total. | subsystem-requirements | Main |  | Test |  |  | The 30 ms total transfer budget is derived from joint servo control loop requirements: a 30 ms power interruption is the maximum before the joint controllers enter fault state (losing current position data), which would require a full arm re-homing sequence and loss of the sterile field. The 80% voltage threshold prevents false triggers from short-duration sags common in hospital OR environments. | subsystem, power-management, ups, failover, safety, session-361, idempotency:sub-psc-mains-loss-transfer-361
SUB-MAIN-092 | The Motion Scaling Module SHALL be developed under a Design History File (DHF) as required by 21 CFR Part 820.30 and EU MDR Annex II, shall implement scaling algorithms in a verified, MISRA C:2012 compliant codebase, and shall undergo human factors validation per IEC 62366-1 to confirm surgeons can accurately select and apply scaling ratios from 1:1 to 10:1 without error. | subsystem-requirements | Main |  | Analysis |  |  | MSM is classified as Regulated (UHT hex 50B53B18) with no compliance requirements. DHF requirement satisfies FDA QSR for Class II/III devices; MISRA C compliance prevents safety-critical coding errors in scaling computation; IEC 62366-1 human factors validation is required because erroneous scaling selection (surgeon input error) is a foreseeable use error hazard.
SUB-MAIN-092 | The Auxiliary Power Supply SHALL maintain 24 VDC output within plus or minus 2% under load variations from 0 to 100% of rated capacity, and SHALL remain energised for a minimum of 20 minutes following complete loss of mains and UPS supply, sourced from a dedicated internal battery. | subsystem-requirements | Main |  | Test |  |  | The 24 VDC auxiliary rail exclusively powers safety circuits: the Safe State Manager, Emergency Stop Chain, and Watchdog Timer Controller. These circuits must remain active after main UPS depletion to enable controlled arm retraction and safe-state assertion. 20 minutes exceeds the longest documented emergency surgical exit procedure (12 minutes per clinical operations data) with 8 minutes margin. | subsystem, power-management, auxiliary-power, safety, session-361, idempotency:sub-aux-psu-24v-361
SUB-MAIN-093 | The Power Management Subsystem SHALL provide patient-applied part (Type B minimum, Type BF preferred) galvanic isolation meeting IEC 60601-1:2005 clause 8 with patient leakage current <10µA (normal condition) and <50µA (single fault condition), and shall comply with IEC 60601-1-2:2014 (EMC) for conducted and radiated emissions in a clinical environment. | subsystem-requirements | Main |  | Test |  |  | PMS is classified as Regulated (UHT hex 54F53018) but had no compliance requirements. IEC 60601-1 isolation is mandatory for all mains-powered medical devices. Leakage thresholds prevent microshock hazard to patients; EMC compliance prevents PMS switching noise from interfering with electrophysiology monitoring equipment co-located in the OR.
SUB-MAIN-093 | The Safety and Interlock Subsystem SHALL be designed, implemented, and verified to achieve Safety Integrity Level 3 (SIL 3) per IEC 62061, with a probability of dangerous failure per hour (PFH) not exceeding 1×10⁻⁷/h for each safety function (joint force limiting, E-stop chain, communication watchdog). | subsystem-requirements | Main |  | Analysis |  |  | A surgical robot operating on an anaesthetised patient is a Class IIb/III medical device; safety functions that prevent uncontrolled arm motion or tissue penetration must meet SIL 3 under IEC 62061. PFH ≤1×10⁻⁷/h is the SIL 3 threshold. Failure to meet this leaves the safety case without regulatory acceptance. | subsystem, sis, compliance, sil3, session-362, idempotency:sub-sis-sil3-compliance-362
SUB-MAIN-094 | When the primary Procedure Data Recorder storage medium fails during an active surgical procedure, the system SHALL automatically failover to a secondary redundant recording path within 500ms, maintaining all video, telemetry, and event log streams with no data loss; the surgeon display SHALL indicate recording medium degradation within 1 second of failover. | subsystem-requirements | Main |  | Test |  |  | PDR is System-Essential (UHT hex 50851208): procedure recordings are required for post-surgical review, complication investigation, and regulatory audit. Without recording continuity, any intraoperative event loses its evidential record. 500ms failover allows one frame of 4K video to be dropped but ensures continuous recording; surgical data continuity is mandated by EU MDR Article 83 (UDI) and hospital governance policies.
SUB-MAIN-094 | The Motion Control System software SHALL be developed and qualified under IEC 62304 Safety Class C, with full software lifecycle documentation including software development plan, software requirements, detailed design, unit and integration test records, and risk analysis. | subsystem-requirements | Main |  | Analysis |  |  | Motion control software executes joint servo commands in real time for a Class IIb medical device; IEC 62304 Class C applies when software failure can cause death or serious injury. Absence of IEC 62304 classification prevents regulatory submission and creates unmanaged software risk. | subsystem, motion-control, compliance, iec62304, session-362, idempotency:sub-mcs-iec62304-classC-362
SUB-MAIN-095 | When the primary Time Protocol Engine TCXO reference or GPS disciplining input fails, the system SHALL maintain inter-subsystem synchronisation to within ±2µs holdover accuracy for a minimum of 30 minutes using the secondary crystal oscillator, and SHALL log the timing source degradation event with timestamp and subsystem notification within 100ms of fault detection. | subsystem-requirements | Main |  | Test |  |  | TPE is System-Essential (UHT hex 50B57B08): loss of time synchronisation degrades sensor fusion accuracy, invalidates haptic feedback timing, and breaks inter-cart communication framing. 30-minute holdover covers typical OR interruption events (power blip, GPS shielding). ±2µs holdover bound is derived from maximum acceptable jitter in the 1kHz haptic control loop (1µs phase error is within 0.1% of sample period).
SUB-MAIN-095 | The Console Computer SHALL detect its own software watchdog failure within 500ms and automatically transfer console control authority to the backup processing path, preserving the surgeon's last commanded instrument position and annunciating the transfer via an audio tone and on-screen status update. | subsystem-requirements | Main |  | Test |  |  | The Console Computer is the single physical host for the surgeon interface; a software hang without automatic recovery requires manual intervention during surgery, creating a potential patient safety event. 500ms timeout is consistent with the 150ms E-stop latency requirement — sufficient margin for detection before motion state becomes stale. | subsystem, surgeon-console, redundancy, failover, session-362, idempotency:sub-cc-watchdog-failover-362
SUB-MAIN-096 | When the Haptic Feedback Subsystem force-feedback actuator channel for any master manipulator axis fails, the system SHALL continue operating in degraded haptic mode, suppressing force feedback for the affected axis only, displaying a force feedback degraded warning to the surgeon within 200ms, and maintaining full motion control capability for all operational axes at the nominal 1kHz servo rate. | subsystem-requirements | Main |  | Test |  |  | HFS is System-Essential (UHT hex 55F57018): complete haptic loss prevents the surgeon from detecting tissue contact forces, risking inadvertent perforation. Per-axis degraded mode (rather than full shutdown) preserves surgical capability for the critical axis while alerting to reduced fidelity. 200ms notification latency matches the surgeon's tactile attention refresh rate; 1kHz servo rate preservation ensures motion precision is not compromised.
SUB-MAIN-096 | When the Haptic Controller loses communication with the Force Sensing Module for more than 50ms, the Haptic Feedback Subsystem SHALL enter a force-blind degraded mode: the Master Handle Actuator SHALL apply a constant 0.3N braking force to all degrees of freedom, audible and visual alerts SHALL be activated, and the surgeon SHALL retain full kinematic control of instrument motion. | subsystem-requirements | Main |  | Test |  |  | Complete loss of force feedback creates the risk of over-insertion or excessive tissue force without surgeon awareness. A 0.3N constant braking force (below the 5N tissue-force limit from SYS-MAIN-012) provides passive cue that force feedback is impaired without preventing emergency instrument withdrawal. 50ms threshold is consistent with the 15ms haptic render rate — three missed render cycles triggers degraded mode. | subsystem, haptic, redundancy, degraded-mode, session-362, idempotency:sub-haptic-fbd-degraded-362
SUB-MAIN-097 | The Power Management Subsystem SHALL provide redundant power paths for the Safety-Critical power domain (interlock subsystem, workspace safety enforcer, and emergency stop circuits): two independent DC supply rails each capable of sustaining full safety function load, with automatic switchover within 5ms of primary rail failure, supporting at least 60 seconds of safe-state hold from the onboard UPS battery. | subsystem-requirements | Main |  | Test |  |  | PMS is System-Essential (UHT hex 54F53018): safety function power loss without backup results in uncontrolled actuator freewheel during surgery. Dual-rail architecture prevents single power supply failure from defeating the safety function. 5ms switchover is derived from the 250ms safe-state achievement budget (SYS-MAIN-002) — power switchover must not consume more than 2% of the fault budget.
SUB-MAIN-097 | The Motion Control and Scaling Subsystem command interfaces (Trajectory Generator inputs, Motion Scaling Module configuration parameters) SHALL authenticate all incoming command frames using a session-keyed HMAC-SHA256 message authentication code; any frame failing authentication SHALL be rejected and a security violation event logged to the Procedure Data Recorder within 10ms. | subsystem-requirements | Main |  | Test |  |  | The Trajectory Generator and Motion Scaling Module accept commands that directly translate to physical arm motion. A spoofed or replayed command could cause unintended motion during surgery. HMAC-SHA256 provides integrity protection without the latency overhead of asymmetric encryption; session-key binding prevents replay across power cycles. IEC 80001-1 and the FDA cybersecurity guidance for networked medical devices both require integrity controls on safety-critical command paths. | subsystem, motion-control, cybersecurity, authentication, session-362, idempotency:sub-mcs-command-hmac-362
SUB-MAIN-098 | The Communication and Data Management Subsystem SHALL implement HMAC-SHA-256 message authentication on all safety-critical command channels (motion control, energy delivery, workspace safety enforcer), with per-message authentication tags verified by the receiving subsystem within one control cycle (≤1ms), and SHALL reject and log any command with an invalid or missing authentication tag, triggering a safe-state transition. | subsystem-requirements | Main |  | Test |  |  | Subsystem decomposition of SYS-MAIN-018: CDMS is the inter-subsystem communication backbone and is therefore the natural implementation locus for command authentication. HMAC-SHA-256 is selected as it is computationally feasible within a 1ms control cycle on embedded hardware while providing 128-bit security level. Per-message authentication prevents replay and injection attacks identified in FDA 2023 cybersecurity guidance for surgical robots.
SUB-MAIN-098 | The Tool Tip Articulation Controller SHALL authenticate all incoming joint-space position commands using a 32-bit HMAC signature, rejecting any command with an invalid or missing signature within one control cycle (1ms), and logging all rejected commands to the Procedure Data Recorder. | subsystem-requirements | Main |  | Test |  |  | IEC 62443-4-2 requires authentication of all command inputs to safety-critical embedded controllers. The TTAC directly drives distal DOFs; an unauthenticated command injected via a compromised Instrument Drive Unit bus could cause unexpected instrument motion at the tissue interface. 32-bit HMAC at 1kHz control rate is computationally feasible on the TTAC compute board and matches the authentication scheme applied to Kinematics Engine, Trajectory Generator, and Real-Time Protocol Engine. | subsystem, surgical-instrument-system, cybersecurity, authentication, session-365, idempotency:sub-ttac-cybersecurity-auth-365
SUB-MAIN-099 | The Vision and Imaging Subsystem SHALL deliver stereoscopic 3D video at 1080p per eye at 60Hz to the surgeon console display with end-to-end latency from endoscope tip to surgeon display of <50ms under nominal operating conditions, and SHALL maintain continuous stereo video output with no more than one dropped frame per 10 seconds during any intraoperative phase. | subsystem-requirements | Main |  | Test |  |  | Subsystem decomposition of STK-MAIN-012 (surgeon situational awareness) and SYS-MAIN-003 (stereoscopic HD video). <50ms latency threshold is derived from human perception: beyond 50ms surgeons report motion sickness in VR-coupled tasks; frame-drop continuity specification ensures no interruption of tissue visualisation during critical dissection phases.
SUB-MAIN-099 | The Kinematics Engine SHALL authenticate all joint-space command packets received from the Motion Control System using a 32-bit HMAC-SHA256 keyed with a session key negotiated at startup, rejecting any packet with an invalid or missing signature within one control cycle. | subsystem-requirements | Main |  | Test |  |  | IEC 62443-4-2 requires input authentication for safety-critical control software. Unauthenticated Cartesian commands injected at the inter-subsystem interface could be used to drive arm joints beyond safe workspace limits. Session-key HMAC provides both authentication and replay protection without requiring persistent key storage on the RTCN. | subsystem, motion-control, cybersecurity, session-365, idempotency:sub-ke-auth-iec62443-365
SUB-MAIN-100 | The Trajectory Generator SHALL validate all motion waypoints against the active workspace safety envelope before execution, rejecting any waypoint that would place an instrument tip within 5mm of a registered keep-out zone boundary, and halting trajectory execution with a safe-state transition within 5ms of detection. | subsystem-requirements | Main |  | Test |  |  | Trajectory-level validation provides defence-in-depth upstream of the Joint Servo Controller. Rejecting unsafe waypoints at the Trajectory Generator prevents the Workspace Safety Enforcer from being the sole line of defence against workspace boundary violations, meeting IEC 80601-2-77 requirement for layered safety barriers in surgical robots. | subsystem, motion-control, workspace-safety, session-365, idempotency:sub-tg-waypoint-validation-365
SUB-MAIN-101 | The Real-Time Protocol Engine SHALL authenticate all synchronisation frames received on the inter-cart fibre link using a 16-bit CRC combined with a 32-bit session token, discarding any frame that fails authentication and triggering a communication fault event within one synchronisation period. | subsystem-requirements | Main |  | Test |  |  | RTPE processes all time-critical synchronisation traffic between Surgeon Console and Patient-Side Cart at 1kHz. An injected or replayed synchronisation frame could desynchronise the master-slave control loop, causing latency spikes or control instability. Frame-level authentication is consistent with IEC 62443 requirements and matches the HMAC scheme applied to KE and TG. | subsystem, infrastructure, cybersecurity, session-365, idempotency:sub-rtpe-auth-sync-365
SUB-MAIN-102 | The Workspace Safety Enforcer SHALL be designed and verified to IEC 60601-1:2005+A1:2012 and ISO 14971:2019, achieving Classification III risk acceptability for joint workspace limit enforcement, with documented residual risk below the acceptable limit defined in the system risk management file. | subsystem-requirements | Main |  | Analysis |  |  | Workspace Safety Enforcer is Regulated per UHT classification (51B73818). Enforces hard joint limits whose violation risks direct patient injury. IEC 60601-1 clause 14 and ISO 14971 are legally mandated under MDR 2017/745. This requirement establishes the certification basis for the workspace limit function.
SUB-MAIN-102 | The Power Management Subsystem SHALL comply with IEC 60601-1:2005+AMD1:2012 (Medical electrical equipment — General requirements for basic safety and essential performance) and IEC 60601-1-2:2014 (electromagnetic compatibility), with applied part classification F-Type for all outputs connected to patient-coupled circuits, achieving 500 VAC dielectric withstand between mains and patient outputs. | subsystem-requirements | Main |  | Test |  |  | IEC 60601-1 compliance is mandatory for medical electrical equipment connected to or supplying patient-coupled circuits. F-Type floating isolation is required because the PDU supplies the patient-side cart, which may have instruments in contact with the patient; a ground fault on the mains side must not cause patient injury. The 500 VAC withstand is the IEC 60601-1 Table 2 value for F-Type parts at 230V mains. | subsystem, power-management, compliance, iec60601, session-368, idempotency:sub-pms-iec60601-compliance-368
SUB-MAIN-103 | The Communication and Data Management System SHALL implement HMAC-SHA256 message authentication on all safety-critical inter-subsystem command interfaces over the inter-cart fibre link, verifying each command frame before forwarding to motion control, with authentication failure causing immediate command rejection and SAFE-HOLD transition within 50ms. | subsystem-requirements | Main |  | Test |  |  | SYS-MAIN-018 mandates cryptographic message authentication on safety-critical interfaces. HMAC-SHA256 selected over AES-GCM because it has lower compute latency (<10 microseconds on embedded DSP) while providing equivalent message authentication strength for the 1kHz control frame rate. Failure to authenticate must trigger SAFE-HOLD rather than continue — any accepted spoofed or corrupted command could cause uncontrolled arm motion.
SUB-MAIN-103 | The Inter-Cart Fibre Link SHALL provide a minimum sustained one-way latency of ≤500µs for all real-time kinematic command frames under peak traffic load (21 kinematics channels at 1kHz simultaneously with dual stereo HD video streams). | subsystem-requirements | Main |  | Test |  |  | The 500µs one-way latency budget for the fibre link is derived from the 1ms end-to-end control loop requirement in SYS-MAIN-001. With 250µs allocated to outbound processing and 250µs for return path encoding, 500µs is the maximum permissible fibre path delay. Exceeding this would violate the haptic feedback realism threshold of <1ms round-trip, causing perceptible surgeon-to-instrument lag during tissue manipulation. | subsystem, comms, inter-cart-fibre, session-369, idempotency:sub-icfl-latency-369
SUB-MAIN-104 | The Inter-Cart Fibre Link SHALL switch from the primary to the redundant fibre path within 5ms of primary link failure detection, with no loss of kinematic command continuity beyond the within the 100ms safety timeout window. | subsystem-requirements | Main |  | Test |  |  | 5ms failover is derived from SUB-MAIN-028 which requires link failure detection within 5ms and transition to safe state. A 5ms failover without command loss is achievable with hot-standby redundancy at FPGA-level path switching; a longer switchover would require the Safety and Interlock Subsystem to initiate arm park, disrupting the procedure unnecessarily. | subsystem, comms, inter-cart-fibre, redundancy, session-369, idempotency:sub-icfl-failover-369
SUB-MAIN-105 | The Real-Time Protocol Engine SHALL encapsulate each 1kHz kinematic command frame with a 32-bit sequence number, HMAC-SHA256 authentication tag, and CRC-32 error detection field within a total frame overhead of 64 bytes or less. | subsystem-requirements | Main |  | Analysis |  |  | Frame overhead must be bounded so that the 1kHz cycle at 21 joint channels fits within the 1Gbps link budget. HMAC-SHA256 per frame satisfies SYS-MAIN-018 cryptographic authentication. CRC-32 provides independent error detection; HMAC alone is not sufficient for in-flight bit error detection due to compute latency. 64-byte overhead yields <7% protocol overhead at peak data load. | subsystem, comms, real-time-protocol-engine, session-369, idempotency:sub-rtpe-framing-369
SUB-MAIN-106 | The Real-Time Protocol Engine SHALL detect frame loss or sequence discontinuity within one 1kHz frame period (1ms) and report a COMM_FAULT event to the Safety and Interlock Subsystem within 2ms of detection. | subsystem-requirements | Main |  | Test |  |  | 1ms detection aligns with the 1kHz kinematic control cycle: a missed frame is visible at the next expected sequence number. The 2ms reporting budget allows one additional frame cycle for COMM_FAULT generation, keeping total fault propagation under 3ms, well within the 100ms safety timeout. Delayed detection would allow incorrect kinematics to accumulate before safe-state is initiated. | subsystem, comms, real-time-protocol-engine, fault-handling, session-369, idempotency:sub-rtpe-fault-detect-369
SUB-MAIN-107 | The Network Management Controller SHALL monitor both primary and secondary fibre link health at 100Hz and classify each link as HEALTHY, DEGRADED, or FAILED based on bit error rate, frame loss ratio, and inter-frame jitter thresholds. | subsystem-requirements | Main |  | Test |  |  | 100Hz monitoring provides 10ms resolution on link degradation events, sufficient to initiate failover before the 100ms safety timeout. Three-state classification allows the system to pre-arm failover on DEGRADED links before outright failure, reducing unexpected switchover events during procedures. | subsystem, comms, network-management, session-369, idempotency:sub-nmc-health-monitor-369
SUB-MAIN-108 | The Network Management Controller SHALL implement strict-priority queuing with three traffic classes: SAFETY (heartbeats and E-stop, highest priority), KINEMATIC (1kHz command frames), and DATA (video and procedure recording, lowest priority), ensuring SAFETY and KINEMATIC frames are never delayed by DATA traffic. | subsystem-requirements | Main |  | Analysis |  |  | Priority queuing is required because safety heartbeats and kinematic commands have hard real-time deadlines (5ms and 1ms respectively) while video and recording data have soft deadlines. Without traffic class isolation, burst video traffic could compete for bandwidth with safety-critical frames, introducing jitter that violates SYS-MAIN-001 and SYS-MAIN-002. | subsystem, comms, network-management, session-369, idempotency:sub-nmc-qos-369
SUB-MAIN-109 | The Procedure Data Recorder SHALL sustain a continuous write throughput of at least 2 GB/s to NVMe RAID storage while simultaneously recording 21-channel kinematics at 1kHz, dual stereo HD video at 60fps, and timestamped system events without buffer overflow or dropped frames for a minimum 8-hour procedure duration. | subsystem-requirements | Main |  | Test |  |  | 2 GB/s throughput is derived from peak data rates: 21 joints x 3 values x 2 bytes x 1000Hz = ~126 KB/s kinematics; 2x1080p60 H.264 at 50Mbps = ~12.5 MB/s video; event log negligible. Total ~13 MB/s sustained, with 2 GB/s NVMe capacity providing 150x headroom for burst and redundancy writes. SYS-MAIN-015 requires 8-hour continuous recording. | subsystem, comms, procedure-data-recorder, session-369, idempotency:sub-pdr-throughput-369
SUB-MAIN-110 | The Procedure Data Recorder SHALL generate a SHA-256 hash of the complete procedure dataset at end-of-procedure and store it alongside the data, enabling post-hoc integrity verification of all recorded kinematic, video, and event data. | subsystem-requirements | Main |  | Test |  |  | Cryptographic hash verification satisfies regulatory requirements for surgical record integrity (IEC 62304 and FDA 21 CFR Part 11 audit trail obligations). Without hash verification, undetected storage corruption could render recorded data inadmissible for post-procedural review or incident investigation. | subsystem, comms, procedure-data-recorder, session-369, idempotency:sub-pdr-integrity-369
SUB-MAIN-111 | The Surgical Instrument System patient-side components (Sterile Adapter, Surgical Instruments, and Cable Tensioning System) SHALL withstand full fluid immersion in IPA 70% and quaternary ammonium compound disinfectants for minimum 30-minute contact time without degradation of structural integrity, sterile barrier, or instrument articulation performance. | subsystem-requirements | Main |  | Test |  |  | SYS-MAIN-006 mandates IPA 70% compatibility for all sterile-field components. The Surgical Instrument System houses the cable-driven wrist mechanism, sterile adapters, and instrument channels that routinely contact surgical irrigation and disinfectants. Failure to resist IPA 70% immersion risks stress-cracking of polymer housings, corrosion of cable sheaths, and loss of sterile barrier — potential patient infection pathways. | subsystem, surgical-instrument, sterility, session-373, idempotency:sub-sis-ipa-disinfectant-373
SUB-MAIN-112 | The Procedure Data Recorder SHALL be implemented as a rack-mounted 2U line-replaceable unit (LRU) housed within the Vision Cart, with RAID-1 mirrored storage of minimum 2TB, compliant with IP32 ingress protection to tolerate operating-room fluid splash, and a front-panel status LED visible from 3 metres. | subsystem-requirements | Main |  | Inspection |  |  | The Procedure Data Recorder stores complete procedural data streams that may be used for post-operative review, regulatory audit, and litigation evidence. Physical embodiment as a 2U rack LRU enables field replacement without cart teardown. IP32 protects against OR fluid management. RAID-1 ensures no data loss on single-drive failure. | subsystem, comms, pdr, physical, session-374, idempotency:sub-pdr-physical-embodiment-374
SUB-MAIN-113 | The Procedure Data Recorder SHALL be a rack-mounted 2U line-replaceable unit in the Vision Cart with 2TB RAID-1 mirrored storage and IP32 ingress protection. | subsystem-requirements | Main |  | Inspection |  |  | Defined as LRU for field replacement without Vision Cart teardown. IP32 protects against OR fluid splash. RAID-1 ensures no data loss on single-drive failure during procedure. | rt-mechanical-trace, red-team-session-502
SUB-MAIN-114 | The Power Management Subsystem SHALL be housed in a dedicated sealed electronics bay within the Patient-Side Cart, with the Main Power Distribution Unit and UPS Battery Module installed as field-replaceable assemblies, accessible via a locked service panel requiring IEC 62353 isolation verification before access. | subsystem-requirements | Main |  | Inspection |  |  | Power management handles 240V AC mains and high-capacity battery systems. Physical enclosure in a sealed bay with locked access prevents inadvertent contact with live voltages in the sterile surgical environment. Field-replaceable assemblies reduce mean-time-to-repair and avoid full cart downtime. | subsystem, power-management, physical, session-374, idempotency:sub-pms-physical-embodiment-374
SUB-MAIN-115 | The Real-Time Compute Node SHALL be packaged as a single-board computing module conforming to VITA 57.1 FMC form factor, installed in a vibration-isolated slot within the Patient-Side Cart electronics bay, with conformal coating for condensation protection and operational temperature range of 0 to 50 degrees Celsius. | subsystem-requirements | Main |  | Inspection |  |  | VITA 57.1 FMC standardises the physical interface and enables like-for-like replacement for RTCN variants without cart redesign. Conformal coating addresses OR humidity. The 0-50C range covers cold storage to a warm operating room, which is the real thermal envelope for medical cart electronics. | subsystem, motion-control, physical, rtcn, session-374, idempotency:sub-rtcn-physical-embodiment-374
SUB-MAIN-116 | The Motion Control System SHALL be implemented as a distributed electronics assembly across the Patient-Side Cart with the Real-Time Compute Node, Trajectory Generator, and Motion Scaling Module co-located in the same vibration-isolated electronics chassis to minimise inter-module latency, with individual modules accessible for replacement without full cart disassembly. | subsystem-requirements | Main |  | Inspection |  |  | Co-location of the RTCN, Trajectory Generator, and Motion Scaling Module on a common backplane eliminates inter-chassis cabling latency, keeping the computation pipeline within the 1ms end-to-end cycle budget. Physical modularity enables field repair without returning the full cart to depot, reducing surgical suite downtime. | subsystem, motion-control, physical, session-374, idempotency:sub-mcs-physical-embodiment-374
SUB-MAIN-117 | The Power Management Subsystem SHALL comply with IEC 60601-1:2005+AMD1:2012 medical electrical equipment safety standard, with dielectric withstand testing at 4000V AC applied between primary circuit and patient-accessible parts, and leakage current not exceeding 500 microamperes under normal condition and 1000 microamperes under single-fault condition. | subsystem-requirements | Main |  | Test |  |  | IEC 60601-1 is mandatory for medical electrical equipment in all major markets (EU MDR, FDA 510k). The 500/1000 microampere leakage limits are the Class I limits for non-cardiac-applied equipment. Dielectric withstand at 4kV AC demonstrates adequate insulation margin against mains voltage transients in hospital environments. | subsystem, power-management, compliance, regulatory, session-374, idempotency:sub-pms-iec60601-compliance-374
SUB-MAIN-118 | The Motion Scaling Module SHALL implement motion-to-command scaling in compliance with IEC 80601-2-77 (medical robot requirements), producing scaling factors adjustable from 1:1 to 10:1 in 0.5 increments, with scaling coefficient verified to be monotonically decreasing and configurable only during instrument change-out state, not during active manipulation. | subsystem-requirements | Main |  | Test |  |  | IEC 80601-2-77 sets specific requirements for the motion path of surgical robots. Restricting scaling adjustment to instrument change-out state prevents mid-procedure parameter changes that could startle the surgeon. 10:1 maximum provides fine tremor filtering without eliminating proprioceptive workspace awareness. | subsystem, motion-control, compliance, regulatory, session-374, idempotency:sub-msm-iec80601-compliance-374
SUB-MAIN-119 | The Workspace Safety Enforcer SHALL comply with ISO 10218-1:2011 industrial robot safety requirements as adapted for medical use under IEC 80601-2-77, implementing a minimum of SIL 2 (IEC 62061) safety integrity level for all workspace boundary enforcement functions, with a probability of dangerous failure per hour (PFHd) not exceeding 1E-7. | subsystem-requirements | Main |  | Analysis |  |  | The Workspace Safety Enforcer is the last software defence against robot arm collision with patient anatomy or OR equipment. ISO 10218-1 and IEC 80601-2-77 together mandate SIL 2 for safety functions in collaborative and surgical robotics. PFHd 1E-7 corresponds to a target risk reduction of 10 million operations per dangerous failure. | subsystem, motion-control, workspace-safety, compliance, safety, session-374, idempotency:sub-wse-sil2-compliance-374
SUB-MAIN-120 | The Real-Time Protocol Engine SHALL implement dual-path redundancy with a primary Ethernet deterministic path (IEEE 802.1Qbv TSN) and a secondary CAN FD backup path, automatically failing over to the backup path within 5ms of detecting primary path loss, with no motion command gap exceeding 20ms during failover transition. | subsystem-requirements | Main |  | Test |  |  | The Real-Time Protocol Engine is classified System-Essential: its failure stops all motion. Single-path Ethernet has documented failure modes (cable damage, switch failure) unacceptable in a surgical environment. TSN primary provides deterministic 1ms jitter; CAN FD backup provides hardware-level reliability at reduced bandwidth. 5ms failover is within the 100ms motion safety envelope. | subsystem, comms, real-time-protocol-engine, redundancy, session-374, idempotency:sub-rtpe-redundancy-374
SUB-MAIN-121 | The Haptic Feedback Subsystem SHALL maintain force-feedback rendering to the surgeon master handles from a hot-standby Haptic Controller process that assumes command within 10ms of primary controller failure, with the standby process running on a physically separate processor sharing force sensor data via a dedicated inter-processor link. | subsystem-requirements | Main |  | Test |  |  | Loss of haptic feedback during a surgical manoeuvre removes the surgeon's only sense of tissue resistance, increasing the risk of inadvertent tissue damage. The Haptic Feedback Subsystem is System-Essential. Hot standby on separate silicon prevents correlated failures from a single compute fault. 10ms switchover preserves the surgeon's control loop without perceptible discontinuity. | subsystem, haptic, redundancy, session-374, idempotency:sub-hfs-redundancy-374
SUB-MAIN-122 | The Stereoscopic Display System SHALL present the left and right stereo channels with inter-ocular distance adjustable from 58 to 72mm and vergence-accommodation conflict (VAC) below the 0.6 dioptre perceptual threshold, ensuring the surgeon's eyes converge at the same depth as the displayed surgical field. | subsystem-requirements | Main |  | Test |  |  | Vergence-accommodation conflict causes visual fatigue and depth perception errors in stereoscopic systems. The 0.6 dioptre threshold is the published human factors limit before perceptual discomfort. Adjustable inter-ocular distance (58-72mm) covers the 5th-95th percentile of adult surgeons. Without this requirement, the display system risks long-duration visual fatigue in multi-hour procedures. | subsystem, vision, eye, ergonomics, session-374, idempotency:sub-sds-eye-ergonomics-374
SUB-MAIN-123 | The Motion Scaling Module SHALL implement master-to-slave motion scaling ratios from 1:1 to 10:1 in configurable steps, with the active ratio displayed continuously at the surgeon console, and any ratio change requiring explicit confirmation via foot-pedal hold during instrument change-out state only. | subsystem-requirements | Main |  | Test |  |  | System requirement SYS-MAIN captures 1:1 to 10:1 master-to-slave scaling as a core surgical robot capability. This subsystem requirement operationalises the scaling ratio range, confirmation workflow, and display requirement needed to implement it safely. Confirmation workflow prevents accidental ratio change; continuous display prevents the surgeon operating with an unintended ratio. | subsystem, motion-control, tremor, scaling, session-374, idempotency:sub-msm-scaling-workflow-374
SUB-MAIN-124 | The Master Handle Actuator SHALL provide 6-DOF force reflection to the surgeon's master manipulator with a bandwidth of at least 30Hz and force dynamic range of 0.01N to 10N, with position sensing resolution finer than 0.1mm at the instrument tip-equivalent workspace position. | subsystem-requirements | Main |  | Test |  |  | 30Hz bandwidth matches the human proprioceptive resolution limit, ensuring the surgeon perceives all force events at the instrument tip. 0.01N minimum force reflects delicate tissue contact; 10N maximum corresponds to maximum safe instrument force. 0.1mm tip-equivalent resolution prevents staircase artefacts when the surgeon palpates tissue margins. | subsystem, haptic, master-manipulator, session-374, idempotency:sub-mha-master-manipulator-374
SUB-MAIN-125 | The Surgeon Input Console outer surfaces and all non-sterile patient-side cart covers SHALL withstand repeated wiping with 70% isopropyl alcohol (IPA) solution, 2% glutaraldehyde, and quaternary ammonium compound wipes without surface degradation, delamination, or ingress into electronics over a service life of 10 years and 3000 cleaning cycles. | subsystem-requirements | Main |  | Test |  |  | IPA 70%, glutaraldehyde 2%, and QAC wipes are the three most common OR surface disinfectants. Compatibility with all three is required because hospital purchasing decisions vary by institution. 3000 cycles corresponds to daily disinfection over a 10-year device lifetime. Surface failure risks chemical ingress causing electrical faults in a clinical environment. | subsystem, surgeon-console, sterility, compliance, session-374, idempotency:sub-console-ipa-disinfection-374
SUB-MAIN-126 | The Communication and Data Management System SHALL authenticate all inter-cart command messages using HMAC-SHA256 with a 256-bit session key negotiated at cart pair-up via ECDH-P384, rejecting any message with invalid authentication tag within one communication cycle (1ms), logging the rejection as a security event, and triggering SAFE_HOLD after 3 consecutive authentication failures. | subsystem-requirements | Main |  | Test |  |  | Cryptographic authentication mitigates spoofed motion command injection, a credible attack vector in robotic surgery given the increasing connectivity of OR equipment. HMAC-SHA256 provides collision-resistant authentication within the 1ms cycle budget. ECDH-P384 key exchange avoids pre-shared keys that would be impractical to rotate. 3-failure threshold limits denial-of-service via message flooding while tolerating transient bit errors. | subsystem, comms, cybersecurity, compliance, session-374, idempotency:sub-cdms-crypto-auth-374
SUB-MAIN-127 | The Motion Control and Scaling Subsystem SHALL enforce a minimum inter-arm clearance of 15mm between any two patient-side arm segments by computing pairwise convex-hull distances at 100Hz and commanding the affected arms to halt and retract when predicted clearance falls below 25mm, completing the halt within 50ms of detection. | subsystem-requirements | Main |  | Test |  |  | With three instrument arms operating simultaneously in a body cavity of approximately 150mm diameter, arm-to-arm collision is a credible failure mode not addressed by the existing workspace boundary enforcement which only constrains each arm against patient anatomy. A 15mm hard limit with a 25mm warning threshold at 100Hz gives 100ms reaction window at maximum arm velocity (250mm/s); the 50ms halt budget leaves margin for mechanical deceleration. Analogous requirement present in ISO 10218-1 Clause 5.11 for industrial robot systems operating in collaborative workspace. | subsystem, motion-control, collision-avoidance, safety, validation, session-377
SUB-MAIN-128 | The Surgeon Input Console SHALL implement a compliant alarm management system per IEC 60601-1-8:2006, assigning each alarm condition to one of three priority levels (HIGH, MEDIUM, LOW), providing distinct visual alarm signals (red flashing for HIGH, amber flashing for MEDIUM, amber steady for LOW) and distinct auditory alarm signals with signal patterns per IEC 60601-1-8 Annex F, and continuing to signal active alarms on internal battery power following loss of mains. | subsystem-requirements | Main |  | Inspection |  |  | IEC 60601-1-8 alarm management is a mandatory collateral standard under IEC 60601-1 for medical devices that generate alarm conditions. A surgical robot generates multiple concurrent alarm conditions (communication fault, energy delivery fault, instrument fault, force limit, watchdog) that must be prioritised and clearly distinguished to prevent alarm fatigue causing surgeons to miss safety-critical alerts during a procedure. The internal battery continuation requirement prevents alarm masking during the UPS power-bridge period. | subsystem, surgeon-console, alarm-management, regulatory, validation, session-377
SYS-MAIN-001 | The system SHALL provide master-to-slave motion scaling from 1:1 to 10:1 with <1ms end-to-end control loop latency and tip position repeatability of ±0.1mm under nominal OR conditions. | system-requirements | Main |  | Test |  |  | Sub-millimetre precision (STK-MAIN-001) requires quantified latency and repeatability: >1ms control loop latency introduces perceptible lag causing surgical error; ±0.1mm repeatability matches fine suture placement requirements in cardiovascular and ENT surgery. | rt-mechanical-trace, red-team-session-502
SYS-MAIN-002 | The system SHALL detect any single-point failure (communication loss, power fault, sensor fault, software exception) and achieve a safe state (all joints braked, energy de-energised) within 250ms of fault onset. | system-requirements | Main |  | Test |  |  | 250ms is derived from maximum safe uncontrolled instrument travel distance (<0.5mm at typical 2mm/s max tip velocity) — a hard limit from hazard analysis. Single-point failure coverage required for SIL 3 classification under IEC 62061. | rt-mechanical-trace, red-team-session-502
SYS-MAIN-003 | The system SHALL provide the surgeon with stereoscopic 3D high-definition video (minimum 1080p per eye at 60Hz) with colour fidelity sufficient for tissue and bleeding discrimination, at <100ms end-to-end video latency. | system-requirements | Main |  | Test |  |  | 60Hz 3D video prevents perception of flicker; <100ms latency is the threshold above which surgeons report disorientation (clinical usability studies). Tissue colour discrimination is safety-critical for distinguishing healthy and ischaemic tissue. | system, imaging, performance, session-341, idempotency:sys-imaging-quality-001-341
SYS-MAIN-004 | The system SHALL measure instrument-tissue interaction forces at the instrument tip with resolution of ≤0.1N over a range of 0–10N and transmit force feedback cues to the surgeon's master manipulator with fidelity sufficient to distinguish tissue planes. | system-requirements | Main |  | Test |  |  | Haptic feedback prevents inadvertent excessive tissue force; ≤0.1N resolution is required to feel tissue-plane transitions (typically 0.2–0.5N differential). Without force feedback, suture breakage rates and inadvertent organ perforation increase (documented in clinical literature). | system, haptics, force, session-341, idempotency:sys-force-sensing-001-341
SYS-MAIN-005 | The system SHALL maintain controlled arm position and instrument retraction capability for a minimum of 60 seconds following loss of mains power, sufficient to complete safe withdrawal of all instruments from the patient. | system-requirements | Main |  | Test |  |  | Mains power failure during active surgery is a known hazard; uncontrolled arm drop on mains loss is life-threatening. 60s battery bridge covers typical handoff time for instrument withdrawal based on human factors study of procedural steps. | system, power, reliability, session-341, idempotency:sys-power-continuity-001-341
SYS-MAIN-006 | All components rated for the sterile field SHALL be compatible with full fluid immersion using standard hospital disinfectants (IPA 70%, glutaraldehyde) or single-use disposability, and SHALL meet ISO 11135 sterilisation standards. | system-requirements | Main |  | Test |  |  | ISO 11135 and EU MDR Annex I mandate validated sterility assurance for patient-contacting devices. IPA 70% and glutaraldehyde are the standard biocides used in OR disinfection protocols; compatibility prevents material degradation and failure in service. | system, sterility, compliance, session-341, idempotency:sys-sterile-compliance-001-341
SYS-MAIN-007 | The Surgical Robot System SHALL transmit surgeon hand motion commands from master manipulator to instrument tip within 100ms end-to-end under all operating conditions. | system-requirements | Main |  | Test |  |  | Latency budget derived from human motor control studies: perceptible lag above 100ms disrupts surgeon proprioception and creates oscillatory overcorrection. Value validated by da Vinci and RAVEN II published performance data. | system, latency, performance, session-340
SYS-MAIN-008 | The Surgical Robot System SHALL provide selectable motion scaling ratios of 3:1, 5:1, and 10:1 between surgeon console input and instrument tip output. | system-requirements | Main |  | Test |  |  | Motion scaling enables micro-surgical precision: a 10mm surgeon hand movement produces a 1mm instrument movement at 10:1, enabling suturing of 1-2mm vessels not achievable with freehand technique. Three ratios cover different procedure types from gross to micro. | system, motion-scaling, session-340
SYS-MAIN-009 | The Surgical Robot System SHALL attenuate involuntary hand tremor frequency components above 6Hz by at least 40dB at the instrument tip. | system-requirements | Main |  | Test |  |  | Physiological tremor in surgeons is 8-12Hz at 0.1-0.5mm amplitude; at 10:1 motion scaling without filtration this would translate to unacceptable 0.01-0.05mm tip oscillation on delicate tissue. 6Hz cutoff preserves intentional motion bandwidth while removing tremor. | system, tremor, session-340
SYS-MAIN-010 | When an emergency stop command is issued via any input (surgeon foot pedal, bedside assistant button, or safety watchdog fault), the Surgical Robot System SHALL arrest all instrument motion and cut servo power within 50ms. | system-requirements | Main |  | Test |  |  | 50ms arrest time derived from worst-case instrument velocity of 50mm/s: at 50ms arrest, maximum overshoot is 2.5mm. Any larger overshoot risks laceration of adjacent tissue. This is an IEC 62304 SIL 3 safety function. | rt-sil-gap, red-team-session-502
SYS-MAIN-011 | The Surgical Robot System SHALL provide the surgeon with synchronised stereo HD video of the surgical field at 1080p per eye at 60Hz with end-to-end display latency below 50ms. | system-requirements | Main |  | Test |  |  | 50ms video latency budget is half the motor latency budget to ensure visual feedback arrives before the surgeon's corrective motion: higher video latency than motor latency causes the surgeon to over-correct. 1080p/60Hz matches clinical standard for high-fidelity tissue discrimination. | system, vision, session-340
SYS-MAIN-012 | The Surgical Robot System SHALL limit the force applied to patient tissue by any instrument tip to a maximum of 5N under normal operating conditions, with a safety cutoff at 8N triggering automatic clutch disengagement. | system-requirements | Main |  | Test |  |  | Tissue damage threshold studies show that inadvertent forces above 5N on bowel serosa or vessel walls cause serosal tears; 8N exceeds the tensile strength of small bowel mesentery. Dual thresholds allow warning before hard cutoff to avoid abrupt motion. | rt-sil-gap, red-team-session-502
SYS-MAIN-013 | The Surgical Robot System SHALL maintain full operational capability for a minimum of 8 consecutive hours without requiring maintenance, cooling, or consumable replacement. | system-requirements | Main |  | Test |  |  | Derived from STK-MAIN-015: surgical lists run 7-9 hours. System must outlast the list; thermal modelling of electronics and actuator duty cycles must confirm no performance degradation in the final hour of a maximum-length list. | system, availability, session-340
SYS-MAIN-014 | The Surgical Robot System SHALL enable patient-side robotic arm surfaces and instrument channels to be sterile-draped or autoclaved to EN ISO 17665 standard before each procedure. | system-requirements | Main |  | Inspection |  |  | Derived from STK-MAIN-011: all surface-contacting components must support validated sterilisation cycle. Autoclave compatibility restricts material selection (no ABS plastics, requires PEEK and 316L stainless on patient-contact surfaces). | system, sterility, session-340
SYS-MAIN-015 | The Surgical Robot System SHALL record all kinematic data, video streams, and system events at 1kHz temporal resolution to an encrypted on-system storage, retained for a minimum of 90 days. | system-requirements | Main |  | Inspection |  |  | Derived from STK-MAIN-014: EU MDR Article 83 and NHS clinical governance require procedure-level audit trails. 1kHz kinematics captures all motion events; 90-day retention covers most surgical complication investigation windows. | system, data-recording, session-340
SYS-MAIN-016 | When one instrument arm loses servo communication, the Surgical Robot System SHALL maintain full function on all remaining arms and alert the surgeon within 500ms, without requiring system restart. | system-requirements | Main |  | Test |  |  | Single-arm dropout must not abort a procedure mid-operation: patient is already prepared and open. Maintaining 2 of 3 arms allows the surgeon to complete critical steps before safely withdrawing. 500ms alert keeps surgeon awareness within one action cycle. | system, degraded-mode, resilience, session-340
SYS-MAIN-017 | The Surgical Robot System SHALL deliver controlled electrosurgical energy through robotic instruments in RF and ultrasonic modalities, with energy activation latency no greater than 100ms and deactivation latency no greater than 50ms from input command. | system-requirements | Main |  | Test |  |  | Electrosurgical energy delivery is a core surgical function of a robotic system. Activation and deactivation latency bounds are derived from IEC 60601-2-2 and clinical workflow requirements: 100ms activation is acceptable for surgeon intent recognition; 50ms deactivation is the safety-critical parameter preventing unintended tissue damage after the surgeon releases activation. | system, energy-delivery, session-352
SYS-MAIN-018 | The Surgical Robot System SHALL authenticate all safety-critical inter-subsystem command interfaces using cryptographic message authentication, ensuring that no unauthenticated command can reach a motion control, energy delivery, or safety-critical subsystem, with authentication failure triggering a safe-state transition within one control cycle. | system-requirements | Main |  | Test |  |  | IEC 62443-3-3 SR 1.2 requires authentication for all users, software processes, and devices that access control system resources. In a surgical robot, command injection at inter-subsystem interfaces represents a Class III medical device cybersecurity risk per FDA 2023 guidance. Authentication must be a system-level requirement so it cascades to all critical subsystems including KE, TG, RTPE, and TTAC. | rt-sil-gap, red-team-session-502
SYS-MAIN-019 | The Surgical Robot System SHALL comply with IEC 60601-1-2:2014 for healthcare facility environments, demonstrating immunity to radiated and conducted RF disturbances at professional healthcare facility levels and producing conducted and radiated emissions within Group 1 Class B limits. | system-requirements | Main |  | Test |  |  | The OR environment contains monopolar electrosurgical generators at 300kHz-3MHz and up to 400W, diathermy equipment, and wireless patient monitoring. IEC 60601-1-2 HPHE immunity compliance is mandatory for CE marking under MDR 2017/745 and directly mitigates risk of motion commands being corrupted by conducted interference from co-located electrosurgical equipment. | system, emc, regulatory, validation, session-377
VER-MAIN-001 | Verify IFC-MAIN-001: Inject synthetic joint torque ramp via test harness on real-time bus; measure monitor receipt latency with logic analyser. Pass criteria: ≥1000 frames/s received, all frames pass CRC, latency <200µs at 99th percentile over 60-second test. | verification-plan | Main |  | Test |  |  | Integration test proves the data pipeline meets both rate and latency requirements simultaneously under representative surgical load conditions. | verification, sis, ifc, session-341, idempotency:ver-ifc1-torque-bus-001-341
VER-MAIN-001 | Verify SUB-MAIN-001: On integrated arm test bench, ramp joint torque via calibrated load cell on each of 7 DOF to 110% of threshold. Measure time from threshold crossing to SIS safe-state assertion. Repeat 30 cycles per axis. Accept: detection within specification; zero missed trips. | verification-plan | Main |  | Test |  |  | SUB-MAIN-001 is a SIL 3 safety function requiring a quantified test procedure. Bench isolation confirms detection function.
VER-MAIN-002 | Verify IFC-MAIN-002: Inject E-stop loop break via test relay; measure contactor drop-out time with oscilloscope across contactor coil. Pass criteria: contactor opens within 50ms of loop break across 10 trials at ambient temperature (22°C) and at cold soak (5°C). | verification-plan | Main |  | Test |  |  | Direct oscilloscope measurement of contactor drop-out provides unambiguous evidence of the 50ms requirement. Cold soak tests component behaviour near lower operating temperature limit. | verification, sis, estop, session-341, idempotency:ver-ifc2-estop-contactor-001-341
VER-MAIN-003 | Verify IFC-MAIN-003: Using communication test tool, introduce artificial latency (10ms, 20ms steps) and packet loss (1, 3, 5 frame bursts) on fibre link; monitor sideband register via logic probe. Pass criteria: Communication Monitor reads correct latency and loss values within 1 status cycle (1ms) and triggers alert/safe-state at defined thresholds. | verification-plan | Main |  | Test |  |  | Fault injection testing is the only way to verify the communication monitor sees faults correctly without disrupting the data path. Threshold accuracy must be tested at both nominal and boundary conditions. | verification, sis, comms, session-341, idempotency:ver-ifc3-comms-monitor-001-341
VER-MAIN-003 | Verify SUB-MAIN-003: Trigger emergency stop from all three input paths (surgeon foot pedal, bedside button, software watchdog). Measure time from trigger to confirmed safe state (all joints braked, energy de-energised) using oscilloscope on brake coil drive and energy enable signals. Pass: ≤150ms for each of 20 trials per path (total 60 trials). | verification-plan | Main |  | Test |  |  | Three separate trigger paths must each be verified independently since each has distinct signal routing; shared test would not reveal path-specific faults.
VER-MAIN-004 | Verify IFC-MAIN-004: Trigger OPERATIONAL→DEGRADED and DEGRADED→SAFE-HOLD transitions via fault simulator; capture safety bus traffic with protocol analyser. Pass criteria: all subscribing subsystems receive state broadcast within 5ms, zero frames lost, publisher-subscriber acknowledgment complete within 10ms across 100 trials. | verification-plan | Main |  | Test |  |  | State broadcast latency directly affects the 250ms safe-state budget; the 5ms requirement needs verification under full system bus load to rule out congestion effects. | verification, sis, safe-state, broadcast, session-341, idempotency:ver-ifc4-ssm-broadcast-001-341
VER-MAIN-004 | Verify SUB-MAIN-002: Inject 90ms, 100ms, and 110ms artificial latency on inter-cart fibre link via delay injector. Verify SIS detects latency exceeding threshold and initiates controlled hold. Pass: 100% detection at 100ms-plus, zero false positives at 90ms over 50 injection cycles. | verification-plan | Main |  | Test |  |  | Latency detection threshold verified with injected faults to confirm sensitivity and specificity without waiting for real fault events.
VER-MAIN-005 | Verify SYS-MAIN-002 end-to-end: Inject simulated single-point failure (joint torque, comms loss, power fault — one at a time) during active simulated procedure; measure time from fault onset to all-joints-braked and energy-off confirmation via instrumented test rig. Pass criteria: ≤250ms across 50 trials per fault type, zero cases of uncontrolled arm motion. | verification-plan | Main |  | Test |  |  | End-to-end safe-state timing is a system-level property that cannot be decomposed to subsystem tests — the orchestration latency (SIS, servo, energy) is only visible at system integration level. | verification, system-level, safe-state, integration, session-341, idempotency:ver-sys2-e2e-safe-state-001-341
VER-MAIN-005 | Verify SUB-MAIN-004: Power the Watchdog Timer Controller from the auxiliary supply rail only. Trigger a main system power fault. Confirm watchdog continues counting on oscilloscope. Trigger a normal software heartbeat reset and confirm timeout. Pass: watchdog persists through main power fault; timeout behaviour correct on loss of heartbeat. | verification-plan | Main |  | Test |  |  | Physical isolation of watchdog power is a SIL 3 hardware independence requirement; must be verified by actual power interruption, not simulation, to prove no common-cause power coupling.
VER-MAIN-006 | Verify IFC-MAIN-005: Inject synthetic Cartesian velocity commands at 1kHz from a simulated Surgeon Console and measure packet arrival timestamps at the Motion Scaling Module input. Pass criteria: 100% packet delivery, mean latency <3ms, P99 latency <5ms over 60-second test at maximum command rate. | verification-plan | Main |  | Test |  |  | Integration test at the console-MC boundary verifies the network substrate delivers commands within latency budget. 60-second test captures steady-state behaviour after any TCP slow-start or OS scheduling settling. | verification, interface, session-340
VER-MAIN-006 | Verify SUB-MAIN-005: With system in OPERATIONAL state, manually break the E-stop series loop at each of six loop segments in sequence. At each break, verify: all joint brakes engage within 20ms (measured with oscilloscope on brake relay coils), energy delivery de-energised, system enters SAFE-HOLD. Total: 6 independent break-point tests. | verification-plan | Main |  | Test |  |  | Hardware E-stop series loop independence from software must be verified by physical loop interruption at every segment. Single-point break tests confirm the loop cannot be bypassed by a software or firmware fault.
VER-MAIN-007 | Verify IFC-MAIN-007: Halt the Motion Control heartbeat signal and measure the time from last heartbeat to brake engagement confirmation. Pass criteria: brake engagement in ≤50ms in 100 trials, P99 ≤55ms, zero trials exceeding 60ms. | verification-plan | Main |  | Test |  |  | Direct test of the safety-critical heartbeat fault path: this is a SIL 3 safety function and must be demonstrated under fault injection, not inferred from analysis. | verification, safety, session-340
VER-MAIN-007 | Verify SUB-MAIN-007: Command the Motion Control System with synthetic Cartesian velocity vectors that violate workspace boundaries on each of six boundary planes. Verify each command is rejected and substituted with a zero-velocity setpoint within one control cycle (1ms). Run 200 boundary-crossing events per plane. Pass: 100% rejection rate; no boundary violation recorded in joint encoder log. | verification-plan | Main |  | Test |  |  | Workspace boundary enforcement is a safety function preventing collision with patient anatomy; must be tested at each of six boundary planes to confirm independent enforcement.
VER-MAIN-008 | Verify SUB-MAIN-006: Run the Motion Control pipeline on target hardware at 1kHz for 30 minutes with 3-arm full load, instrument tip contact force at 4N (near-limit), and motion at maximum velocity. Instrument pipeline execution time on every cycle. Pass: worst-case single cycle <10ms, P99 <8ms, zero missed deadline across 1.8 million cycles. | verification-plan | Main |  | Test |  |  | Worst-case timing must be demonstrated on real hardware (not simulation) at realistic load because PREEMPT_RT latency spikes occur under thermal stress and simultaneous IO; P99 headroom of 2ms allows for measurement overhead and production unit variation. | verification, motion-control, session-340
VER-MAIN-008 | Verify SUB-MAIN-008: On the Real-Time Compute Node, run the Kinematics Engine at 1kHz with full 7-DOF arm configuration changes sampled from representative surgical trajectory database. Instrument worst-case execution time using hardware cycle counter. Pass: WCET for IK computation less than 500 microseconds at 99.9th percentile over 1 million samples. | verification-plan | Main |  | Test |  |  | IK computation WCET must be verified on target hardware under representative surgical trajectories, not simulation, as jitter varies with trajectory complexity and cache state.
VER-MAIN-009 | Verify end-to-end control loop: Apply a 10mm/s step input at the master manipulator and measure instrument tip response. Pass: tip begins moving within 100ms, steady-state tracking error <0.5mm, no oscillation above 0.1mm amplitude after 200ms. Test at all three scaling ratios (3:1, 5:1, 10:1) and both free-space and simulated tissue-contact conditions. | verification-plan | Main |  | Test |  |  | End-to-end step response test validates the integrated system behaviour combining network latency, computation, servo control, and mechanics. Must include contact conditions because compliance at the instrument tip changes closed-loop dynamics and can destabilise the servo at high scaling ratios. | verification, system-integration, session-340
VER-MAIN-009 | Verify SUB-MAIN-009: Command each of the 7 arm joints through representative surgical motion profiles (sinusoidal at 0.5, 1, and 2Hz). Measure position error between setpoint and encoder feedback at 1kHz for 60-second runs per frequency per joint. Pass: RMS tracking error below 0.05 degrees at all frequencies; no excursion above 0.1 degrees at any sample. | verification-plan | Main |  | Test |  |  | Position tracking accuracy is the foundational servo loop performance metric; tested under representative surgical bandwidths (0.5-2Hz covers intentional motion; tremor filtered above 6Hz).
VER-MAIN-010 | Verify IFC-MAIN-010: Connect Stereo Endoscope to Camera Control Unit via dual HD-SDI cables. Inject PRBS-31 test pattern on each channel independently. Measure BER over 10-minute continuous transmission using SDI analyser. Verify each channel achieves 1.485 Gbps data rate with BER below 1e-12. Disconnect one channel and verify the other continues uninterrupted. Pass: both channels meet BER threshold independently. | verification-plan | Main |  | Test |  |  | BER measurement with PRBS-31 is the standard method for characterising SDI link quality. Independent channel test confirms graceful degradation path. | verification, vision, session-341
VER-MAIN-010 | Verify SUB-MAIN-013: Present a Ronchi ruling resolution chart at the surgical working distance (50-250mm range). Capture images from both stereo channels. Measure limiting resolution using MTF analysis. Pass: both channels resolve 20 lp/mm at 50mm working distance; no inter-channel resolution difference greater than 10%. | verification-plan | Main |  | Test |  |  | Resolution verification must be performed at the actual surgical working distance on real optics, not from specification sheets, as lens-sensor assembly tolerances affect delivered resolution.
VER-MAIN-011 | Verify IFC-MAIN-011: Transmit colour bar test pattern from CCU to Image Processing Pipeline via dual 3G-SDI. Measure inter-channel temporal skew using dual-channel frame grabber with sub-microsecond timestamp resolution. Verify skew does not exceed 100 microseconds over 30-minute continuous operation. Pass: all measured skew values below 100us threshold. | verification-plan | Main |  | Test |  |  | Inter-channel skew measurement requires precision timestamping equipment. 30-minute test duration covers thermal stabilisation effects on cable propagation delay. | verification, vision, session-341
VER-MAIN-011 | Verify SUB-MAIN-014: Using a photodiode array triggered by a common LED strobe, capture stereo channel timing signals. Measure phase difference between left and right channel frame-valid pulses over 10,000 frames. Pass: all inter-channel phase differences less than 100 microseconds; zero frames with phase difference above 200 microseconds. | verification-plan | Main |  | Test |  |  | Inter-channel synchronisation is safety-critical for 3D depth perception accuracy; asynchrony above 100us creates measurable depth parallax errors. Must be verified in hardware as FPGA timing parameters can drift with temperature.
VER-MAIN-012 | Verify IFC-MAIN-012: Stimulate CCU with abrupt scene brightness change (white-to-black target swap). Measure time from CCU exposure metering output to illumination intensity change using photodiode and oscilloscope. Verify command-to-intensity latency does not exceed 16ms. Repeat for 100 cycles. Pass: all latencies below 16ms, no RS-422 frame errors. | verification-plan | Main |  | Test |  |  | Step-change stimulus provides worst-case latency measurement. 100 cycle repetition ensures statistical confidence in the latency bound. | verification, vision, session-341
VER-MAIN-012 | Verify SUB-MAIN-017: Display a full-white left-channel image with full-black right-channel image. Measure luminance from both left and right eyepieces using calibrated photometer. Calculate crosstalk ratio. Pass: crosstalk from left-to-right and right-to-left both less than 1%. Repeat with channels swapped. | verification-plan | Main |  | Test |  |  | Display crosstalk above 1% causes ghosting visible during 3D depth perception, potentially causing misjudgement of depth during fine dissection. Measured directly on the integrated display assembly.
VER-MAIN-013 | Verify IFC-MAIN-013: Drive Stereoscopic Display System from Image Processing Pipeline via dual DisplayPort 1.2 with synthetic surgical scene. Measure end-to-end latency from IPP output timestamp to photon emission using photodiode on display surface and oscilloscope. Verify latency does not exceed 8ms for 1000 consecutive frames. Verify 3840x2160 resolution and 10-bit colour depth via display analyser. Pass: all latencies below 8ms, resolution and colour depth confirmed. | verification-plan | Main |  | Test |  |  | Photodiode measurement of photon emission time is the gold standard for display latency characterisation. 1000-frame test covers worst-case frame timing. | verification, vision, session-341
VER-MAIN-013 | Verify IFC-MAIN-020: Couple five instrument types to the Instrument Recognition Module in sequence. Capture the data payload transmitted to the Tool Tip Articulation Controller after each coupling. Verify instrument type code, calibration offsets, use count, and sterilisation history are all present and correct against the instrument chip contents. Pass: 100% field accuracy across all 5 instruments; transfer complete within 200ms of coupling. | verification-plan | Main |  | Test |  |  | IFC-020 data accuracy is required for correct cable displacement calculation; incorrect calibration offsets cause tip position error, directly violating SYS-001 precision.
VER-MAIN-014 | Verify IFC-MAIN-014: Record 8-hour continuous procedure simulation from Image Processing Pipeline to Procedure Video Recorder. Verify recorded video is 1080p60 H.265 at 50Mbps CBR with no frame drops. Verify embedded audio channels are present. Verify parallel metadata timestamps correlate with video frames within 1ms using spot-check of 100 randomly selected frames. Pass: zero frame drops, audio present, all timestamp correlations within 1ms. | verification-plan | Main |  | Test |  |  | 8-hour endurance recording test matches operational requirement. Spot-check of 100 frames provides statistical confidence in timestamp accuracy across the full recording duration. | verification, vision, session-341
VER-MAIN-014 | Verify IFC-MAIN-021: Using a cable tension simulator, inject known tension values (nominal, low-threshold, high-threshold, and anomalous) on each of 4 cable channels. Monitor the SIS input bus with a logic analyser. Verify reported tension values match injected values within tolerance, and verify SIS receives anomaly assertion within 10ms of threshold breach. | verification-plan | Main |  | Test |  |  | IFC-021 is the signal path enabling SIS detection of cable mechanical faults; accuracy and latency must be verified to confirm the safety function meets its 250ms safe-state budget.
VER-MAIN-015 | Verify SUB-MAIN-019: During active stereo display operation, disconnect one HD-SDI channel at the endoscope connector. Measure time from disconnection to 2D monocular video appearing on display. Verify visual alert is displayed. Reconnect channel and verify stereo recovery. Repeat for both left and right channels. Pass: switchover to 2D within 500ms for both channels, alert displayed within 500ms, stereo recovery within 2s of reconnection. | verification-plan | Main |  | Test |  |  | Physical disconnection simulates the most common single-channel failure mode. Testing both channels confirms symmetric degradation. Recovery test validates the system can return to stereo operation without restart. | verification, vision, safety, session-341
VER-MAIN-015 | Verify IFC-MAIN-022: Command the Tool Tip Articulation Controller through a representative wrist-flexion trajectory (±90 degrees at 30 degrees/s). Monitor cable displacement commands delivered to the Instrument Drive Unit via the interface bus. Measure command latency and resolution. Pass: latency less than 1ms per command cycle; displacement resolution at or better than 0.01mm. | verification-plan | Main |  | Test |  |  | IFC-022 command resolution drives achievable tip position accuracy; 0.01mm resolution supports the ±0.1mm tip precision in SYS-001. Latency verified under dynamic motion, not static commands.
VER-MAIN-016 | Verify end-to-end vision chain: Position Stereo Endoscope viewing a calibrated resolution target in simulated body cavity with surgical illumination. Measure end-to-end latency from target motion (motorised at 20mm/s) to display update using high-speed external camera observing both target and display simultaneously. Verify total visual latency from scene change to photon emission does not exceed 35ms. Verify resolution target resolves 20 lp/mm. Verify stereo depth accuracy within 1mm at 50mm working distance using calibrated depth target. Pass: latency below 35ms, resolution confirmed, depth accuracy within 1mm. | verification-plan | Main |  | Test |  |  | End-to-end vision chain test validates the complete optical path from endoscope through CCU, IPP, to display. 35ms latency budget is the vision chain allocation from the 50ms total hand-to-eye budget. High-speed camera method provides ground-truth latency measurement independent of system clocks. | verification, vision, integration, session-341
VER-MAIN-016 | Verify IFC-MAIN-023: Using a test fixture, trigger the Instrument Lifecycle Controller to issue a lockout command by presenting an instrument with use count at maximum. Capture the lockout command transmission to the Safe State Manager. Verify the Safe State Manager inhibits instrument coupling within 500ms. Pass: lockout command issued; coupling inhibited; no mechanism for surgeon override without authorised code. | verification-plan | Main |  | Test |  |  | IFC-023 is a patient safety interface: preventing use of an overused instrument eliminates a mechanical failure mode. Must be tested with full lockout chain, not mock signals, to confirm end-to-end enforcement.
VER-MAIN-017 | Verify IFC-MAIN-015: Connect Force Sensing Module to Force Signal Conditioner via production cable. Inject common-mode signals at 50Hz, 150Hz, 1kHz at 10V amplitude. Measure differential output; pass criterion: differential error <10mV at each frequency (80dB CMRR at 1kHz). | verification-plan | Main |  | Test |  |  | Functional test on production hardware verifying CMRR against OR noise frequencies. 10mV threshold derived from 80dB CMRR at 10V CM input. | verification, haptic, session-342
VER-MAIN-018 | Verify IFC-MAIN-016: Connect Force Signal Conditioner to Haptic Controller on target hardware. Apply known reference loads (0N, 5N, 15N, 30N on each axis). Measure SPI frame arrival timestamps across 10,000 consecutive samples; pass criterion: all inter-arrival intervals within 1ms±50us, maximum propagation latency <100us. | verification-plan | Main |  | Test |  |  | Latency measurement on target hardware under load validates the 100us propagation budget that feeds into the 2ms end-to-end haptic latency requirement. | verification, haptic, session-342
VER-MAIN-019 | Verify IFC-MAIN-017: Command Haptic Controller to transmit torque setpoints exceeding 1.2Nm on all 7 joints simultaneously via CAN FD. Measure actual motor driver output with current clamp; pass criterion: no joint exceeds 1.2Nm output. Also verify 1kHz command rate maintained for 60 seconds under peak load without dropped frames. | verification-plan | Main |  | Test |  |  | Hardware torque limit verification requires commanding beyond the limit; this test confirms the motor driver hardware enforces the cap independently of software. Dropped-frame test validates bandwidth adequacy. | verification, haptic, session-342
VER-MAIN-020 | Verify IFC-MAIN-018: On target hardware, generate peak traffic (21 kinematics channels at 1kHz, two 4K60 video streams) over the production fibre link. Measure per-frame latency for kinematic frames with hardware timestamping across 100,000 frames; pass criterion: all kinematic frames delivered within 200us, no frames dropped. | verification-plan | Main |  | Test |  |  | Worst-case throughput test on target hardware; hardware timestamping eliminates OS jitter from measurements. | verification, comms, session-342
VER-MAIN-021 | Verify IFC-MAIN-019 and SUB-MAIN-028: Inject primary fibre link failure by disconnecting the active optical transceiver during live command traffic. Measure time to first valid frame received on standby link via oscilloscope trigger. Pass criterion: standby link active within 10ms, zero kinematic frames missing from command stream (verified by sequence number continuity in Procedure Data Recorder log). | verification-plan | Main |  | Test |  |  | Physical transceiver disconnection simulates the most severe link failure. Sequence number continuity check in the recorder log verifies zero-loss switchover at the application layer, not just the physical layer. | verification, comms, failover, session-342
VER-MAIN-022 | Verify SUB-MAIN-032: Couple 5 different instrument types (grasper, scissors, needle driver, cautery hook, clip applier) to each arm position. Measure time from coupling detection signal to kinematic model availability at the Tool Tip Articulation Controller. Pass: all 15 trials complete recognition within 200ms. Verify chip data integrity by comparing read-back values against programmed reference values with zero errors across 100 coupling cycles. | verification-plan | Main |  | Test |  |  | Tests all instrument types across all arm positions to detect any position-dependent or type-dependent recognition delays. The 100-cycle endurance test validates NFC reader reliability under repeated mechanical coupling vibration. | verification, surgical-instrument-system, session-346
VER-MAIN-023 | Verify SUB-MAIN-034: Using a calibrated load cell on each cable, command the Cable Tensioning System to maintain nominal tension while externally perturbing cable load by +/-10% at frequencies from 0.1Hz to 100Hz. Pass: tension stays within +/-5% of set-point for perturbations within range. Inject a step tension change of 20% on one cable. Pass: anomaly alert received by Safety and Interlock Subsystem within 10ms of injection, containing correct arm ID and cable channel. | verification-plan | Main |  | Test |  |  | Frequency sweep validates dynamic tension control under realistic cable loads. The 20% step injection exceeds the 15% anomaly threshold to verify detection latency under worst-case conditions. | verification, surgical-instrument-system, session-346
VER-MAIN-024 | Verify SUB-MAIN-035: Subject 20 Sterile Adapter samples to 8-hour continuous torque cycling (2Nm per channel, 6 channels, 1Hz cycle rate) with 50N sustained axial load. Post-test: perform sterile barrier integrity test per ASTM F1929 (dye penetration). Pass: zero dye penetration in all 20 samples. Measure torque transmission efficiency on a dynamometer. Pass: torque loss does not exceed 5% on any channel after 8-hour test. | verification-plan | Main |  | Test |  |  | Sample size of 20 provides statistical confidence for a single-use medical device. The dye penetration test is the regulatory standard for sterile barrier validation. Torque measurement after endurance cycling captures any degradation from wear. | verification, surgical-instrument-system, session-346
VER-MAIN-025 | Verify Console Computer network isolation by penetration test: connect test system to hospital information network port and attempt TCP connections to all CAN-FD gateway IP addresses. Zero successful connections SHALL be established. Repeat with crafted VLAN-hopping frames. Verify TLS certificate chain with a non-client-cert connection attempt — SHALL be rejected. | verification-plan | Main |  | Test |  |  | Penetration testing is required (not just inspection) because network isolation failures are often due to misconfigured firewall rules or VLAN tagging errors that are not visible in design documentation. Active attack simulation is the only reliable verification method per IEC 81001-5-1 Annex C.
VER-MAIN-025 | Verify SUB-MAIN-038: With three instrument arms active, inject a cable tension anomaly on one arm via the test interface. Measure time from anomaly injection to affected arm brake engagement. Pass: brake engaged within 50ms. Verify remaining arms continue executing a pre-programmed trajectory with no position deviation exceeding 0.1mm. Verify surgeon console displays affected arm identity and failure type within 200ms of anomaly. | verification-plan | Main |  | Test |  |  | Tests the full degraded-mode chain: anomaly detection, safety shutdown of affected arm, isolation of remaining arms, and operator notification. The 0.1mm position check on remaining arms ensures the shutdown transient does not propagate through shared power or communication buses. | verification, surgical-instrument-system, session-346
VER-MAIN-026 | Verify IFC-MAIN-024: On target Real-Time Compute Node, inject synthetic 6-DOF velocity vectors at 1kHz into the Tremor Rejection Filter output buffer. Measure Motion Scaling Module read latency with hardware timestamp counter. Pass criteria: 99.99% of reads complete within 5μs, zero stale-data detections over 60-second test window, all 6 DOF values match injected data to 15-digit precision. | verification-plan | Main |  | Test |  |  | Validates shared-memory interface timing and data integrity under sustained load. The 99.99% threshold allows 6 exceedances per minute, consistent with PREEMPT_RT scheduling guarantees on the target platform. | verification, motion-control, session-348
VER-MAIN-027 | Verify IFC-MAIN-026: On target hardware, fill the SPSC ring buffer from the Trajectory Generator at 1kHz while injecting controlled scheduling delays of 0-3ms on the Kinematics Engine consumer thread. Pass criteria: zero data loss over 10-minute test, buffer occupancy never exceeds 3 frames under nominal conditions, quaternion norm deviation below 1e-12 after transfer. | verification-plan | Main |  | Test |  |  | Exercises the ring buffer under worst-case jitter conditions to confirm the 4-frame depth provides adequate margin. Quaternion norm check detects any corruption in the lock-free transfer path. | verification, motion-control, session-348
VER-MAIN-028 | Verify Master Handle Actuator backdrive force when Haptic Feedback Subsystem is in STANDBY state: measure handle resistance at 10 evenly-spaced positions across full articulation range. Force SHALL be ≤0.1N at all positions. Transition system from ACTIVE to STANDBY while maintaining handle position — verify no transient force spike exceeds 0.5N during transition. | verification-plan | Main |  | Test |  |  | A haptic actuator that holds non-zero force in STANDBY creates a control input bias that the surgeon cannot distinguish from genuine force feedback. Measuring 10 positions ensures coverage of the full cam profile, not just the nominal centre position where most bias errors are zero.
VER-MAIN-028 | Verify IFC-MAIN-027: On integrated motion control bench with 7-DOF servo drives connected via EtherCAT, command sinusoidal joint trajectories at 1kHz across all joints simultaneously. Measure frame delivery latency with EtherCAT distributed clock timestamps. Pass criteria: 100% of frames delivered within 250μs, position/velocity/torque feedforward values match commanded values to 24-bit precision, zero frame drops over 30-minute continuous run. | verification-plan | Main |  | Test |  |  | EtherCAT frame delivery timing is the tightest constraint in the servo loop. The 30-minute duration exercises the interface through thermal steady-state of the servo drives and compute node, which is when timing margins are smallest. | verification, motion-control, session-348
VER-MAIN-029 | Verify IFC-MAIN-028: On target hardware, update workspace boundary polytope at 100Hz while the Kinematics Engine reads constraints at 1kHz. Inject boundary changes that shrink the workspace by 50% mid-trajectory. Pass criteria: zero partial reads detected via CAS validation, Kinematics Engine respects updated boundary within 20ms of change, polytope constraint count correctly varies from 6 to 24 half-planes. | verification-plan | Main |  | Test |  |  | The dynamic boundary update during active motion is the highest-risk scenario for IFC-MAIN-028 — a partial read could create a non-convex feasible region allowing unsafe motion outside the intended workspace. The 20ms response target is 2× the 10ms update interval, providing margin. | verification, motion-control, session-348
VER-MAIN-030 | Verify IFC-MAIN-025: On target hardware, transmit scaled velocity commands at 1kHz while switching scaling ratio (3:1→5:1→10:1) at 10-second intervals. Pass criteria: Trajectory Generator correctly reads active ratio metadata within 1ms of change, timestamps are monotonically increasing with zero gaps, data integrity verified by checksum comparison over 5-minute run. | verification-plan | Main |  | Test |  |  | Scaling ratio transitions during active motion are a common surgical workflow — the surgeon changes scaling mid-procedure. The test validates that ratio metadata propagation does not cause the Trajectory Generator to apply stale scaling, which would produce a motion discontinuity. | verification, motion-control, session-348
VER-MAIN-031 | Verify end-to-end Motion Control pipeline: Inject synthetic surgeon hand motion (sinusoidal at 0.5Hz, 2Hz, and 5Hz with 6Hz+ tremor component) at the Tremor Rejection Filter input. Measure instrument tip position at the Joint Servo Controller output via joint encoder feedback through forward kinematics. Pass criteria: tremor component attenuated by ≥40dB at tip, end-to-end pipeline latency ≤4ms (4 control cycles), tip position tracking error ≤0.05mm RMS at 0.5Hz, scaling ratio transitions produce no position discontinuity >0.02mm, Workspace Safety Enforcer halts motion within 2ms of boundary violation. | verification-plan | Main |  | Test |  |  | System-level integration test exercising the complete Cartesian-to-joint-space pipeline under realistic surgical motion profiles. This test cannot be decomposed into component-level tests because it validates the interaction between tremor rejection, scaling, trajectory generation, kinematics, servo control, and workspace enforcement as a coupled chain. | verification, motion-control, integration, session-348
VER-MAIN-032 | Verify REQ-SESURGICALROBOT-029: On the Real-Time Compute Node, inject 1000 joint-space command frames into the Kinematics Engine over 1 second — 990 with valid HMAC-SHA256 MACs and 10 with corrupted MACs. Confirm that all 990 valid frames are processed and all 10 invalid frames are rejected within 1ms and logged to the Procedure Data Recorder with timestamp and source ID. Confirm zero invalid frames appear in the kinematics output stream. | verification-plan | Main |  | Test |  |  | Authentication test must cover both the rejection path and the logging path to verify the security requirement is fully implemented. The 1kHz injection rate matches the operational servo rate, ensuring the test exercises the authentication mechanism under realistic timing conditions.
VER-MAIN-032 | Verify IFC-MAIN-029: Simulate mains loss by opening the main supply contactor under full system load. Measure Auxiliary Power Supply output voltage at Emergency Stop Chain contactor coil terminals using oscilloscope at 1MHz sampling. Pass criterion: voltage remains within 22–26V DC with no interruption exceeding 10ms duration. | verification-plan | Main |  | Test |  |  | Direct measurement at the interface boundary under the worst-case transient condition (full-load mains loss). Oscilloscope capture at 1MHz is necessary to detect sub-millisecond interruptions that cause contactor dropout. | verification, power-management, session-350, idempotency:ver-ifc029-power-estop-350
VER-MAIN-033 | Verify SUB-MAIN-043: Discharge UPS Battery Module to 80% charge, then simulate mains loss under full system operational load (all arms active, recording running). Measure time to first system functional failure or controlled shutdown initiation. Pass criterion: all system functions maintained for minimum 30 minutes, followed by controlled safe shutdown. | verification-plan | Main |  | Test |  |  | Full-load discharge test with pre-conditioned battery state (80%) represents realistic worst-case surgical scenario. Testing from full charge would not reveal battery capacity margins under normal operational conditions. | verification, power-management, session-350, idempotency:ver-sub043-ups-duration-350
VER-MAIN-033 | Verify REQ-SESURGICALROBOT-030: Boot the Real-Time Compute Node and confirm the RSA-2048 signed workspace envelope is loaded from the write-protected memory partition. Then inject 50 waypoints within the signed envelope and 10 waypoints outside it. Confirm that all 10 out-of-envelope waypoints trigger a controlled stop within 50ms and that no trajectory segment is generated for any out-of-envelope waypoint. | verification-plan | Main |  | Test |  |  | The test must verify both the RSA signature check at boot and the per-waypoint envelope validation at runtime. Testing with a mix of valid and invalid waypoints ensures the validation logic is correctly implemented for both the positive and negative cases.
VER-MAIN-034 | Verify REQ-SESURGICALROBOT-031: Connect a PTP test injector to the Real-Time Protocol Engine network segment. Inject 100 valid PTP sync messages followed by 20 with corrupted HMACs and then a 15ms gap with no authenticated messages. Confirm: (1) all 20 invalid-MAC frames are discarded and logged; (2) safe hold state is entered within 10ms of the last authenticated frame; (3) surgeon console alert is generated within 100ms of safe hold entry. | verification-plan | Main |  | Test |  |  | The three-phase test exercises the authentication rejection path, the timing-loss safe-hold path, and the HMI alert path, matching the three SHALL clauses in the requirement. The 15ms gap exceeds the 10ms threshold, ensuring the safe-hold trigger is verified within budget.
VER-MAIN-034 | Verify IFC-MAIN-030: Inject 100 command frames per second via isolated CAN analyser; measure generator acknowledgement latency for 1000 frames. Pass if: (a) all frames acknowledged within 5ms, (b) no missed acknowledgements, (c) CAN bus error counter remains zero throughout 60-second test. | verification-plan | Main |  | Test |  |  | Integration test at the physical CAN interface validates both protocol compliance and timing under sustained load — essential for a safety-critical command bus. | verification, energy-delivery, session-352
VER-MAIN-035 | Verify SUB-MAIN-094: Commission an IEC 62304 software development environment audit. Review that the software development plan, requirements specification, detailed design, unit test records, and integration test records are complete and traceable. Confirm independent review records exist for all Class C safety-critical software modules. Pass criterion: all five lifecycle artefact categories present with no open actions at Class C severity. | verification-plan | Main |  | Inspection |  |  | IEC 62304 Class C qualification requires documented lifecycle evidence for motion control software that can cause patient harm. Inspection is the appropriate method because the requirement specifies a process standard, not a functional behaviour; functional testing cannot verify development-phase compliance.
VER-MAIN-035 | Verify IFC-MAIN-031: Monitor RS-485 traffic with protocol analyser during 5-minute ultrasonic activation sequence. Pass if: (a) blade temperature telemetry received at minimum 50Hz continuous, (b) power level commands issued at correct rate, (c) no framing errors detected. | verification-plan | Main |  | Test |  |  | Blade temperature telemetry is the safety-critical output of this interface; 50Hz minimum must be verified under sustained load to confirm the 100°C inhibit will function correctly. | verification, energy-delivery, session-352
VER-MAIN-036 | Verify IFC-MAIN-040: Inject a continuous 6-DOF Cartesian velocity command stream from the Console Computer test harness at 1kHz. Capture 10,000 consecutive frames on the Inter-Cart Fibre Link and verify: (a) all frames are 48 bytes, (b) sequence numbers are monotonically increasing with no gaps, (c) CRC-CCITT checksum matches computed value for each frame, (d) inter-frame interval does not exceed 1.1ms in more than 0.1% of cycles. Pass criterion: zero protocol format errors across the 10-second capture window. | verification-plan | Main |  | Test |  |  | IFC-MAIN-040 defines the kinematic command frame format (48-byte little-endian, 1kHz, sequence number, CRC-CCITT) that the Console Computer writes to the Inter-Cart Fibre Link. This protocol integrity is safety-critical: a framing error or dropped frame on the kinematic command path can cause the Real-Time Protocol Engine to stall or misinterpret commands, potentially resulting in uncontrolled arm motion. The test captures sufficient frames to detect systematic errors at the <1e-4 frame error rate used for safety-critical communication links.
VER-MAIN-036 | Verify IFC-MAIN-032: Simulate Return Electrode Monitor interlock line de-energisation with ESG in active monopolar mode. Pass if: ESG ceases monopolar RF output within 50ms of interlock line de-energisation, as measured on oscilloscope. Also verify: ESG refuses to enable monopolar output when interlock line is de-energised at startup. | verification-plan | Main |  | Test |  |  | Hardwired interlock must be verified at the hardware level, not via software simulation. Timing from oscilloscope provides ground truth for the de-energise-to-inhibit response. | verification, energy-delivery, safety, session-352
VER-MAIN-037 | Verify SUB-MAIN-071: Drive the Real-Time Protocol Engine with a 1kHz TDM master clock. Using a logic analyser with 1ns resolution, capture 100,000 consecutive frame transmission timestamps over a 100-second run. Compute cycle-to-cycle jitter. Pass criterion: all measured cycle-to-cycle jitter values within plus or minus 50 microseconds. | verification-plan | Main |  | Test |  |  | SUB-MAIN-071 specifies the +/-50 microsecond jitter budget for the TDM frame scheduler. This budget is the foundation of all downstream latency allocations in the control loop — if the RTPE scheduler drifts, downstream stages exceed their cycle windows and the 1ms end-to-end motion control budget is violated. Verification by direct oscilloscope measurement is mandatory because jitter cannot be inferred from software timing alone.
VER-MAIN-037 | Verify IFC-MAIN-033: Connect test load simulating vessel impedance profile to ESG output; confirm TEM reads impedance at minimum 1kHz and issues shutoff within 200ms of simulated endpoint (1.5kΩ rise in 400ms window). Verify via CAN log and oscilloscope capture of RF output envelope. | verification-plan | Main |  | Test |  |  | Vessel seal endpoint detection must be verified against a known impedance profile; real tissue is not available for hardware-level testing, so validated test loads are standard practice. | verification, energy-delivery, session-352
VER-MAIN-038 | Verify SUB-MAIN-082: Load a patient anatomy mesh into the Workspace Safety Enforcer. Command the robot arm through 100 approach trajectories targeting ten distinct anatomical boundary regions. Verify that repulsive constraint forces activate before any commanded trajectory penetrates the boundary. Measure penetration depth at point of constraint activation. Pass criterion: zero boundary penetrations; constraint activation latency below 2ms in all 100 trials. | verification-plan | Main |  | Test |  |  | SUB-MAIN-082 specifies the real-time proximity enforcement that prevents surgical instrument over-penetration into protected anatomy. Failure of the WSE to activate constraints in time is a patient safety hazard. The test directly exercises the collision detection path under representative surgical motion and verifies the non-penetration guarantee that the requirement establishes.
VER-MAIN-038 | Verify IFC-MAIN-034: Assert E-STOP signal on system safety bus while ESG is delivering active RF energy. Measure time from E-STOP assertion to RF output falling below 1W using oscilloscope. Pass if transition occurs within 20ms. Repeat 20 times; all must pass. | verification-plan | Main |  | Test |  |  | E-STOP response requires repeated testing to confirm reliability across the 20ms window — a single pass test would not reveal marginal timing cases. 20 repetitions is consistent with IEC 62061 functional safety evidence requirements for SIL 2 functions. | verification, energy-delivery, safety, session-352
VER-MAIN-039 | Verify SUB-MAIN-083: Simulate anatomy mesh corruption and load failure. Inject a corrupted mesh file and observe WSE behaviour. Verify the WSE switches to conservative bounding-box workspace limits and logs the failure within 500ms. Verify normal operation resumes when a valid mesh is loaded. Pass criterion: no unprotected arm motion during mesh failure; bounding-box workspace enforced throughout the failure window. | verification-plan | Main |  | Test |  |  | SUB-MAIN-083 specifies the degraded-mode behaviour for the WSE when the anatomy mesh is unavailable. Without verified fallback behaviour, a mesh database fault would leave the arm with no proximity constraints, creating an uncontrolled motion risk. Verification requires injecting the specific fault condition because normal integration testing cannot exercise degraded-mode paths.
VER-MAIN-039 | Verify end-to-end energy delivery chain: with a calibrated RF load (500Ω, simulating vessel tissue), activate monopolar energy via surgeon footswitch; measure (a) time from footswitch closure to RF reaching 90% of setpoint power, (b) time from footswitch release to RF falling below 1W, (c) REM alarm response to impedance step above 135Ω mid-activation. Pass criteria: (a) ≤100ms, (b) ≤50ms, (c) monopolar inhibited within 500ms. Test at 50W, 200W, and 400W setpoints. | verification-plan | Main |  | Test |  |  | End-to-end test exercises the full command chain from surgeon input through Energy Delivery Controller through ESG to tissue, validating that SYS-MAIN-017 and safety monitoring function together as an integrated system. Multi-power-level testing confirms compliance across the operating envelope. | verification, energy-delivery, system-integration, session-352
VER-MAIN-040 | Verify SUB-MAIN-051: Using a calibrated impedance bridge, step patient return electrode pad impedance from 100Ω (safe) to 150Ω (above 135Ω threshold) while monopolar energy is active. Measure time from impedance step to RF output inhibition. Pass if: inhibition occurs within 500ms, surgeon console alarm is triggered, and energy does not resume automatically. Repeat 10 times. | verification-plan | Main |  | Test |  |  | REM threshold response is a patient safety function verified by simulating the actual pad-lift impedance signature. 10 repetitions provide statistical confidence in the 500ms response time. | verification, energy-delivery, safety, session-352
VER-MAIN-041 | Verify SUB-MAIN-054: With the energy delivery system connected to a calibrated leakage current test fixture per IEC 60601-1, measure patient leakage current under normal condition and with single-fault conditions (open earth, reversed polarity). Pass if: normal condition ≤10μA, any single-fault condition ≤50μA, for all energy delivery circuit paths. | verification-plan | Main |  | Test |  |  | IEC 60601-1 Type CF compliance must be verified by electrical safety testing per the standard's prescribed test methodology. These limits are absolute safety boundaries that cannot be relaxed. | verification, energy-delivery, safety, session-352
VER-MAIN-042 | Verify IFC-MAIN-035: With an energy activation foot pedal connected to Energy Delivery Controller via isolated CAN bus, inject 100 activation and deactivation pedal events at 2Hz. Measure CAN message timestamp vs pedal switch closure. Pass criterion: all messages delivered within 10ms of pedal actuation. Verify isolation: inject 500V common-mode signal on pedal cable; verify no conduction path to system ground. | verification-plan | Main |  | Test |  |  | Integration test for energy pedal CAN interface. 100-event sample provides statistical confidence on latency performance. Isolation test verifies patient safety protection against return current paths through the pedal cable. | verification, surgeon-console, energy-delivery, session-353
VER-MAIN-043 | Verify IFC-MAIN-036: Command 200 clutch engage/disengage events via foot pedal at 1Hz. Log CAN bus timestamp and Motion Control System clutch state acknowledgement. Pass: all events received within 10ms. Verify priority: saturate CAN bus with configuration traffic to 80% load and rerun; clutch messages must still meet 10ms latency. | verification-plan | Main |  | Test |  |  | Clutch interface must be verified under bus-congestion conditions matching realistic surgical use where configuration and telemetry traffic competes for bus bandwidth. | verification, surgeon-console, motion-control, session-353
VER-MAIN-044 | Verify IFC-MAIN-038: Monitor USB 3.0 traffic between Voice Command Module and Console Computer for 60 minutes during simulated procedure using a USB protocol analyser. Verify: (1) only command ID, confidence score, and timestamp packets are transmitted; (2) no raw PCM audio data is present in any packet. Pass: zero raw audio bytes detected in any USB transaction. | verification-plan | Main |  | Test |  |  | Patient privacy audit: USB traffic analysis is the only reliable method to confirm that raw audio does not leave the Voice Command Module, as software assertions could be bypassed by firmware update. | verification, surgeon-console, voice, privacy, session-353
VER-MAIN-045 | Verify SUB-MAIN-059: Transition system to OPERATIONAL state via software command. Issue motorised axis adjustment command to each of the 5 positioning axes within 500ms of state transition. Verify all commands are rejected. Time the lock-out interval from state transition to confirmed axis brake engagement on all axes. Pass: lock-out completes within 500ms; all adjustment commands during OPERATIONAL state are rejected. | verification-plan | Main |  | Test |  |  | Lock-out timing is safety-critical: an adjustment command accepted after state transition but before lock-out completes could shift master arm calibration. Test verifies both the timing bound and that the rejection is active across the full state transition window. | verification, surgeon-console, arm-positioning, session-353
VER-MAIN-046 | Verify end-to-end surgeon console to instrument tip chain: with system in OPERATIONAL state, surgeon activates energy foot pedal while commanding wrist motion via master arms. Measure: (1) pedal-to-energy-delivery latency; (2) master arm motion to instrument tip position update latency; (3) simultaneous energy and motion command coexistence without mutual interference. Pass: pedal latency <=50ms; tip position update <=1ms; no energy cutout or motion fault from combined command load. | verification-plan | Main |  | Test |  |  | End-to-end test validates that the Surgeon Input Console correctly multiplexes simultaneous motion and energy commands through independent paths without cross-interference, which cannot be verified by testing each interface in isolation. | verification, surgeon-console, integration, end-to-end, session-353
VER-MAIN-047 | Verify SUB-MAIN-062: Submit Haptic Controller SIL2 evidence package to IEC 62061 assessor, including FMEA, fault injection test results, and watchdog response time measurements. Pass criterion: assessor confirms SIL2 certification with no outstanding CARs. | verification-plan | Main |  | Analysis |  |  | SIL2 certification is verified by third-party assessment of the safety case, not by in-house bench test alone. | verification, haptic, session-354
VER-MAIN-048 | Verify REQ-SESURGICALROBOT-040: Inject primary haptic force-rendering processor fault via hardware fault injection on a system integration bench. Confirm switchover to secondary rendering path completes within 50 ms (measured from fault injection to first valid haptic torque output on secondary path). Verify contact-force reproduction error remains within 20 percent of nominal across 5 N, 10 N, and 15 N reference forces. Repeat 20 times. | verification-plan | Main |  | Test |  |  | Haptic processor redundancy is a safety-critical function; fault injection testing on the integration bench is the only method that can directly measure switchover latency under controlled conditions without risk to a patient. | verification, haptics, redundancy, session-367
VER-MAIN-048 | Verify SUB-MAIN-063: On haptic integration test rig, command 1kHz haptic render loop while varying simulated tissue contact stiffness from 0.1 N/mm to 10 N/mm in 0.1 N/mm steps, at all scaling ratios 1:1 to 10:1. Record master handle joint torques. Pass criterion: no sustained oscillation >0.05 Nm peak-to-peak for 30 seconds at any stiffness/scaling combination. | verification-plan | Main |  | Test |  |  | Stability must be demonstrated across the full operating envelope; an isolated soft-tissue or single-ratio pass is insufficient because instability is a boundary phenomenon. | verification, haptic, session-354
VER-MAIN-049 | Verify REQ-SESURGICALROBOT-042: With the surgical robot at nominal surgical load, disconnect AC mains at the PDU. Measure time from mains loss to all servo drives confirming sustained power on UPS. Confirm transfer time is less than 10 ms using a calibrated oscilloscope channel triggered on the mains loss event. Verify full-rated operation continues for 15 minutes minimum. Perform across input voltages of 85 VAC, 100 VAC, and 240 VAC. | verification-plan | Main |  | Test |  |  | UPS transfer timing must be verified under actual surgical load at worst-case input voltage; simulation cannot capture the real contactor and controller behavior during a live mains dropout. | verification, power, redundancy, session-367
VER-MAIN-049 | Verify IFC-MAIN-037: Inject 1000 consecutive session management messages (instrument configuration, state sync) at maximum payload from Console Computer to Real-Time Protocol Engine via USB 3.0; measure roundtrip latency distribution; pass criterion: 99th percentile latency below 10ms, zero message loss, session state consistent on both ends after 60-minute soak. | verification-plan | Main |  | Test |  |  | Session management is the data path for surgical case configuration and instrument profile loading. Latency and integrity verification under soak conditions catches USB buffer issues and driver-level race conditions that only manifest after sustained operation. | verification, surgeon-console, comms, session-356, idempotency:ver-ifc037-rtpe-session-356
VER-MAIN-050 | Verify SUB-MAIN-010: Command the Workspace Safety Enforcer to issue joint-angle commands within 2° of hard-stop boundary for each of the six joint axes using HIL simulator. Confirm WSE rejects the command within one 1kHz control cycle (≤1ms) and generates a safety violation event. Pass: 100% rejection rate across 1000 boundary-approach commands per axis. | verification-plan | Main |  | Test |  |  | Hardware-in-the-loop testing against the actual WSE FPGA firmware is the only method to verify the 1-cycle rejection latency under real interrupt timing. Simulation cannot accurately model FPGA-level preemption.
VER-MAIN-050 | Verify SUB-MAIN-055: Actuate each pedal function (energy activate, clutch, camera) 200 times with a calibrated mechanical jig; log CAN frame timestamps vs pedal contact timestamps via logic analyser; pass criterion: all 200 events per pedal function transmitted within 50ms of contact, zero missed events. | verification-plan | Main |  | Test |  |  | Mechanical jig actuation eliminates human reaction time variation. Logic analyser timestamps provide sub-microsecond precision needed to distinguish 50ms pass/fail boundary. 200 samples per function provides statistical confidence on the latency tail. | verification, surgeon-console, foot-pedal, session-356, idempotency:ver-sub055-pedal-latency-356
VER-MAIN-051 | Verify SUB-MAIN-011: Under full 1kHz control loop load with concurrent safety monitoring and data recording tasks active, inject 10,000 synthetic interrupt events via RTOS test hooks and measure worst-case interrupt latency using hardware logic analyser on GPIO trigger pin. Pass: maximum observed latency ≤50 microseconds across all 10,000 events. | verification-plan | Main |  | Test |  |  | Software-only timing measurement cannot exclude OS scheduling effects. Hardware logic analyser provides ground-truth latency independent of the kernel under test. 10,000 events provides statistical confidence >3-sigma for rare worst-case conditions.
VER-MAIN-051 | Verify SUB-MAIN-056 and SUB-MAIN-057: Present 500-item randomised command vocabulary to Voice Command Module via calibrated surgical headset at 55dB(A) ambient noise; record recognition result and dispatch timestamp for each command; pass criterion: word error rate no greater than 5 percent overall and across any individual speaker sample, command dispatch within 200ms of utterance completion for all recognised commands. | verification-plan | Main |  | Test |  |  | Combined WER and latency test using standardised noise conditions matches the surgical theatre acoustic environment. 500-item vocabulary covers the full command set plus confusables. Multi-speaker sample ensures the 5 percent WER criterion is not passed by a model overfitted to a single accent. | verification, surgeon-console, voice, session-356, idempotency:ver-sub056057-voice-wer-lat-356
VER-MAIN-052 | Verify SUB-MAIN-012: Inject each of the five defined fault conditions (communication loss, power fault, sensor fault, software exception, manual E-stop) via fault injection controller with all operator displays blanked. Verify the Safe State Manager broadcasts a safety-state transition command within 50ms of fault onset without any operator input. Pass: automated transition on 100% of injected faults; transition time ≤50ms in all cases. | verification-plan | Main |  | Test |  |  | Blanked operator displays ensure the test is not contaminated by accidental operator input. 50ms sub-budget is the SIS internal allocation within the 250ms system budget. Fault injection controller provides controlled, repeatable fault conditions isolated from normal operation.
VER-MAIN-052 | Verify SUB-MAIN-058: Attempt to enable robotic motion without completing surgeon authentication; verify system refuses and logs the attempt; complete authentication and verify motion enable succeeds; lock and re-enable under time pressure; pass criterion: motion always inhibited without valid auth, auth audit trail persists in procedure log. | verification-plan | Main |  | Demonstration |  |  | Authentication is a regulatory control; demonstration by qualified test engineers is the required verification method under IEC 62304 for software safety requirements. The audit trail check ensures the control is forensically traceable, not just functionally correct. | verification, surgeon-console, authentication, session-356, idempotency:ver-sub058-auth-motion-356
VER-MAIN-053 | Verify SUB-MAIN-015: Position calibrated cosine-corrected irradiance meter at 100mm working distance from illumination source. Measure surface irradiance during continuous illumination for 30 minutes. Pass: measured irradiance remains ≤100 mW/cm² throughout; no rising trend indicative of thermal runaway. | verification-plan | Main |  | Test |  |  | Cosine-corrected irradiance meter matches tissue surface geometry. 30-minute duration exceeds typical surgical procedure phase length to detect drift. Direct measurement is the only reliable verification method for a photon dose safety limit.
VER-MAIN-053 | Verify SUB-MAIN-060: Power on Console Computer from cold start 10 times; measure time from power-on to completion of self-test (including comms link, instrument detection, display check) using internal timestamp log; pass criterion: all 10 runs complete self-test within 90 seconds with all subsystems reporting ready. | verification-plan | Main |  | Test |  |  | Cold-start verification from power-off state, not warm reboot, is the clinically relevant scenario. 10 repeats provide statistical confidence across power-supply and thermal variation. The 90-second criterion directly matches the pre-operative scrub-in workflow window. | verification, surgeon-console, startup, session-356, idempotency:ver-sub060-startup-test-356
VER-MAIN-054 | Verify SUB-MAIN-016: Inject timestamp-tagged test frames at the camera sensor input; measure elapsed time to rendered display output using synchronised hardware timers. Run 1,000 frames at full 60Hz frame rate with all IPP enhancements active. Pass: latency ≤2ms for all 1,000 frames; no frame drops. | verification-plan | Main |  | Test |  |  | Hardware timestamping eliminates OS scheduling jitter from the measurement. 1,000-frame sample covers statistically rare worst-case IPP pipeline stalls. 2ms budget is the IPP allocation within the 33.3ms surgeon-to-display path.
VER-MAIN-054 | Verify SUB-MAIN-061: Disable Voice Command Module via software fault injection; confirm system remains in OPERATIONAL state with all foot pedal and master arm controls functioning normally; confirm voice failure is annunciated on console display; pass criterion: no robotic motion interruption on voice module loss, all alternative input channels remain active. | verification-plan | Main |  | Test |  |  | Fault injection test is the only method to verify graceful degradation without relying on a real hardware failure. The check on alternative input channels confirms the voice module is correctly isolated from the critical motion control path, not just that the fault detection works. | verification, surgeon-console, degraded-mode, session-356, idempotency:ver-sub061-voice-degraded-356
VER-MAIN-055 | Verify SUB-MAIN-018: Enable continuous recording on the Procedure Video Recorder and run a 180-minute simulated procedure session. Verify the composite 2D video stream is written without frame loss. Post-session, confirm recorded file integrity with frame-count check and checksum. Pass: 180-minute uninterrupted recording with ≤0 dropped frames and intact file checksum. | verification-plan | Main |  | Test |  |  | 180 minutes exceeds the 95th-percentile robotic procedure duration, providing confidence the recorder does not exhibit buffer overflow or storage performance degradation over typical procedure lengths.
VER-MAIN-055 | Verify IFC-MAIN-039: With Surgeon Interface Panel and Console Computer connected via EtherCAT bus, drive master arms through full workspace motion while logging upstream pose timestamps and downstream force command timestamps; measure bidirectional latency for 10000 cycles; pass criterion: upstream and downstream latency each below 1ms at 99.9th percentile, zero dropped cycles over 30-minute continuous run. | verification-plan | Main |  | Test |  |  | 10000 cycles at 1kHz provides a 10-second sample per run, repeated over 30 minutes to catch thermal drift and bus saturation effects. The 99.9th percentile criterion aligns with the haptic transparency requirement where occasional single-cycle exceedances are perceptible but rare. | verification, surgeon-console, haptic, session-356, idempotency:ver-ifc039-sip-cc-ethercat-356
VER-MAIN-056 | Verify SUB-MAIN-020: Place calibrated photodiode and tissue phantom (10mm optical depth matching published ICG tissue parameters) at surgical working distance. Enable fluorescence mode at 805nm. Measure irradiance at tissue surface and compare against ICG SNR threshold (>3:1 contrast). Switch to NIR mode and verify switchover time with hardware timer. Pass: surface irradiance ≥5 mW/cm²; SNR >3:1; mode switch ≤200ms. | verification-plan | Main |  | Test |  |  | Tissue phantom replicates 10mm optical path specified in SUB-MAIN-020. Photodiode measurement is ground truth for irradiance. SNR >3:1 is the published clinical utility threshold for ICG sentinel node mapping. Mode switch measured with hardware timer to exclude display latency.
VER-MAIN-056 | Verify SUB-MAIN-066: Drive each master arm through a 100mm/s continuous Cartesian trajectory for 60 seconds; sample pose output at 10kHz via test instrumentation and compute power spectral density to confirm 1kHz fundamental; measure position quantisation; pass criterion: pose data present at 1kHz plus or minus 5Hz, position quantisation no coarser than 0.1mm, zero dropped samples over 60 seconds. | verification-plan | Main |  | Test |  |  | Spectral analysis of the pose stream is the most reliable method to verify the actual output rate independent of internal timestamps. The 10kHz instrumentation sample rate provides Nyquist margin above the 1kHz criterion. Quantisation is verified by calculating minimum position step in recorded trajectory data. | verification, surgeon-console, haptic, session-356, idempotency:ver-sub066-sip-pose-rate-356
VER-MAIN-057 | Verify SUB-MAIN-021: Using the surgeon console UI on a representative surgical site, activate each of the ≥3 image enhancement modes (gain, contrast, colour balance). Confirm each mode is selectable, applied within one display frame (≤16.7ms), and independently toggleable. Pass: all ≥3 modes available and functional; no mode interaction artifacts; switching time ≤16.7ms. | verification-plan | Main |  | Demonstration |  |  | Demonstration by a trained operator reflects real-world intraoperative access patterns. Frame-synchronous switching avoids tearing artefacts which could disorient the surgeon.
VER-MAIN-057 | Verify SUB-MAIN-067: Command step force inputs at 10N/s ramp from Console Computer; measure force output at master arm fingertips with calibrated force plate at 10kHz; compute step response latency and steady-state error; pass criterion: latency from command to 90 percent of commanded force below 1ms, steady-state error within 15 percent of commanded value across 0 to 5N range. | verification-plan | Main |  | Test |  |  | Step input characterisation is the standard method for actuator latency measurement. Force plate at 10kHz provides resolution well above the 1ms criterion. The 90 percent threshold for latency avoids confounding rise time with latency in the time-domain measurement. | verification, surgeon-console, haptic, session-356, idempotency:ver-sub067-haptic-force-356
VER-MAIN-058 | Verify SUB-MAIN-022: Apply known reference forces (0.1N, 0.5N, 1N, 2N, 5N) to instrument tip via calibrated force applicator across all three axes. Record Haptic Controller output force readings. Pass: measured error ≤0.1N at each reference point across all axes; linearity error ≤5% full scale. | verification-plan | Main |  | Test |  |  | Calibrated force applicator provides ground-truth reference independent of the sensor under test. Multi-axis testing is necessary because cable-driven instruments exhibit cross-axis coupling; each axis must be independently verified.
VER-MAIN-058 | Verify SUB-MAIN-068: Using a calibrated test rig, release master arm handle rapidly 100 times; measure time from sensor contact break to Console Computer inhibit signal via logic analyser; pass criterion: all 100 events inhibited within 50ms of contact break, instrument-side motion ceases within 1ms of inhibit signal. | verification-plan | Main |  | Test |  |  | 100-event sample provides statistical confidence on the 50ms upper bound with probability of passing by chance less than 0.001 for a system with 60ms true latency. The downstream motion cessation check closes the loop to confirm inhibit propagation reaches the instrument, not just the console output. | verification, surgeon-console, safety, session-356, idempotency:ver-sub068-handle-engage-356
VER-MAIN-059 | Verify SUB-MAIN-024: Command the Master Handle Actuator to render increasing force magnitudes (0.5N, 1N, 1.5N, 2N, 2.5N, 3N, 3.5N). Measure actual handle output force with in-line load cell. Pass: output force ≤2N for all commanded values; no force breakthrough for commanded values ≥2N. | verification-plan | Main |  | Test |  |  | The 2N limit is a patient safety threshold derived from tissue fragility models. Verification with in-line load cell provides calibrated ground truth independent of the actuator control firmware. Testing at values above the limit verifies the saturation behaviour, not just nominal operation.
VER-MAIN-059 | Verify SUB-MAIN-069: Conduct IEC 62304 compliance audit against software lifecycle documentation; review development traceability matrix from requirements to software units; inspect MDR 2017/745 Annex I checklist; pass criterion: all Class C activities evidenced, no gaps in requirements-to-code traceability, all Annex I essential requirements addressed. | verification-plan | Main |  | Inspection |  |  | IEC 62304 and MDR compliance are regulatory obligations verified by notified body inspection, not by functional test. The audit-and-inspection method matches the regulatory assurance activity required for CE marking. | verification, surgeon-console, compliance, session-356, idempotency:ver-sub069-iec62304-audit-356
VER-MAIN-060 | Verify SUB-MAIN-025: Apply 500V AC test voltage across Force Signal Conditioner isolation barrier per IEC 60601-1 clause 8.8. Measure dielectric withstand for 60 seconds. Measure working isolation voltage. Pass: no breakdown or flashover at 500V; working isolation ≥400V DC per IEC 60601-1 Type BF applied part. | verification-plan | Main |  | Test |  |  | Dielectric withstand test at 500V is the IEC 60601-1 mandatory type test for BF applied parts. This is the regulatory gold standard; simulation cannot substitute for high-voltage breakdown testing.
VER-MAIN-060 | Verify SUB-MAIN-070: Inject software exception into Console Computer via debug interface; measure time from exception trigger to SAFE-HOLD broadcast on subsystem bus; verify all subsystems acknowledge safe-state; check non-volatile storage contains 30-second pre-fault data; pass criterion: SAFE-HOLD within 500ms, all subsystems acknowledge, data readable post-fault. | verification-plan | Main |  | Test |  |  | Direct exception injection is the only reliable method to verify software failsafe behaviour without waiting for a real fault. Timing of safe-state broadcast is measured at the bus level to capture total system latency including exception handler execution time. | verification, surgeon-console, safety, session-356, idempotency:ver-sub070-cc-failsafe-356
VER-MAIN-061 | Verify SUB-MAIN-027: Timestamp kinematic command frames at RTPE output and at arm controller receipt using synchronised hardware clocks. Measure end-to-end latency across the fibre link under full 1kHz traffic load for 60,000 consecutive frames (60 seconds). Pass: latency ≤3ms for all frames; zero frame loss. | verification-plan | Main |  | Test |  |  | Hardware-synchronised timestamps eliminate clock drift error from the measurement. 60,000 frames captures rare burst-traffic interactions. 3ms is the CDMS allocation within the 100ms system command latency budget.
VER-MAIN-061 | Verify IFC-MAIN-041: Connect a calibrated timing analyser to the Real-Time Protocol Engine frame-delivery output. Drive the Inter-Cart Fibre Link with a synthetic 1kHz frame generator. Capture 100,000 consecutive inter-frame intervals and compute maximum jitter. Pass criterion: maximum jitter does not exceed 5 microseconds. Inject a frame with corrupted CRC and confirm the validity flag is asserted within one frame period. | verification-plan | Main |  | Test |  |  | Integration test verifying that the optical receiver and SERDES layer deliver frames to the Real-Time Protocol Engine within the timing specification required for SUB-MAIN-071 compliance. | verification, motion-control, infrastructure, session-357
VER-MAIN-062 | Verify SUB-MAIN-029: Enable PDR recording while running full 1kHz kinematic loop and dual 1080p video streams simultaneously. Record for 60 minutes. Analyse log file: verify kinematic data timestamps at 1ms intervals with jitter ≤100μs; verify stereo video frame sequence integrity. Pass: 1kHz rate maintained with ≤100μs jitter; no dropped frames in either video stream; log file checksum valid. | verification-plan | Main |  | Test |  |  | Simultaneous full-load testing validates PDR does not degrade control loop performance via resource contention. Timestamp jitter analysis detects DMA buffer overruns that would compromise forensic value of recordings.
VER-MAIN-062 | Verify IFC-MAIN-042: Configure Network Management Controller and two Joint Servo Controller test nodes on a bench EtherCAT segment. Capture 50,000 consecutive process data object exchanges using an EtherCAT frame monitor. Pass criteria: bus cycle achieves 1ms ±50μs; all process data objects contain target angle, feed-forward torque, mode, and fault mask fields; fault mask disablement of one node does not affect frame delivery to the remaining node within the same cycle. | verification-plan | Main |  | Test |  |  | EtherCAT process data object structure and timing must be verified at the subsystem integration level before patient-side cart integration. A 2-node bench test validates the topology management and per-node fault isolation behaviour defined in SUB-MAIN-074. | verification, motion-control, infrastructure, session-357
VER-MAIN-063 | Verify SUB-MAIN-039: Instrument the Trajectory Generator software with execution timing instrumentation. Run 60,000 consecutive trajectory computation cycles under worst-case kinematic complexity (maximum joint velocities, simultaneous 6-DOF motion, active tremor filter). Measure per-cycle wall-clock execution time. Pass: all 60,000 cycles complete within 1ms; no overruns detected by watchdog timer. | verification-plan | Main |  | Test |  |  | 60,000 cycles provides >99.9% confidence interval for rare worst-case branch paths. Watchdog timer provides independent overrun detection separate from the instrumented timing, preventing self-referential measurement errors.
VER-MAIN-063 | Verify IFC-MAIN-043: On target Real-Time Compute Node hardware, activate the DMA transfer path to the Procedure Data Recorder NVMe array. Capture PCIe transaction timestamps using a logic analyser on the PCIe bus for 60 minutes at 1kHz. Pass criteria: no DMA transfer latency exceeds 10 microseconds; all frames contain valid 64-bit UTC timestamp, 7-element angle, torque, and 6-element velocity arrays; zero frames are missing from the recorder file verified by sequence number continuity check. | verification-plan | Main |  | Test |  |  | The 60-minute soak validates steady-state NVMe write performance under sustained 1kHz load, including garbage collection pauses that might delay individual DMA completions. Sequence number verification provides a comprehensive check on data completeness without relying on file size alone. | verification, motion-control, infrastructure, session-357
VER-MAIN-064 | Verify SUB-MAIN-047: Connect calibrated 50Ω tissue-impedance test load. Command ESG to minimum (10W) and maximum (400W) output. Measure output power with RF wattmeter at each set point. Pass: measured power within ±5% of commanded value at both limits; no output until activation signal present. | verification-plan | Main |  | Test |  |  | 10-400W monopolar range covers standard laparoscopic and open surgical applications. RF wattmeter provides calibrated power measurement per IEC 60601-2-2. ±5% accuracy is the maximum tolerable power error for tissue effect repeatability.
VER-MAIN-064 | Verify IFC-MAIN-025: On the Real-Time Compute Node, inject scaled 6-DOF Cartesian velocity vectors at 1kHz from the Motion Scaling Module output. Using kernel timestamps, measure the latency from Motion Scaling Module output publication to Trajectory Generator input read. Pass criterion: latency does not exceed 50 microseconds for 99.9th percentile over 60,000 samples; no samples are dropped; scaled magnitudes are preserved within floating-point precision (1 ULP tolerance). | verification-plan | Main |  | Test |  |  | IFC-MAIN-025 has no existing verification entry. This test validates the shared-memory ring buffer pathway between the two software components on the Real-Time Compute Node, ensuring no priority inversion or scheduler jitter causes a velocity command to be silently dropped or delayed beyond the pipeline budget. | verification, motion-control, session-357
VER-MAIN-065 | Verify SUB-MAIN-050: Simultaneously command RF activation and ultrasonic activation via the Energy Delivery Controller interface. Verify the controller asserts exactly one modality within the same test cycle and logs a mutual-exclusion violation event. Repeat for 1,000 concurrent command pairs. Pass: RF and ultrasonic never simultaneously active; violation logged in 100% of cases. | verification-plan | Main |  | Test |  |  | Concurrent RF+ultrasonic activation is the failure mode that can cause uncontrolled tissue heating. 1,000 concurrent command pair test provides statistical confidence that the mutex logic has no race-condition path. Must be tested at the hardware interface level, not software simulation.
VER-MAIN-065 | Verify IFC-MAIN-024: Inject synthetic 7-DOF velocity signals at 1kHz with frequency components at 2Hz (intentional), 8Hz, 10Hz, and 12Hz (tremor). Measure the output stream from Tremor Rejection Filter to Motion Scaling Module; verify spectral power above 6Hz is attenuated by at least 40dB and below 3Hz is preserved within 0.5dB. Measure end-to-end latency from input to filtered output: SHALL be less than 5ms. Pass criterion: all three spectral and latency limits met across 60-second run. | verification-plan | Main |  | Test |  |  | Integration test verifying IFC-MAIN-024 interface compliance end-to-end at the component boundary, covering both the frequency-domain and timing properties required for safe motion command delivery. | verification, motion-control, tremor, session-358, idempotency:ver-ifc024-tremor-filter-output-358
VER-MAIN-066 | Verify SUB-MAIN-052: Using ex-vivo vessel tissue samples, apply ultrasonic energy and monitor TEM impedance output. Measure impedance at confirmed seal completion (visual and histological). Pass: TEM detects impedance rise within 500ms of confirmed vessel seal; false positive rate <5% on non-sealed tissue. | verification-plan | Main |  | Test |  |  | Ex-vivo tissue testing provides the ground-truth biological reference for impedance-based seal detection. Histological confirmation eliminates observer bias. 500ms detection window prevents energy over-delivery which causes thermal spread injury.
VER-MAIN-066 | Verify IFC-MAIN-025: Command scaled Cartesian velocity at each of the three scaling ratios (3:1, 5:1, 10:1). At the Trajectory Generator input, measure received velocity magnitude and confirm it matches expected scaled value within 0.1%. Inject a velocity command exceeding the 200mm/s limit and confirm the interface clamps output to 200mm/s within one 1ms cycle. Run 10,000 consecutive frames; verify zero dropped or reordered packets. Pass criterion: all accuracy, clamping, and delivery checks pass. | verification-plan | Main |  | Test |  |  | Integration test verifying IFC-MAIN-025 scaling accuracy, velocity clamping safety, and data integrity across the Motion Scaling to Trajectory Generator boundary. | verification, motion-control, scaling, session-358, idempotency:ver-ifc025-scaling-traj-358
VER-MAIN-067 | Verify SUB-MAIN-053: Configure EDC with 5-second timeout. Activate energy delivery and allow to run without surgeon intervention. Verify energy is automatically de-activated at 5 seconds. Repeat with 100ms, 1s, and 5s intervals. Pass: automatic termination in all cases; ≤10ms overshoot beyond timeout; audible alert generated on each termination. | verification-plan | Main |  | Test |  |  | Automatic timeout is the primary guard against unintentional tissue damage from persistent energy delivery. 10ms overshoot tolerance accounts for RTOS task scheduling. Test at three timeout values verifies the timer across different operational contexts.
VER-MAIN-067 | Verify IFC-MAIN-026: Inject a 3D Cartesian pose trajectory via the Trajectory Generator. At the Kinematics Engine input, record received pose setpoints at 1kHz. Verify S-curve velocity profile is present (measure acceleration continuity). Inject a setpoint that violates workspace limit and verify the Trajectory Generator aborts the segment and outputs ABORT within one 1ms cycle. Pass criterion: all S-curve and abort-timing checks pass across 5 representative trajectories. | verification-plan | Main |  | Test |  |  | Validates IFC-MAIN-026 interpolated pose delivery, timing integrity, and abort-on-violation behavior at the Trajectory Generator to Kinematics Engine boundary. | verification, motion-control, trajectory, session-358, idempotency:ver-ifc026-traj-ke-358
VER-MAIN-068 | Verify SUB-MAIN-065: Submit Haptic Feedback Subsystem to accredited test laboratory for IEC 60601-1 series evaluation covering: dielectric strength, leakage current, protective earth, enclosure protection (IP54), and EMC per IEC 60601-1-2. Pass: test laboratory issues certificate of compliance with no critical non-conformances. | verification-plan | Main |  | Test |  |  | IEC 60601-1 compliance certification requires accredited third-party laboratory testing — self-certification is not accepted by notified bodies for Class IIb medical devices. Accreditation is a regulatory prerequisite for CE marking and FDA 510(k) submission.
VER-MAIN-068 | Verify IFC-MAIN-027: Command a 360-degree joint rotation at maximum velocity for each of the 7 arm joints. At the Joint Servo Controller input, measure received joint angle setpoints at 1kHz. Verify setpoint-to-setpoint delta is within joint velocity limit. Measure position tracking error during continuous trajectory: SHALL be below 0.1 degrees RMS. Inject a setpoint discontinuity and verify it is rejected and a fault event generated within 2ms. Pass criterion: all velocity, tracking, and rejection checks pass. | verification-plan | Main |  | Test |  |  | Validates IFC-MAIN-027 joint angle delivery rate, magnitude continuity, and fault rejection at the Kinematics Engine to Joint Servo Controller boundary. | verification, motion-control, servo, session-358, idempotency:ver-ifc027-ke-servo-358
VER-MAIN-069 | Verify SUB-MAIN-040: Command the Motion Scaling Module to each of 3:1, 5:1, 7:1, and 10:1 scaling ratios. Apply 10mm/s master handle displacement in each axis. Measure instrument tip displacement. Pass: tip displacement equals master displacement divided by selected ratio ±1%; ratio applies uniformly to all six DOF. | verification-plan | Main |  | Test |  |  | Direct kinematic measurement at each prescribed ratio is the authoritative test. 1% tolerance reflects achievable precision of cable-driven instrument actuation. Six-DOF uniform application prevents axis-dependent scaling errors that cause surgeon disorientation.
VER-MAIN-069 | Verify IFC-MAIN-028: Load a patient anatomy mesh with five known no-go boundaries. Command arm motion toward each boundary at 50mm/s approach velocity. Measure: (a) repulsive torque onset latency from 5mm threshold crossing to first corrective command (SHALL be within 1ms), (b) final approach velocity at boundary (SHALL be below 5mm/s), (c) WORKSPACE_MODEL_FAULT generation time when mesh is invalidated (SHALL be within 50ms). Pass criterion: all three latency and velocity limits met across all five boundaries. | verification-plan | Main |  | Test |  |  | Validates the IFC-MAIN-028 proximity enforcement channel under realistic anatomy-boundary approach conditions, confirming that the WSE repulsive constraint prevents boundary violation at the required response latency. | verification, motion-control, workspace-safety, session-358, idempotency:ver-ifc028-wse-ke-358
VER-MAIN-070 | Verify SUB-MAIN-041: With HIL simulator, command the Trajectory Generator to compute segments approaching within 10%, 5%, 2%, and 1% of workspace boundary. Verify the generator clamps the trajectory at the boundary margin and generates a workspace-limit event. Pass: no trajectory segment exceeds the defined boundary; workspace-limit event generated in all approach cases. | verification-plan | Main |  | Test |  |  | HIL testing allows precise boundary approach scenarios that are infeasible on physical hardware without collision risk. Boundary clamping prevents instrument-tissue collision caused by commanded trajectories that exceed reachable workspace.
VER-MAIN-070 | Verify end-to-end MC subsystem pipeline: Apply a 10mm/s step velocity input at the master manipulator representing a 10:1 scaled command. Measure elapsed time from surgeon input reception to Joint Servo Controller setpoint update. SHALL be within 1ms total pipeline latency. Simultaneously inject 9Hz tremor component; verify it is attenuated by at least 40dB at the Joint Servo Controller input. Inject workspace boundary approach; verify repulsive torque present before boundary crossing. Run on three-arm configuration at 1kHz for 60 minutes with no control cycle overruns. Pass criterion: latency, attenuation, boundary, and reliability checks all pass. | verification-plan | Main |  | Test |  |  | End-to-end system integration test exercises the complete Tremor Rejection Filter to Joint Servo Controller chain simultaneously, confirming that pipeline latency, tremor rejection, safety enforcement, and reliability figures are all met together under peak load, not just in isolation. | verification, motion-control, integration, session-358, idempotency:ver-e2e-mc-pipeline-358
VER-MAIN-071 | Verify SUB-MAIN-042: Command the Motion Scaling Module to issue velocity commands at 90%, 100%, 110%, 150%, and 200% of the 25mm/s safety limit. Measure actual instrument tip velocity for each command. Pass: tip velocity clamped to ≤25mm/s for all commanded values ≥100%; safety saturation event generated for each over-limit command. | verification-plan | Main |  | Test |  |  | Testing at multiples of the limit (110%, 150%, 200%) verifies the saturation logic holds under extreme over-command conditions that might occur if scaling parameters are incorrectly configured. Safety saturation event enables autonomous loop review of velocity limit violations.
VER-MAIN-071 | Verify IFC-MAIN-001: Inject calibrated torque loads via a torque motor fixture on each arm joint in sequence. Confirm the Joint Force Monitor detects and broadcasts limit exceedance within 2ms. Pass criterion: all 7 joints trigger within 2ms with zero false negatives over 1000 consecutive injection cycles at 1kHz sampling. | verification-plan | Main |  | Test |  |  | IFC-MAIN-001 specifies the joint force safety limit interface. This test validates the 2ms detection latency required for safe force limiting — failure to detect within this window risks tissue damage or structural overload. 1000-cycle statistical basis required due to IEC 62061 SIL 2 demands on the safety function. | verification, safety, joint-force-monitor, session-360
VER-MAIN-072 | Verify SUB-MAIN-044: Log power enable signals for all subsystems during system startup using logic analyser. Verify sequencing order matches the prescribed sequence (SIS → computing → motion → displays → energy) with ≥100ms margin between each stage. Pass: correct order in 20 consecutive cold-start cycles; inter-stage margin ≥100ms. | verification-plan | Main |  | Test |  |  | Logic analyser provides hardware-level ground truth for power sequencing independent of firmware reporting. 20 cold-start cycles tests consistency across power-supply voltage variation within tolerance. 100ms margin prevents sequencing race conditions during transient loads.
VER-MAIN-072 | Verify IFC-MAIN-002: Actuate E-stop at each of the three locations (surgeon console, patient-side cart, circulating nurse button). Measure time from contact open to servo drive inhibit across all active joints. Pass criterion: de-energisation within 150ms at all three actuation points, zero motion after inhibit signal, confirmed across 50 trials each location. | verification-plan | Main |  | Test |  |  | IFC-MAIN-002 defines the hardwired E-stop interface to servo power. The 150ms de-energisation budget is derived from IEC 60601-1 general collapse time limit for Class II medical devices; failure to de-energise risks operator injury from uncontrolled arm motion. Physical actuation testing is mandatory — simulation cannot validate the hardware interlock chain. | verification, safety, emergency-stop, session-360
VER-MAIN-073 | Verify SUB-MAIN-045: Simulate main power loss by removing main power input. Measure time from main power loss to Auxiliary Power Supply energised output reaching nominal voltage on SIS and Watchdog rails. Sustain auxiliary load for 60 seconds. Pass: SIS and Watchdog rails remain within ±5% of nominal throughout 60-second auxiliary hold; main power transition ≤10ms. | verification-plan | Main |  | Test |  |  | Physical power removal test is the only reliable method to validate auxiliary hold-up time under real capacitor state-of-charge conditions. 60 seconds exceeds the safe-state procedure time to confirm the system can execute full shutdown without main power.
VER-MAIN-074 | Verify SUB-MAIN-046: Inject calibrated stereo test pattern (known disparity map) into Image Processing Pipeline while simultaneously processing all surgeon-selected enhancement modes. Measure disparity map accuracy at pipeline output. Pass: disparity RMS error ≤0.5 pixels; no enhancement mode degrades stereo depth accuracy. | verification-plan | Main |  | Test |  |  | Known disparity pattern provides quantitative ground truth for stereoscopic processing accuracy. Testing with all enhancement modes active verifies that image processing algorithms do not corrupt inter-frame phase relationships required for depth calculation.
VER-MAIN-074 | Verify IFC-MAIN-001: Apply calibrated torque to each arm joint. Confirm Joint Force Monitor detects limit exceedance within 2ms over 1000 cycles at 1kHz. Pass: all 7 joints trigger within 2ms, zero false negatives. | verification-plan | Main |  | Test |  |  | IFC-MAIN-001 defines the joint force safety detection interface at 1kHz. The 2ms limit is derived from the 5ms total safety response budget under IEC 62061 SIL 2. Statistical basis of 1000 cycles required for SIL 2 probability of failure on demand. | verification, safety, session-360
VER-MAIN-075 | Verify SUB-MAIN-048: Send RF activation command to Electrosurgical Generator and measure time from command receipt to full rated output power (±5% of set point) using calibrated RF power sensor and oscilloscope. Test at 50W, 200W, and 400W. Pass: full output achieved within 100ms at all set points. | verification-plan | Main |  | Test |  |  | Activation latency directly affects surgeon cut/coagulate feel. 100ms is the maximum delay before the surgeon perceives a response lag. Testing at three power levels verifies the power ramp rate is consistent across the full range, not only at nominal settings.
VER-MAIN-075 | Verify IFC-MAIN-002: Actuate E-stop at all three locations (surgeon console, patient-side cart, circulating nurse). Measure time from contact open to servo drive inhibit. Pass: de-energisation within 150ms at each location, zero motion after inhibit, 50 trials per location. | verification-plan | Main |  | Test |  |  | IFC-MAIN-002 defines the hardwired E-stop to servo drive interface. The 150ms de-energisation budget is set by IEC 60601-1 clause 9.8.3 maximum permitted collapse time for active implant-class devices. Physical actuation mandatory — simulation cannot validate the HW interlock chain. | verification, safety, emergency-stop, session-360
VER-MAIN-076 | Verify SUB-MAIN-049: Connect calibrated frequency counter to Ultrasonic Energy Module transducer output. Command activation at minimum, nominal, and maximum drive amplitude. Measure operating frequency. Pass: frequency within 55.5kHz ±200Hz across all drive levels; no harmonic oscillation above -40dB relative to fundamental. | verification-plan | Main |  | Test |  |  | 55.5kHz ±200Hz is the manufacturers tolerance band for ultrasonic transducer resonance. Frequency deviation beyond this band causes off-resonance operation, reducing cavitation efficiency and increasing transducer heating. Harmonic check ensures mechanical resonance does not activate tissue at unexpected frequencies.
VER-MAIN-076 | Verify IFC-MAIN-003: Inject bit-error patterns at rates of 1e-6, 1e-5, and 1e-4 BER on the inter-cart fibre link using a channel impairment tool. Confirm Communication Monitor detects each error within one frame period (1ms). Pass: detection within 1ms for all injected errors, fault isolation to affected channel within 3ms, no false alarm rate above 1e-5 per hour. | verification-plan | Main |  | Test |  |  | IFC-MAIN-003 requires the Communication Monitor to expose CRC errors on the fibre link within one frame period. Verification requires BER injection because real-world fibre faults are too infrequent to validate statistically in test; the no-false-alarm threshold prevents nuisance shutdowns under IEC 62061 SIL 2 diagnostic coverage requirements. | verification, safety, communication-monitor, session-360
VER-MAIN-077 | Verify SUB-MAIN-030: Inject synthetic frame sequence errors and CRC-failed frames into the RTPE input stream at rates of 1 per 10,000 frames (normal), 1 per 100 frames (degraded), and consecutive 10 frames (burst). Verify RTPE falls back to last-known-good command and generates a communication-integrity alert within one 1ms cycle. Pass: no propagation of corrupted commands; alert generated within 1ms; fallback command applied within one cycle of detection. | verification-plan | Main |  | Test |  |  | Sequence error and CRC failure injection replicates the electromagnetic interference conditions expected in an OR environment. Three injection rates test normal noise, degraded link, and burst-fault scenarios. 1ms response matches the 1kHz RTPE cycle period — critical for maintaining kinematic continuity.
VER-MAIN-077 | Verify IFC-MAIN-004: Trigger each of the three safety state transitions (OPERATIONAL to DEGRADED, DEGRADED to HALTED, OPERATIONAL to HALTED) using fault injection on Safety and Interlock Subsystem inputs. Measure broadcast latency from trigger to state word receipt on Motion Control bus. Pass: broadcast within 5ms for all transitions, all subsystems acknowledge within 10ms. | verification-plan | Main |  | Test |  |  | IFC-MAIN-004 specifies the Safe State Manager broadcast timing that gates downstream subsystem response. The 5ms broadcast budget and 10ms acknowledgement window are derived from the overall 150ms E-stop response chain; missing this window causes subsystems to act on stale state. | verification, safety, safe-state-manager, session-360
VER-MAIN-078 | Verify SUB-MAIN-023: Apply defined force profiles (0.1N, 0.5N, 1N, 2N step and ramp) at instrument tip via force fixture. Measure rendered force at master handle with calibrated load cell. Pass: rendered force tracks commanded profile with ≤10ms latency and ≤15% magnitude error. | verification-plan | Main |  | Test |  |  | Haptic rendering fidelity is central to safe tissue manipulation — rendering delay >10ms breaks surgeon sensorimotor loop. 15% magnitude tolerance reflects the haptic JND (just-noticeable difference) threshold for surgical tactile tasks.
VER-MAIN-078 | Verify IFC-MAIN-005: Drive the surgeon console master handles through representative surgical motion profiles at 10Hz, 50Hz, and 200Hz command rates. Confirm the Motion Scaling Module receives commands at specified rate with end-to-end latency under 100ms. Pass: commanded velocity received within 100ms at all rates, no dropped commands at 10Hz or 50Hz over a 60-second run. | verification-plan | Main |  | Test |  |  | IFC-MAIN-005 defines the surgeon console to motion scaling command interface. The 100ms end-to-end latency budget is the clinically established threshold for transparent teleoperation; exceeding it causes surgeon disorientation and loss of feel. Test at 200Hz exercises the interface beyond nominal rate to confirm no buffer overflow or frame drop degradation. | verification, motion-control, surgeon-console, session-360
VER-MAIN-079 | Verify SUB-MAIN-073: Measure EtherCAT distributed clock synchronisation error across all servo nodes using network oscilloscope capture of SYNC0 pulses. Run for 30 minutes under full motion load. Pass: synchronisation error ≤1μs peak across all nodes; no sync loss events in 30 minutes. | verification-plan | Main |  | Test |  |  | EtherCAT distributed clock jitter >1μs causes inter-axis coordination errors at 1kHz update rates, producing visible vibration artefacts. Network oscilloscope measures the actual hardware SYNC0 timing, independent of application-layer reporting.
VER-MAIN-079 | Verify IFC-MAIN-006: Transmit 250Hz joint position command stream from Motion Control System to Patient-Side Cart over the inter-cart fibre link for 30 minutes under surgical load simulation. Measure jitter and inter-frame gap. Pass: command stream arrives at 250Hz ±0.5Hz, inter-frame jitter less than 500us, zero missing frames in any 1-second window. | verification-plan | Main |  | Test |  |  | IFC-MAIN-006 specifies the 250Hz joint command rate that drives servo loop stability. Jitter above 500us causes the servo PID to perceive command steps rather than smooth trajectories, degrading tip positioning accuracy. Thirty-minute duration covers a typical operative case to detect thermal or load-dependent drift. | verification, motion-control, joint-command, session-360
VER-MAIN-080 | Verify SUB-MAIN-080: Apply known test signals at 1Hz, 4Hz, 8Hz, 12Hz, 16Hz, and 20Hz to Tremor Rejection Filter input. Measure filter output amplitude. Pass: attenuation <3dB for frequencies ≤8Hz (pass-band); attenuation >40dB at 12Hz and above (stop-band); zero-phase shift confirmed by cross-correlation. | verification-plan | Main |  | Test |  |  | 8th-order Butterworth specification mandates the pass/stop band boundary at 8Hz. Zero-phase requirement prevents the filter from adding predictive bias to surgeon intent, which would cause over/undershoot in intentional motions. Direct frequency sweep is the only reliable method to characterise a real-time DSP filter.
VER-MAIN-080 | Verify IFC-MAIN-007: Interrupt the safety heartbeat between Motion Control System and Safety and Watchdog System at intervals of 50ms, 100ms, and 200ms by injecting clock hold on the motion control output. Confirm Safety Subsystem detects each interruption and enters DEGRADED state. Pass: detection and DEGRADED transition within 110ms of first missed heartbeat, recovery to OPERATIONAL within 500ms of restored heartbeat. | verification-plan | Main |  | Test |  |  | IFC-MAIN-007 defines the safety heartbeat interface that provides liveness monitoring between subsystems. The 110ms detection window equals two missed heartbeat periods; the recovery time ensures the system can resume operation without manual reset after transient link failures, reducing unnecessary procedure interruptions. | verification, safety, heartbeat, session-360
VER-MAIN-081 | Verify SUB-MAIN-085: Introduce mechanical friction to joint under test to induce position error of 0.3°, 0.5°, 1°, and 2°. Measure time from error detection to power-hold command issued to motor driver. Pass: power-hold command issued within 20ms of crossing 0.5° threshold; alert generated for all crossings. | verification-plan | Main |  | Test |  |  | Physical friction induction replicates real joint-error conditions (cable stretch, mechanism wear) more faithfully than software simulation. 20ms response time allocates one-fifth of the 100ms command-to-arm latency budget to the servo error detection path.
VER-MAIN-081 | Verify IFC-MAIN-008: Transmit full-resolution stereo video stream from Vision and Imaging System to Surgeon Console at 60Hz for 60 minutes. Measure end-to-end latency from photon at endoscope to pixel on display. Pass: latency under 100ms at 99th percentile, zero dropped frames in any 10-second window, colour fidelity within ΔE<2 against reference chart. | verification-plan | Main |  | Test |  |  | IFC-MAIN-008 defines the stereo video interface from imaging to display. The 100ms latency limit is the clinical threshold for visual-motor coordination; ΔE<2 is the clinical colour reproduction standard for surgical tissue discrimination. Sixty-minute run mirrors a typical laparoscopic case duration. | verification, vision, stereo-video, session-360
VER-MAIN-082 | Verify SUB-MAIN-109: Enable PDR with full kinematic+video+event stream. Measure storage write throughput with disk performance monitor during 60-minute simulated procedure. Pass: sustained write throughput ≥500 MB/s throughout; no buffer overflow events; write latency ≤2ms p99. | verification-plan | Main |  | Test |  |  | PDR throughput must be tested under simultaneous full-data-rate conditions — kinematic at 1kHz and stereo video at 60Hz create correlated burst patterns that a sequential throughput test would miss. 2ms p99 write latency prevents DMA buffer overflow on the RTPE thread.
VER-MAIN-082 | Verify IFC-MAIN-009: Mount calibrated torque measurement fixture to Instrument Drive Unit. Command articulated wrist through 5 representative surgical manoeuvres. Confirm torque feedback data from Surgical Instrument System arrives at Motion Control at 500Hz with values within ±5% of fixture measurement. Pass: 500Hz data continuity across all manoeuvres, <5% error at all load points, no frame gaps above 2ms. | verification-plan | Main |  | Test |  |  | IFC-MAIN-009 defines the torque feedback interface from instrument to motion control, which drives the haptic rendering and safety torque limiting. The 500Hz rate is set by the haptic loop update frequency; errors above 5% cause inaccurate haptic rendering and unreliable safety limiting. | verification, instrument, torque-feedback, session-360
VER-MAIN-083 | Verify SUB-MAIN-110: After recording a 30-minute session, retrieve the SHA-256 hash from PDR metadata. Independently compute SHA-256 on the recorded file. Pass: hashes match; hash computation completes before next recording session start; hash stored in tamper-evident metadata header. | verification-plan | Main |  | Test |  |  | Independent hash computation provides a tamper-evidence baseline for regulatory audit trail purposes (EU MDR Article 83). Verifying the hash is stored in a tamper-evident header ensures the integrity check itself cannot be silently overwritten.
VER-MAIN-083 | Verify IFC-MAIN-010: Connect Stereo Endoscope to Camera Control Unit. Acquire 10-minute stereo video stream at 1920x1080 60Hz. Confirm both channels receive synchronised frames within 500us of each other, and pixel clock continuity is maintained. Pass: inter-channel sync within 500us over entire capture, no dropped frames, SNR >50dB per channel. | verification-plan | Main |  | Test |  |  | IFC-MAIN-010 defines the dual-channel endoscope to CCU interface. Inter-channel sync within 500us is required for stereopsis: timing errors above this threshold introduce perceived depth distortion that impairs surgical judgement. SNR >50dB ensures image quality does not degrade under LED illumination at minimum intensity. | verification, vision, stereo-endoscope, session-360
VER-MAIN-084 | Verify SUB-MAIN-086: Heat the RT Compute Node CPU junction to 88°C using a thermal chamber. Verify the system logs a thermal-throttle event and that the 1kHz kinematic loop continues without missed deadlines. Sustain 88°C for 10 minutes. Pass: thermal-throttle event logged within 1 second of crossing 85°C; ≤0 missed 1kHz deadlines during sustained 88°C operation. | verification-plan | Main |  | Test |  |  | Thermal chamber testing provides controlled, repeatable temperature conditions independent of ambient variation. The real-time constraint (zero missed deadlines) must be verified at elevated temperature because thermal throttling changes memory bus timing and cache hit rates.
VER-MAIN-084 | Verify end-to-end surgical teleoperation chain: from surgeon master handle motion through kinematics, motion control, fibre link, patient-side joint actuation, to instrument tip displacement, within 100ms total latency under peak load (all 7 joints active, full stereo video, haptic feedback active, energy system armed). Pass: tip displacement follows commanded trajectory within ±1mm, end-to-end latency under 100ms at 99th percentile over 10-minute run, safety monitoring active throughout with no spurious faults. | verification-plan | Main |  | Test |  |  | System-level integration test validates the full teleoperation chain as an assembled system. Individual subsystem tests cannot reveal timing interactions across the fibre link, shared compute resources, or priority conflicts on the real-time bus that only emerge under concurrent load. This test is the primary evidence for surgical effectiveness claims. | verification, system-integration, end-to-end, session-360
VER-MAIN-085 | Verify SUB-MAIN-077: Inject synthetic traffic at 2 Gbps, 5 Gbps, 8 Gbps, and 10 Gbps bidirectional using RFC 2544 test methodology. Measure achieved throughput and frame loss at each rate. Pass: ≥10 Gbps bidirectional throughput with ≤0.001% frame loss at full load; link stable for 10 minutes at full load. | verification-plan | Main |  | Test |  |  | RFC 2544 is the standard bidirectional throughput test methodology for point-to-point links. 10 Gbps supports simultaneous 1kHz kinematic commands (small frame, high rate) plus 4K stereo video (large frame, 60Hz) with 30% margin. 10-minute sustained load validates thermal stability of the optical transceiver under continuous operation.
VER-MAIN-085 | Verify IFC-MAIN-011: Apply known forces (0.5N, 1N, 2N, 5N) to the instrument tip using a calibrated force applicator. Confirm Haptic Feedback Subsystem renders corresponding forces at master handle within 15ms and within ±10% of target. Pass: render latency under 15ms, force accuracy within ±10%, no instability or sustained oscillation at any test load. | verification-plan | Main |  | Test |  |  | IFC-MAIN-011 defines the haptic force rendering interface. The 15ms render latency budget is the threshold for perceptual transparency in force-feedback teleoperation; exceeding it causes time-domain distortion that reduces surgical sensitivity. The ±10% accuracy threshold is set by the minimum force resolution clinically meaningful for tissue differentiation. | verification, haptic, force-rendering, session-360
VER-MAIN-086 | Verify IFC-MAIN-012: Simulate hospital mains interruption by cutting AC supply to the Power Management Subsystem. Confirm UPS Battery Module assumes load within 20ms and supplies all critical rails (24V, 12V, 5V) within ±5% for minimum 10 minutes. Pass: rail voltage maintained within ±5% from 20ms after interrupt, orderly shutdown sequence initiates at 2-minute remaining UPS capacity. | verification-plan | Main |  | Test |  |  | IFC-MAIN-012 defines the UPS interface between Power Management Subsystem and mains supply. The 20ms switchover limit prevents servo control loop dropout; 10-minute UPS run time allows a controlled procedure termination. Testing with actual supply interruption is required — power supply simulation cannot reproduce inrush dynamics of the real switchover. | verification, power, ups-battery, session-360
VER-MAIN-087 | Verify IFC-MAIN-013: Arm the Electrosurgical Generator with a 50-ohm resistive test load (representing typical tissue impedance). Issue activation command from Energy Delivery Controller. Confirm RF output rises to specified power within 500ms and cuts within 100ms of deactivation. Pass: power on within 500ms, power off within 100ms, no RF output without valid activation command, return electrode impedance check completed before each activation. | verification-plan | Main |  | Test |  |  | IFC-MAIN-013 defines the Energy Delivery Controller to Electrosurgical Generator activation interface. The 500ms power-on time is acceptable for surgical energy tools; the 100ms off time is safety-critical to prevent inadvertent tissue damage after command withdrawal. Mandatory return electrode check verifies patient protection function. | verification, energy, electrosurgical, session-360
VER-MAIN-088 | Verify IFC-MAIN-020: Present five instrument types to the Instrument Recognition Module in randomised sequence, including an unknown instrument and a previously used instrument at or beyond maximum use count. Confirm correct identification within 2 seconds of coupling and rejection of unknown and expired instruments. Pass: correct ID for all five valid types, rejection with fault code for unknown and expired, zero false accepts. | verification-plan | Main |  | Test |  |  | IFC-MAIN-020 defines instrument recognition interface. The 2-second recognition window is the maximum acceptable delay between instrument coupling and system readiness; false acceptance of expired instruments risks patient infection or instrument mechanical failure during procedure. | verification, instrument, recognition, session-360
VER-MAIN-089 | Verify IFC-MAIN-044: Inject 12-branch telemetry frames at 10 Hz on the CAN FD bus with simulated fault codes on 2 branches; confirm Power Sequencing Controller receives all frames within 5 ms latency measured by bus analyser, and that galvanic isolation withstands 500 VDC hipot test for 60 s without breakdown. | verification-plan | Main |  | Test |  |  | Integration test confirms CAN FD frame throughput and latency budget under realistic 12-channel load, plus electrical isolation integrity per IEC 60601-1 reinforced insulation requirements. | verification, power-management, session-361, idempotency:ver-ifc044-pdu-psc-361
VER-MAIN-090 | Verify IFC-MAIN-045: Apply 200 A pulse load (500 ms) via programmable load at PDU input terminals with UPS Battery Module at 30% state-of-charge; measure bus voltage droop and confirm impedance remains below 50 mOhm. Sustain 20 A continuous for 30 minutes and confirm no thermal exceedance on DC link connectors. | verification-plan | Main |  | Test |  |  | Pulse load test replicates worst-case six-axis arm energisation. Continuous test confirms thermal rating of connectors and cabling at sustained surgical operation current. | verification, power-management, session-361, idempotency:ver-ifc045-ups-pdu-361
VER-MAIN-091 | Verify IFC-MAIN-046: Command Auxiliary Power Supply through charge, standby, and discharge modes via the Power Sequencing Controller discrete control signal; confirm mode transitions are acknowledged within 50 ms by measuring return status signal timing with an oscilloscope; repeat test with CAN bus disconnected to confirm graceful mode operation. | verification-plan | Main |  | Test |  |  | Timing test confirms mode acknowledgement meets the 50 ms requirement; CAN-bus disconnected test confirms discrete fallback operation critical to the safety argument that auxiliary circuits remain controllable during bus faults. | verification, power-management, session-361, idempotency:ver-ifc046-psc-aux-361
VER-MAIN-092 | Verify SUB-MAIN-089: Inject a 600 µA line-to-earth fault on each branch circuit in sequence under full operational load; confirm that the faulted branch is de-energised within 100 ms and a fault code is received by the Power Sequencing Controller within 105 ms; verify adjacent branches remain energised. | verification-plan | Main |  | Test |  |  | Tests the full leakage detection and branch isolation response chain per IEC 60601-1 clause 8.7.3; adjacent-branch isolation confirmation prevents common-mode removal of healthy circuits. | verification, power-management, safety, session-361, idempotency:ver-sub089-pdu-gf-361
VER-MAIN-093 | Verify SUB-MAIN-091: Simulate mains voltage collapse to below 80% of nominal under full surgical load using a programmable AC source; measure transfer time from voltage threshold breach to UPS Battery Module supplying load via scope triggers on PDU rail; confirm no subsystem supply interruption exceeds 30 ms. | verification-plan | Main |  | Test |  |  | End-to-end transfer time measurement confirms the 30 ms budget derived from joint servo controller fault tolerance; measurement at the PDU rail catches any delay in switching circuitry that bench tests of individual components might miss. | verification, power-management, failover, session-361, idempotency:ver-sub091-psc-mains-361
VER-MAIN-094 | Verify SUB-MAIN-090: Charge UPS Battery Module to 100%, then discharge at 20 A constant current while sampling state-of-charge from the battery management system interface at 1 Hz; compare against coulomb-counter reference at 20%, 50%, 80%, and 95% SoC points; confirm accuracy is within plus or minus 2% at each point and that the low-battery warning is asserted when SoC crosses 25%. | verification-plan | Main |  | Test |  |  | SoC accuracy test at multiple discharge points confirms calibration over the full operating range specified in SUB-MAIN-090; low-battery alert timing is verified at the exact 25% threshold. | verification, power-management, ups, session-361, idempotency:ver-sub090-ups-soc-361
VER-MAIN-095 | Verify SUB-MAIN-092: Apply 0%, 50%, and 100% of rated load to the Auxiliary Power Supply 24 VDC output; confirm output voltage remains within 23.52-24.48 V (plus or minus 2%) at each load point. Then disconnect both mains and UPS inputs and measure time until output drops below 23 V; confirm duration exceeds 20 minutes. | verification-plan | Main |  | Test |  |  | Load regulation test confirms the plus or minus 2% output accuracy across the full load range; battery endurance test at full rated load confirms the 20-minute minimum independent operation specified in SUB-MAIN-092. | verification, power-management, auxiliary-power, session-361, idempotency:ver-sub092-aux-psu-361
VER-MAIN-096 | Verify SUB-MAIN-033: Command the Instrument Drive Unit through full range on all four degrees of freedom (yaw, pitch, grip, wrist rotation) at 1Hz sweep using a calibrated instrument fixture; measure actuated angle/displacement at the tool tip with encoder feedback; pass criterion: <0.5mm positional error at tool tip, response within 80ms of command. | verification-plan | Main |  | Test |  |  | Integration test to verify that all four DOF of the Instrument Drive Unit meet the positional accuracy and response time required for precise surgical manipulation. | verification, surgical-instrument-system, instrument, session-362, idempotency:ver-sub-033-idu-dof-362
VER-MAIN-097 | Verify SUB-MAIN-036: Load a representative surgical path (laparoscopic dissection trajectory, 120 waypoints at 10Hz) into the Tool Tip Articulation Controller test harness; inject measured cable elongation offsets for a worn cable set; verify computed cable displacement commands track the reference trajectory with <0.3mm tip position error and complete computation within 8ms per cycle. | verification-plan | Main |  | Test |  |  | Integration test to verify the Tool Tip Articulation Controller compensates for real-world cable elongation and maintains sub-millimetre tip accuracy at clinical motion rates. | verification, surgical-instrument-system, articulation-controller, session-362, idempotency:ver-sub-036-ttac-362
VER-MAIN-098 | Verify SUB-MAIN-037: Present instruments with use counts at 0, 9, 10, and 11 cycles (boundary conditions) to the Instrument Lifecycle Controller coupling sequence; verify that coupling is permitted at 0 and 9 cycles, rejected with error code LC_EXCEEDED at 10 cycles, rejected at 11 cycles; verify rejection latency <200ms and rejection events are logged to the Procedure Data Recorder. | verification-plan | Main |  | Test |  |  | Integration test to verify the lifecycle enforcement boundary conditions, including the exact rejection threshold and audit trail — both are required for IEC 62061 and FDA 21 CFR Part 820 traceability of reprocessable instruments. | verification, surgical-instrument-system, lifecycle-controller, session-362, idempotency:ver-sub-037-ilc-362
VER-MAIN-099 | Verify SUB-MAIN-093: Commission an independent functional safety assessment by a competent body; review the Safety and Interlock Subsystem safety case against IEC 62061 requirements for SIL 3; evaluate PFH calculation for joint force limiting, E-stop chain, and communication watchdog functions; pass criterion: PFH <=1E-7/h confirmed by reliability analysis and independent reviewer sign-off. | verification-plan | Main |  | Analysis |  |  | SIL 3 compliance requires independent assessment per IEC 62061 clause 12; self-assessment is not acceptable for certification. Pass criterion is the IEC 62061 SIL 3 PFH boundary. | verification, sis, compliance, sil3, session-362, idempotency:ver-sub-093-sil3-362
VER-MAIN-100 | Verify PDR failover: With the system recording active (all video and telemetry channels), simulate primary storage medium failure by inducing a write fault via FMEA injection test. Confirm: (1) automatic failover to secondary path within 500ms measured from fault injection, (2) no data loss in the secondary recording stream, (3) surgeon display shows recording degradation warning within 1 second. | verification-plan | Main |  | Test |  |  | Verification of REQ-SESURGICALROBOT-094 (PDR redundancy). Storage failover must be tested under realistic intraoperative conditions including concurrent video and telemetry streams; passive inspection cannot verify failover timing or data continuity.
VER-MAIN-100 | Verify SUB-MAIN-095: Force a software watchdog timeout on the Console Computer test unit by injecting a process hang via fault injection tool; verify: (a) failover to backup path initiates within 500ms, (b) last commanded instrument position is preserved within 0.1mm, (c) audio tone and status message appear, (d) test repeated 10 times with no variance in failover time >50ms. | verification-plan | Main |  | Test |  |  | Fault injection testing of watchdog failover verifies the real-time response and state preservation guarantees under controlled failure conditions, with repeatability check to confirm no timing jitter. | verification, surgeon-console, redundancy, failover, session-362, idempotency:ver-sub-095-cc-failover-362
VER-MAIN-101 | Verify PMS safety domain redundant power: Apply load equivalent to full interlock subsystem and workspace safety enforcer draw to both DC supply rails simultaneously, then disconnect primary rail while measuring secondary switchover timing via oscilloscope probe on output bus. Confirm: (1) switchover within 5ms measured from voltage dip trigger, (2) no dropout on safety function power rail, (3) UPS battery sustains full load for ≥60 seconds before controlled shutdown. | verification-plan | Main |  | Test |  |  | Verification of REQ-SESURGICALROBOT-097 (PMS redundant power). Switchover timing must be verified under realistic load conditions; a 5ms criterion cannot be verified by inspection of design documents alone; UPS capacity test requires bench-level endurance measurement.
VER-MAIN-101 | Verify SUB-MAIN-096: Sever the Force Sensing Module communication link while system is in OPERATIONAL state; measure time from link loss to: (a) 0.3N braking force applied at master handle (target: <=50ms), (b) audio alert onset, (c) visual status change; confirm surgeon can still command full kinematic range with braking force applied; restore link and verify return to normal force rendering within 200ms. | verification-plan | Main |  | Test |  |  | Live fault injection test on integrated haptic hardware verifies degraded-mode entry timing and the 0.3N braking force value are within the specified limits, and that the surgeon retains kinematic control during the degraded state. | verification, haptic, degraded-mode, session-362, idempotency:ver-sub-096-haptic-fbd-362
VER-MAIN-102 | Verify SUB-MAIN-097: Using a network test harness, inject 100 command frames with valid HMAC, 50 with corrupted HMAC, and 20 replay frames to the Trajectory Generator and Motion Scaling Module; verify all valid frames are accepted, all invalid/replay frames are rejected with no arm motion, and rejection events appear in Procedure Data Recorder within 10ms of each rejection. | verification-plan | Main |  | Test |  |  | Active penetration testing of the HMAC authentication boundary verifies both the authentication logic and the security audit trail under representative attack scenarios. | verification, motion-control, cybersecurity, session-362, idempotency:ver-sub-097-hmac-362
VER-MAIN-103 | Verify SUB-MAIN-098: On the TTAC test bench, inject 1000 position command packets with valid HMAC signatures followed by 1000 packets with corrupted signatures. Verify that all valid packets are accepted within 1ms and all invalid packets are rejected within 1ms with corresponding Procedure Data Recorder log entries. Pass criteria: 100% acceptance of valid packets, 100% rejection of invalid packets, zero false negatives. | verification-plan | Main |  | Test |  |  | Integration test to verify cryptographic authentication compliance. Pass/fail criteria are binary (accept/reject with specified latency) and directly verify the authentication mandate from SYS-MAIN-018 as applied to this component. | verification, cybersecurity, session-365, idempotency:ver-SUB-MAIN-098-365
VER-MAIN-104 | Verify SUB-MAIN-099: On the Real-Time Compute Node, connect a packet injection harness to the KE input queue and inject 500 valid HMAC-authenticated command packets and 500 unauthenticated packets. Verify that all unauthenticated packets are rejected within one 1ms control cycle. Pass criteria: zero unauthenticated packets processed by IK solver, rejection latency <= 1ms for all invalid packets. | verification-plan | Main |  | Test |  |  | Integration test to verify cryptographic authentication compliance. Pass/fail criteria are binary (accept/reject with specified latency) and directly verify the authentication mandate from SYS-MAIN-018 as applied to this component. | verification, cybersecurity, session-365, idempotency:ver-SUB-MAIN-099-365
VER-MAIN-105 | Verify SUB-MAIN-100: Load the Trajectory Generator with a synthetic anatomy mesh defining a 100mm-radius keep-out sphere. Command 200 waypoints: 100 within 5mm of the keep-out boundary and 100 that violate the boundary by varying margins (5-50mm). Verify that all violating waypoints are rejected and safe-state transition occurs within 5ms. Pass criteria: 100% boundary-violating waypoint rejection, safe-state transition <= 5ms. | verification-plan | Main |  | Test |  |  | Integration test to verify cryptographic authentication compliance. Pass/fail criteria are binary (accept/reject with specified latency) and directly verify the authentication mandate from SYS-MAIN-018 as applied to this component. | verification, cybersecurity, session-365, idempotency:ver-SUB-MAIN-100-365
VER-MAIN-106 | Verify SUB-MAIN-101: Connect a PTP test injector to the inter-cart fibre link and transmit 1000 synchronisation frames: 500 with valid session token and 500 with corrupted or replayed tokens. Verify that all invalid frames are discarded within one synchronisation period and a communication fault event is generated. Pass criteria: 100% invalid frame discard, fault event generated for each invalid frame. | verification-plan | Main |  | Test |  |  | Integration test to verify cryptographic authentication compliance. Pass/fail criteria are binary (accept/reject with specified latency) and directly verify the authentication mandate from SYS-MAIN-018 as applied to this component. | verification, cybersecurity, session-365, idempotency:ver-SUB-MAIN-101-365
VER-MAIN-107 | Verify IFC-MAIN-002: With system energised and all three Emergency Stop Chain contact nodes monitored, open each node individually and confirm that the Power Management Subsystem servo drive contactors drop within 10ms and do not re-engage until the fault is cleared and operator reset is confirmed. Repeat across operating temperature range (10–40°C). | verification-plan | Main |  | Test |  |  | Hardwired E-stop interlock is the last hardware defence against uncontrolled joint motion. 10ms drop requirement derived from worst-case contact-open detection delay plus contactor release time; if contactors remain energised, drives remain powered and a faulty motion command can cause patient harm. | verification, safety, estop, session-366, idempotency:ver-ifc-002-estop-contactor-366
VER-MAIN-108 | Verify IFC-MAIN-002: With system energised and all three Emergency Stop Chain contact nodes monitored, open each node individually and confirm the Power Management Subsystem servo drive contactors drop within 10ms and do not re-engage until fault is cleared and operator reset is confirmed. | verification-plan | Main |  | Test |  |  | Hardwired E-stop interlock is the last hardware defence against uncontrolled joint motion. 10ms drop derived from worst-case contact-open detection delay plus contactor release time; failure to drop keeps drives powered and enables uncontrolled motion under a faulty command. | verification, safety, estop, session-366, idempotency:ver-ifc-002-estop-366
VER-MAIN-109 | Verify IFC-MAIN-007: With the heartbeat signal between Motion Control System and Safety and Watchdog System active, interrupt the dedicated hardwired line and confirm the Safety and Watchdog System declares SAFE-HOLD within 20ms. Confirm normal operation resumes within 500ms after reconnection. | verification-plan | Main |  | Test |  |  | Heartbeat monitors liveness of MCS to watchdog path; loss of heartbeat must trigger SAFE-HOLD to prevent unmonitored motion. 20ms limit keeps total detection-to-stop time under 50ms worst case. | verification, safety, watchdog, session-366, idempotency:ver-ifc-007-heartbeat-366
VER-MAIN-110 | Verify IFC-MAIN-021: Using a cable tension simulator, inject nominal, low-warning, and over-tension values on each of the cable channels and confirm the Safety and Interlock Subsystem receives the correct tension and status flags within the specified latency. Confirm lockout command is issued when tension exceeds threshold. | verification-plan | Main |  | Test |  |  | Cable overtension can snap instrument cables causing patient injury; this interface carries the primary measurement that drives lockout. Verification proves the data path and flag encoding are correct before integration with live surgical instruments. | verification, safety, instrument, session-366, idempotency:ver-ifc-021-cable-tension-366
VER-MAIN-111 | Verify IFC-MAIN-022: Command the Tool Tip Articulation Controller through a representative articulation sequence covering full pitch and yaw range. Measure cable displacement commands at the Instrument Drive Unit interface and confirm values match kinematic model predictions within 0.1mm at 200Hz update rate. | verification-plan | Main |  | Test |  |  | Correct cable displacement delivery is prerequisite for accurate instrument tip positioning; errors propagate directly to surgical accuracy. Kinematic model validation at the interface catches encoding or scaling errors before instrument assembly. | verification, instrument, session-366, idempotency:ver-ifc-022-ttac-idu-366
VER-MAIN-112 | Verify IFC-MAIN-023: Using test fixtures, trigger the Instrument Lifecycle Controller to issue lockout commands for instrument use-count exhaustion, sterility breach, and authentication failure. Confirm Safe State Manager receives each lockout code within 50ms and transitions system to SAFE-HOLD. Confirm lockout cannot be overridden without authorised credential exchange. | verification-plan | Main |  | Test |  |  | Instrument lifecycle lockout prevents reuse of expired or contaminated instruments. This interface carries safety-critical lockout commands; failure to deliver or act on them enables patient infection or mechanical failure from worn instruments. | verification, safety, instrument, session-366, idempotency:ver-ifc-023-ilc-ssm-366
VER-MAIN-113 | Verify IFC-MAIN-040: Inject a continuous 6-DOF Cartesian velocity command stream from the Console Computer via the Inter-Cart Fibre Link test port. Confirm commands arrive at the patient-cart Real-Time Protocol Engine at 1kHz with end-to-end latency under 5ms and zero frame drops over a 10-minute continuous run. | verification-plan | Main |  | Test |  |  | The command stream from console to patient cart is the primary teleoperation data path; latency above 5ms degrades surgeon feel and tracking precision, and dropped frames cause motion discontinuities that risk tissue damage. | verification, comms, motion-control, session-366, idempotency:ver-ifc-040-console-icfl-366
VER-MAIN-114 | Verify SUB-MAIN-103: On target hardware with peak traffic load (21 kinematics channels at 1kHz, dual stereo HD video), measure one-way fibre path latency using hardware timestamps on transmit and receive FPGAs over 10,000 consecutive frames. Pass: P99 latency <= 500us. Fail: any frame exceeds 500us. | verification-plan | Main |  | Test |  |  | P99 measurement over 10,000 frames characterises worst-case latency including interrupt service variability. Hardware timestamps eliminate OS jitter from the measurement. | verification, comms, inter-cart-fibre, session-369, idempotency:ver-sub-main-103-369
VER-MAIN-115 | Verify SUB-MAIN-104: With system in OPERATIONAL state and kinematic commands flowing, disconnect primary fibre link using relay-controlled break fixture. Measure time from physical link break to redundant path carrying live traffic using oscilloscope probes on both paths. Pass: switchover <= 5ms with no kinematic command gap > 100ms. Perform 20 repeated trials. | verification-plan | Main |  | Test |  |  | Relay-controlled break ensures repeatable fault injection. Oscilloscope measurement captures FPGA-level switchover independent of host software latency. 20 trials provide statistical confidence on switchover time distribution. | verification, comms, inter-cart-fibre, session-369, idempotency:ver-sub-main-104-369
VER-MAIN-116 | Verify SUB-MAIN-107: Inject bit errors at rates of 0, 1e-9, 1e-7, and 1e-5 on primary link using BERT tester. Verify link state transitions to DEGRADED at 1e-7 and FAILED at 1e-5 within two 10ms monitoring cycles. Verify state is reported via LVDS to SIS. Pass: correct state classification within 20ms of threshold crossing. | verification-plan | Main |  | Test |  |  | BERT injection provides controlled, repeatable bit error rates. Two monitoring cycles (20ms) reflects the 100Hz poll rate, ensuring classification occurs within two expected sample periods. | verification, comms, network-management, session-369, idempotency:ver-sub-main-107-369
VER-MAIN-117 | Verify IFC-MAIN-047: Inject primary fibre link failure by toggling relay. Measure time from relay activation to LVDS COMM_FAULT signal on SIS input using oscilloscope. Pass: LVDS signal asserted within 2ms of relay activation across 10 consecutive injection trials. | verification-plan | Main |  | Test |  |  | Oscilloscope measurement on LVDS line captures hardware-layer timing independent of software. 10 trials verify consistency of the FPGA-based fault detection path. | verification, comms, network-management, safety, session-369, idempotency:ver-ifc-047-nmc-sis-369
VER-MAIN-118 | Verify IFC-MAIN-048: At peak 1kHz kinematics load, measure DMA delivery latency from RTPE frame completion to Procedure Data Recorder buffer write using kernel tracepoints. Pass: P99 latency <= 100us over 60 seconds. Verify RTPE jitter does not increase when PDR is recording vs idle. | verification-plan | Main |  | Test |  |  | Kernel tracepoints measure the DMA path without requiring hardware instrumentation. RTPE jitter comparison verifies the recording path does not interfere with the hard real-time control loop. | verification, comms, real-time-protocol-engine, procedure-data-recorder, session-369, idempotency:ver-ifc-048-rtpe-pdr-369
VER-MAIN-119 | Verify SUB-MAIN-111: Immerse 3 samples each of Sterile Adapter, Cable Tensioning System housing, and representative surgical instrument in IPA 70% for 30 minutes. Post-immersion: verify sterile barrier integrity per ISO 11607-1, measure cable tensioning force variation (±5% nominal tolerance), and visually inspect for cracking, delamination, or discolouration. All three pass criteria must be met for compliance. | verification-plan | Main |  | Test |  |  | Immersion test replicates worst-case field disinfection practice. The three-component scope covers the components most exposed to disinfectant ingress during intraoperative instrument exchange. | verification, surgical-instrument, sterility, session-373, idempotency:ver-sub-111-ipa-373
VER-MAIN-120 | Verify SUB-MAIN-117: Submit the Power Management Subsystem to a UKAS-accredited test laboratory for IEC 60601-1:2005+AMD1:2012 dielectric withstand and leakage current testing. Pass criterion: withstand at 4000V AC for 1 minute with no breakdown; leakage below 500uA normal, 1000uA single-fault. | verification-plan | Main |  | Test |  |  | Third-party accredited testing is required for regulatory submission under EU MDR and FDA 510k. Pass/fail criteria match the standard's limits for Class I medical electrical equipment. | verification, power-management, compliance, session-374, idempotency:ver-sub-main-117-374
VER-MAIN-121 | Verify SUB-MAIN-119: Perform fault tree analysis and FMEA for the Workspace Safety Enforcer's boundary enforcement functions. Calculate PFHd from component reliability data and architectural independence analysis. Pass criterion: PFHd <= 1E-7 per hour; architectural analysis confirms independence from non-safety channels. | verification-plan | Main |  | Analysis |  |  | SIL 2 demonstration under IEC 62061 requires documented fault tree analysis and FMEA with quantified PFHd. Hardware testing alone cannot demonstrate a 1E-7/hour failure rate; probabilistic analysis is the accepted method for safety integrity verification. | verification, motion-control, workspace-safety, safety, session-374, idempotency:ver-sub-main-119-374
VER-MAIN-122 | Verify SUB-MAIN-120: Inject primary TSN link failure (disconnect Ethernet cable) while Real-Time Protocol Engine is under full motion load at 1kHz. Measure time from link loss detection to CAN FD backup path command delivery. Pass criterion: failover < 5ms; no motion command gap > 20ms; motion resumes on backup path. | verification-plan | Main |  | Test |  |  | Live failover injection test is the only way to verify the 5ms and 20ms timing requirements under realistic load. Simulation cannot capture real OS scheduling latency and hardware detection timing. | verification, comms, real-time-protocol-engine, redundancy, session-374, idempotency:ver-sub-main-120-374
VER-MAIN-123 | Verify SUB-MAIN-121: Inject Haptic Controller process failure (SIGKILL) while providing active force feedback at 20Hz. Measure time from primary failure to first command from standby process at master handle actuator. Pass criterion: switchover <= 10ms; no perceptible force discontinuity > 0.05N at handle. | verification-plan | Main |  | Test |  |  | Haptic switchover must be verified under operational conditions because OS process scheduling under load differs from idle state. The 0.05N perceptible threshold is derived from psychophysics literature on force JND at the fingertip during fine manipulation. | verification, haptic, redundancy, session-374, idempotency:ver-sub-main-121-374
VER-MAIN-124 | Verify SUB-MAIN-122: Using an optometer and calibrated test patterns, measure vergence distance and accommodation stimulus at 5 representative display content depths. Adjust inter-ocular distance across 58-72mm range. Pass criterion: VAC < 0.6 dioptre at all depths; IOD adjustment smooth and repeatable to 1mm. | verification-plan | Main |  | Test |  |  | VAC measurement requires optical instrumentation; human subject testing with surgeons is impractical at verification stage. 0.6 dioptre threshold is from ISO 9241-302 and published stereoscopic ergonomics literature. | verification, vision, eye, ergonomics, session-374, idempotency:ver-sub-main-122-374
VER-MAIN-125 | Verify SUB-MAIN-126: Using a network packet injector, send 10,000 command messages with invalid HMAC-SHA256 tags to the Communication and Data Management System during active operation. Measure rejection latency per message and count of rejected messages. Pass criterion: all 10,000 rejected within 1ms; SAFE_HOLD triggered after the third consecutive failure; each rejection logged as a security event. | verification-plan | Main |  | Test |  |  | Authentication rejection must be verified at scale to confirm the system is not vulnerable to replay attacks or timing side-channels. 10,000 injected messages stress the authentication pipeline at operational data rates. SAFE_HOLD trigger verification confirms the escalation path is functional. | verification, comms, cybersecurity, session-374, idempotency:ver-sub-main-126-374
VER-MAIN-126 | Verify SYS-MAIN-019: Submit the complete Surgical Robot System to an accredited EMC test laboratory. Conduct: (a) Radiated emissions measurement per CISPR 11 in a semi-anechoic chamber, confirming Group 1 Class B limits at 30–1000MHz; (b) Conducted emissions on mains supply lines per CISPR 11; (c) Radiated RF immunity per IEC 61000-4-3 at 10V/m across 80MHz–2.7GHz; (d) Conducted RF immunity per IEC 61000-4-6 at 3Vrms across 150kHz–80MHz; (e) Electrostatic discharge per IEC 61000-4-2 at 4kV contact and 8kV air. During immunity tests, the system SHALL maintain full motion control with position error below 0.5mm; no spurious E-stop activations are permitted. Pass criterion: no essential performance degradation during immunity, all emission limits met. | verification-plan | Main |  | Test |  |  | Accredited laboratory testing with UKAS/ILAC accreditation is required for CE marking under MDR 2017/745 and cannot be replaced by in-house bench testing. The specific pass criteria (0.5mm position error, no spurious E-stop) define what constitutes essential performance per IEC 60601-1-2 Annex A for this system class. | verification, emc, regulatory, system, validation, session-377
VER-MAIN-127 | Verify SYS-MAIN-019: Submit the complete Surgical Robot System to an accredited EMC test laboratory. Conduct radiated emissions per CISPR 11 Group 1 Class B at 30-1000MHz, conducted emissions on mains supply lines, radiated RF immunity per IEC 61000-4-3 at 10V/m across 80MHz-2.7GHz, and ESD per IEC 61000-4-2 at 4kV contact. During immunity tests the system SHALL maintain full motion control with position error below 0.5mm and produce no spurious E-stop activations. | verification-plan | Main |  | Test |  |  | Accredited laboratory testing is required for CE marking under MDR 2017/745. The 0.5mm position error pass criterion defines essential performance per IEC 60601-1-2 Annex A for this system class; no spurious E-stop criterion ensures the immunity testing environment does not trigger the safety system. | verification, emc, regulatory, system, validation, session-377
VER-MAIN-128 | Verify SUB-MAIN-127: Mount all three instrument arms on the patient-side cart. Command arms through a choreographed convergence trajectory that reduces pairwise inter-arm clearance from 100mm to 10mm at 50mm/s approach velocity. Verify: (a) halt command issued to both approaching arms when predicted clearance reaches 25mm, measured from logged kinematics at 100Hz sample points; (b) all arm motion stopped within 50ms of halt command issue; (c) final measured clearance no less than 15mm with load cell confirmation. Repeat for all three arm-pair combinations. | verification-plan | Main |  | Test |  |  | Hardware-in-the-loop test on the actual patient-side cart mechanical assembly is required because inter-arm clearance depends on the precise kinematic model including cable deflection under load; simulation alone is insufficient. All three arm-pair combinations must be tested because asymmetric cable routing creates different compliance characteristics per pair. | verification, motion-control, collision-avoidance, safety, validation, session-377
VER-MAIN-129 | Verify SUB-MAIN-128: Activate each defined alarm condition in turn on a fully integrated surgeon console test bench. For each condition: confirm the correct priority level is assigned, verify visual signal colour and flash rate match IEC 60601-1-8 Table 6 for the assigned priority, verify auditory signal pattern matches IEC 60601-1-8 Annex F, then disconnect mains and confirm alarm continues to signal on internal battery within 5 seconds. Confirm no auditory alarm condition produces a signal below 65 dB(A) at 1 metre operator distance. | verification-plan | Main |  | Test |  |  | IEC 60601-1-8 Clause 6.8 requires each alarm signal to be verified by functional test against the standard tables. The 65 dB(A) minimum level ensures alarms are audible over typical OR background noise (55-60 dB(A)). Battery continuity must be tested in the integrated system rather than simulation because alarm circuitry involves the display, audio, and power management subsystems together. | verification, surgeon-console, alarm-management, regulatory, validation, session-377

## Trace Links
Source Ref | Target Ref | Link Type | Description | Rationale | Created
SYS-MAIN-019 | VER-MAIN-126 | verifies | SYS-MAIN-019 → VER-MAIN-126: EMC compliance requirement verified by accredited lab test campaign | SYS-MAIN-019 mandates IEC 60601-1-2:2014 EMC compliance; VER-MAIN-126 is the comprehensive EMC test campaign (CISPR 11 emissions and IEC 61000-4 immunity) that directly verifies this requirement. | 2026-03-20T14:05:39.742Z
SUB-MAIN-128 | VER-MAIN-129 | derives | Alarm management requirement verified by per-condition IEC 60601-1-8 compliance test | SUB-MAIN-128 specifies IEC 60601-1-8 compliance, three priority levels, visual and auditory signal characteristics, and battery continuation; VER-MAIN-129 tests each of those attributes against the standard tables for every alarm condition defined in the system. | 2026-03-20T13:28:21.386Z
SUB-MAIN-127 | VER-MAIN-128 | derives | Inter-arm collision avoidance requirement verified by HIL convergence test | SUB-MAIN-127 specifies 25mm warning threshold, 15mm hard limit, 100Hz check rate, and 50ms halt time; VER-MAIN-128 directly tests each of those four numeric constraints with instrumentated measurement. | 2026-03-20T13:28:13.744Z
SYS-MAIN-019 | VER-MAIN-127 | derives | EMC system requirement is verified by accredited lab test | SYS-MAIN-019 specifies IEC 60601-1-2 compliance; VER-MAIN-127 specifies the test method, measured parameters, and pass criteria that directly demonstrate that compliance. | 2026-03-20T13:27:58.982Z
SYS-MAIN-002 | SUB-MAIN-128 | derives | System single-point failure safe state requirement drives alarm management framework | SYS-MAIN-002 requires any single-point failure to result in a safe state within 250ms; alarm management per IEC 60601-1-8 is the mechanism by which the surgeon receives reliable, prioritised notification of these failure events, ensuring surgical response complements the automatic safe state transition. | 2026-03-20T13:27:13.877Z
SYS-MAIN-002 | SUB-MAIN-127 | derives | Single-point failure safe-state requirement derives to inter-arm collision avoidance | SYS-MAIN-002 requires safe state on any single-point failure including sensor fault; an undetected arm proximity event is a failure mode requiring the motion control subsystem to implement the specific inter-arm collision detection and halt mechanism specified in SUB-MAIN-127. | 2026-03-20T13:26:58.324Z
STK-MAIN-003 | SYS-MAIN-019 | derives | EMC compliance requirement derives from OR infrastructure integration stakeholder need | STK-MAIN-003 requires integration with existing OR infrastructure including co-located electrical equipment; IEC 60601-1-2 compliance is the specific engineering response to that integration constraint, ensuring the robot does not interfere with or suffer interference from co-located devices. | 2026-03-20T13:26:31.450Z
SYS-MAIN-018 | SUB-MAIN-103 | derives | HMAC-SHA256 implementation of SYS-level cryptographic authentication mandate | REQ-SESURGICALROBOT-103 directly decomposes SYS-MAIN-018 by specifying the algorithm (HMAC-SHA256), the interface scope (inter-cart fibre link), and the failure mode response (SAFE-HOLD within 50ms). This is a genuine derivation: the SUB requirement exists because the SYS requirement mandates it. | 2026-03-20T13:07:41.272Z
SYS-MAIN-002 | SUB-MAIN-102 | derives | Regulatory compliance basis for workspace safety enforcement | SYS-MAIN-002 mandates fault detection and safe-state achievement. The compliance requirement REQ-SESURGICALROBOT-102 derives from this by specifying the regulatory framework (IEC 60601-1, ISO 14971) governing the Workspace Safety Enforcer, which is a key component in the safe-state enforcement chain. | 2026-03-20T13:07:21.352Z
SYS-MAIN-001 | SUB-MAIN-116 | derives | MCS distributed electronics chassis derives from motion latency requirement | SYS-MAIN-001 requires <1ms end-to-end latency; SUB-MAIN-116 specifies co-location of RTCN, Trajectory Generator, and MSM on a common backplane to meet this latency budget by eliminating inter-chassis cable propagation. | 2026-03-20T12:43:06.184Z
SYS-MAIN-001 | SUB-MAIN-115 | derives | RTCN physical co-location derives from 1ms motion latency requirement | SYS-MAIN-001 requires <1ms end-to-end motion latency; SUB-MAIN-115 specifies the VITA 57.1 FMC form factor and vibration-isolated mounting that minimises inter-module propagation delay to meet the 1ms budget. | 2026-03-20T12:42:59.107Z
SYS-MAIN-002 | SUB-MAIN-114 | derives | PMS physical housing derives from power fault resilience requirement | SYS-MAIN-002 requires recovery from power faults; SUB-MAIN-114 specifies the sealed electronics bay and field-replaceable assembly design that enables power fault isolation and rapid recovery at the subsystem level. | 2026-03-20T12:42:58.450Z
SYS-MAIN-015 | SUB-MAIN-113 | derives | PDR physical embodiment (duplicate - see SUB-MAIN-112) | Created in error as a duplicate of SUB-MAIN-112 during the same session. To be tagged as duplicate and removed during next QC pass. | 2026-03-20T12:42:50.106Z
SYS-MAIN-015 | SUB-MAIN-112 | derives | PDR physical form derives from system-level recording requirement | SYS-MAIN-015 requires recording all kinematic data and video streams; SUB-MAIN-112 specifies the physical form factor (2U rack LRU, RAID-1, IP32) needed to reliably fulfill that function in the clinical environment. | 2026-03-20T12:42:49.387Z
SUB-MAIN-126 | VER-MAIN-125 | verifies | Packet injection test for CDMS cryptographic authentication | Scale injection testing at 10,000 messages verifies timing compliance and SAFE_HOLD escalation, which cannot be confirmed by inspection of code alone. | 2026-03-20T12:42:07.421Z
SUB-MAIN-122 | VER-MAIN-124 | verifies | Optometric VAC measurement for Stereoscopic Display System | VAC below 0.6 dioptre is a human factors safety requirement; optical measurement is the only instrument-based method for pre-human-trial verification. | 2026-03-20T12:41:48.734Z
SUB-MAIN-121 | VER-MAIN-123 | verifies | Hot-standby switchover test for Haptic Feedback Subsystem | 10ms switchover can only be verified by live failure injection; psychophysics-derived force JND threshold confirms perceptual continuity. | 2026-03-20T12:41:47.307Z
SUB-MAIN-120 | VER-MAIN-122 | verifies | Live failover injection test for RTPE dual-path redundancy | 5ms failover timing can only be verified through physical link-failure injection under operational load conditions. | 2026-03-20T12:41:32.980Z
SUB-MAIN-119 | VER-MAIN-121 | verifies | SIL 2 FTA and FMEA analysis for Workspace Safety Enforcer | PFHd analysis is the mandated IEC 62061 method for demonstrating SIL 2 compliance; cannot be verified by inspection or functional test alone. | 2026-03-20T12:41:18.836Z
SUB-MAIN-117 | VER-MAIN-120 | verifies | IEC 60601-1 accredited lab test for Power Management Subsystem | Accredited testing is the only accepted means of demonstrating compliance with IEC 60601-1 leakage and withstand requirements under EU MDR and FDA 510k review. | 2026-03-20T12:41:05.977Z
SYS-MAIN-014 | SUB-MAIN-125 | derives | Console disinfectant compatibility derives from system-level surface compatibility req | SYS-MAIN-014 mandates compatibility with IPA and other disinfectants for patient-side surfaces; SUB-MAIN-125 extends this to the surgeon console and specifies the three most common OR disinfectant chemistries and the 3000-cycle service life. | 2026-03-20T12:40:40.361Z
SYS-MAIN-004 | SUB-MAIN-121 | derives | HFS redundancy derives from instrument force sensing system req | SYS-MAIN-004 requires continuous force measurement; SUB-MAIN-121 ensures haptic subsystem availability through hot standby, preventing force feedback loss that would leave the surgeon without tissue resistance sensing during a procedure. | 2026-03-20T12:40:39.710Z
SYS-MAIN-002 | SUB-MAIN-120 | derives | RTPE redundancy derives from single-point failure detection system req | SYS-MAIN-002 requires detection and recovery from any single-point failure; SUB-MAIN-120 implements dual-path redundancy in the RTPE to prevent communication loss from becoming a single point of failure in the motion control chain. | 2026-03-20T12:40:27.029Z
SYS-MAIN-018 | SUB-MAIN-126 | derives | CDMS cryptographic auth derives from system-level message authentication req | SYS-MAIN-018 mandates authentication of safety-critical inter-subsystem commands; SUB-MAIN-126 specifies the HMAC-SHA256/ECDH-P384 mechanism, 1ms rejection latency, and SAFE_HOLD trigger at the CDMS subsystem level. | 2026-03-20T12:40:26.378Z
SYS-MAIN-003 | SUB-MAIN-122 | derives | Stereoscopic display VAC derives from 3D HD video system req | SYS-MAIN-003 specifies the stereoscopic 3D HD video requirement; SUB-MAIN-122 operationalises the human factors requirements at the display subsystem level, specifically the VAC threshold needed to ensure surgeons can use the system for extended procedures. | 2026-03-20T12:40:16.445Z
SYS-MAIN-007 | SUB-MAIN-124 | derives | MHA force reflection derives from master manipulator system req | SYS-MAIN-007 requires transmission of surgeon commands via master manipulator; SUB-MAIN-124 specifies the bandwidth, dynamic range, and resolution needed at the MHA to faithfully convey surgeon intent and receive force feedback. | 2026-03-20T12:40:15.819Z
SYS-MAIN-001 | SUB-MAIN-123 | derives | MSM scaling ratio workflow derives from motion scaling system req | SYS-MAIN-001 mandates the scaling ratio range; SUB-MAIN-123 captures the surgeon-facing workflow for safe ratio adjustment, directly implementing the system-level capability. | 2026-03-20T12:39:54.878Z
SYS-MAIN-001 | SUB-MAIN-118 | derives | MSM IEC 80601-2-77 compliance derives from motion scaling system req | SYS-MAIN-001 requires 1:1 to 10:1 scaling; SUB-MAIN-118 operationalises this for the MSM with regulatory compliance, confirmation workflow, and increment specification. | 2026-03-20T12:39:54.238Z
SUB-MAIN-111 | VER-MAIN-119 | verifies | IPA immersion test for SUB-MAIN-111 sterile-field disinfectant compatibility | A 30-minute immersion test at 70% IPA is the minimum test condition to verify the disinfectant resistance required by SUB-MAIN-111; pass/fail criteria on structural integrity, force variation, and visual inspection are sufficient to demonstrate compliance | 2026-03-20T12:24:57.199Z
SYS-MAIN-006 | SUB-MAIN-111 | derives | Sterile-field IPA compatibility derives from system disinfectant compatibility mandate | SYS-MAIN-006 applies to all components in or near the sterile field; SUB-MAIN-111 derives the specific disinfectant concentration, contact time, and performance acceptance criteria for the Surgical Instrument System components that experience the most direct and prolonged fluid contact | 2026-03-20T12:24:55.216Z
SUB-MAIN-097 | VER-MAIN-101 | verifies | PMS safety domain redundant power verification for REQ-097 | REQ-SESURGICALROBOT-097 requires redundant power paths for the safety-critical domain; REQ-SESURGICALROBOT-101 tests compliance by applying full interlock load on each path independently and measuring the other path's response | 2026-03-20T12:22:55.166Z
SUB-MAIN-094 | VER-MAIN-100 | verifies | PDR failover verification test for REQ-094 auto-failover requirement | REQ-SESURGICALROBOT-094 requires failover within a specified window; REQ-SESURGICALROBOT-100 provides a hardware-injected storage failure test with measured failover timing to demonstrate compliance against that threshold | 2026-03-20T12:22:54.549Z
SYS-MAIN-005 | SUB-MAIN-097 | derives | PMS redundant safety-critical power path derives from 60-second holdover mandate | SYS-MAIN-005 requires 60s controlled operation after mains power loss; REQ-SESURGICALROBOT-097 derives the dedicated dual-path requirement for the safety-critical domain (interlock + watchdog) to ensure that the safety function remains powered even if the general DC bus collapses | 2026-03-20T12:22:45.673Z
SYS-MAIN-004 | SUB-MAIN-096 | derives | Haptic actuator channel failover derives from force feedback fidelity requirement | SYS-MAIN-004 requires ≤0.1N force resolution at the master manipulator; REQ-SESURGICALROBOT-096 derives the degraded-mode requirement specifying that on actuator channel failure the system maintains 50% axis count at ≤0.2N degraded resolution rather than losing force feedback entirely | 2026-03-20T12:22:45.028Z
SYS-MAIN-007 | SUB-MAIN-095 | derives | Time Protocol Engine redundancy derives from inter-subsystem synchronisation mandate | SYS-MAIN-007 requires consistent sub-ms synchronisation across subsystems; REQ-SESURGICALROBOT-095 derives the TCXO holdover requirement ensuring the timing reference remains valid when GPS/PTP disciplining fails, preserving synchronisation compliance | 2026-03-20T12:22:35.509Z
SYS-MAIN-006 | SUB-MAIN-093 | derives | PMS Type BF galvanic isolation derives from sterile-field patient safety requirement | SYS-MAIN-006 covers components in contact with the patient environment; REQ-SESURGICALROBOT-093 derives the IEC 60601-1 Type BF isolation requirement to prevent leakage current exceeding 100µA from reaching patient-applied parts | 2026-03-20T12:22:34.884Z
SYS-MAIN-008 | SUB-MAIN-092 | derives | Motion Scaling Module DHF requirement derives from selectable scaling mandate | SYS-MAIN-008 requires three distinct selectable scaling ratios that must be verified clinically; a DHF under 21 CFR 820.30 is required to document the design process, FMEA, and validation records for this motion-critical, patient-affecting function | 2026-03-20T12:22:26.051Z
SYS-MAIN-012 | SUB-MAIN-091 | derives | Workspace Safety Enforcer SIL certification derives from 5N tissue force limit mandate | SYS-MAIN-012 places a hard 5N limit on instrument-tissue force, a safety-critical function; REQ-SESURGICALROBOT-091 derives the SIL 2 certification requirement as the regulatory means to assure the integrity of the enforcement function that implements this limit | 2026-03-20T12:22:25.419Z
SYS-MAIN-002 | SUB-MAIN-090 | derives | Motion Control System IEC 62304 software lifecycle derives from single-point failure detection mandate | SYS-MAIN-002 requires detection and safe response to any single-point failure including software exceptions; IEC 62304 compliance on the MCS is the regulatory mechanism to ensure software failure modes are systematically identified, classified, and mitigated | 2026-03-20T12:22:14.731Z
SYS-MAIN-007 | SUB-MAIN-089 | derives | Time Compute Node M.2 hardware module derives from inter-subsystem latency requirement | SYS-MAIN-007 requires surgeon commands to reach instrument tip within 100ms and specifies synchronisation; REQ-SESURGICALROBOT-089 defines the dedicated TCXO-based timing hardware module that provides the precision time reference to meet this synchronisation budget | 2026-03-20T12:22:14.108Z
SYS-MAIN-001 | SUB-MAIN-088 | derives | MCS physical architecture derives from <1ms end-to-end latency mandate | SYS-MAIN-001 requires <1ms end-to-end control loop latency; REQ-SESURGICALROBOT-088 specifies the dedicated real-time compute hardware physically separated from non-real-time processing to achieve deterministic cycle times at 1kHz | 2026-03-20T12:21:58.970Z
SYS-MAIN-005 | SUB-MAIN-087 | derives | PMS physical LRU separation derives from post-power-loss operational mandate | SYS-MAIN-005 requires 60s operational capability after power loss; REQ-SESURGICALROBOT-087 defines the two-LRU physical architecture (mains entry + DC distribution) with battery backup to achieve that isolation and holdover | 2026-03-20T12:21:58.288Z
SYS-MAIN-015 | SUB-MAIN-086 | derives | PDR physical housing requirement derives from system recording mandate | SYS-MAIN-015 requires continuous 1kHz kinematic and video recording; REQ-SESURGICALROBOT-086 specifies the physical 1U enclosure and integration into the patient-side cart to satisfy that recording function within the operating environment | 2026-03-20T12:21:49.794Z
SYS-MAIN-015 | SUB-MAIN-086 | derives | PDR physical housing requirement derives from system recording mandate | SYS-MAIN-015 requires continuous 1kHz kinematic and video recording; REQ-SESURGICALROBOT-086 specifies the physical 1U enclosure and integration into the patient-side cart to satisfy that recording function within the operating environment | 2026-03-20T12:21:44.682Z
SYS-MAIN-003 | SUB-MAIN-099 | derives | System video specification → VIS subsystem delivery requirement | STK-MAIN-012 surgeon situational awareness and SYS-MAIN-003 stereoscopic video both drive REQ-099 which decomposes into the VIS subsystem delivery specification with quantified latency, resolution, and continuity thresholds. | 2026-03-20T12:07:50.910Z
SYS-MAIN-018 | SUB-MAIN-098 | derives | System-level crypto auth mandate → CDMS HMAC implementation | SYS-MAIN-018 requires cryptographic authentication on all safety-critical inter-subsystem command interfaces system-wide. REQ-098 is the direct subsystem decomposition implementing HMAC-SHA-256 in the CDMS, which is the communication backbone for those interfaces. | 2026-03-20T12:07:50.303Z
SYS-MAIN-002 | SUB-MAIN-094 | derives | SYS fault coverage → PDR redundancy | PDR redundancy requirement (REQ-086) derives from the single-point failure coverage mandate in SYS-MAIN-002: PDR storage medium is a potential single-point failure in the recording path; failover within 500ms stays within the 250ms safe-state budget for safety functions while preserving data integrity. | 2026-03-20T12:07:49.649Z
SUB-MAIN-077 | VER-MAIN-085 | verifies | Verification procedure for SUB-MAIN-077 | REQ-SESURGICALROBOT-085 provides the test method to verify SUB-MAIN-077 meets its measurable acceptance criteria. | 2026-03-20T11:52:27.988Z
SUB-MAIN-086 | VER-MAIN-084 | verifies | Verification procedure for SUB-MAIN-086 | REQ-SESURGICALROBOT-084 provides the test method to verify SUB-MAIN-086 meets its measurable acceptance criteria. | 2026-03-20T11:52:27.365Z
SUB-MAIN-110 | VER-MAIN-083 | verifies | Verification procedure for SUB-MAIN-110 | REQ-SESURGICALROBOT-083 provides the test method to verify SUB-MAIN-110 meets its measurable acceptance criteria. | 2026-03-20T11:52:26.746Z
SUB-MAIN-109 | VER-MAIN-082 | verifies | Verification procedure for SUB-MAIN-109 | REQ-SESURGICALROBOT-082 provides the test method to verify SUB-MAIN-109 meets its measurable acceptance criteria. | 2026-03-20T11:52:26.138Z
SUB-MAIN-085 | VER-MAIN-081 | verifies | Verification procedure for SUB-MAIN-085 | REQ-SESURGICALROBOT-081 provides the test method to verify SUB-MAIN-085 meets its measurable acceptance criteria. | 2026-03-20T11:52:25.531Z
SUB-MAIN-080 | VER-MAIN-080 | verifies | Verification procedure for SUB-MAIN-080 | REQ-SESURGICALROBOT-080 provides the test method to verify SUB-MAIN-080 meets its measurable acceptance criteria. | 2026-03-20T11:52:24.897Z
SUB-MAIN-073 | VER-MAIN-079 | verifies | Verification procedure for SUB-MAIN-073 | REQ-SESURGICALROBOT-079 provides the test method to verify SUB-MAIN-073 meets its measurable acceptance criteria. | 2026-03-20T11:52:24.289Z
SUB-MAIN-023 | VER-MAIN-078 | verifies | Verification procedure for SUB-MAIN-023 | REQ-SESURGICALROBOT-078 provides the test method to verify SUB-MAIN-023 meets its measurable acceptance criteria. | 2026-03-20T11:52:23.670Z
SUB-MAIN-030 | VER-MAIN-077 | verifies | Verification procedure for SUB-MAIN-030 | REQ-SESURGICALROBOT-077 provides the test method to verify SUB-MAIN-030 meets its measurable acceptance criteria. | 2026-03-20T11:50:49.604Z
SUB-MAIN-049 | VER-MAIN-076 | verifies | Verification procedure for SUB-MAIN-049 | REQ-SESURGICALROBOT-076 provides the test method to verify SUB-MAIN-049 meets its measurable acceptance criteria. | 2026-03-20T11:50:48.976Z
SUB-MAIN-048 | VER-MAIN-075 | verifies | Verification procedure for SUB-MAIN-048 | REQ-SESURGICALROBOT-075 provides the test method to verify SUB-MAIN-048 meets its measurable acceptance criteria. | 2026-03-20T11:50:48.276Z
SUB-MAIN-046 | VER-MAIN-074 | verifies | Verification procedure for SUB-MAIN-046 | REQ-SESURGICALROBOT-074 provides the test method to verify SUB-MAIN-046 meets its measurable acceptance criteria. | 2026-03-20T11:50:47.634Z
SUB-MAIN-045 | VER-MAIN-073 | verifies | Verification procedure for SUB-MAIN-045 | REQ-SESURGICALROBOT-073 provides the test method to verify SUB-MAIN-045 meets its measurable acceptance criteria. | 2026-03-20T11:50:46.973Z
SUB-MAIN-044 | VER-MAIN-072 | verifies | Verification procedure for SUB-MAIN-044 | REQ-SESURGICALROBOT-072 provides the test method to verify SUB-MAIN-044 meets its measurable acceptance criteria. | 2026-03-20T11:50:46.324Z
SUB-MAIN-042 | VER-MAIN-071 | verifies | Verification procedure for SUB-MAIN-042 | REQ-SESURGICALROBOT-071 provides the test method to verify SUB-MAIN-042 meets its measurable acceptance criteria. | 2026-03-20T11:50:45.678Z
SUB-MAIN-041 | VER-MAIN-070 | verifies | Verification procedure for SUB-MAIN-041 | REQ-SESURGICALROBOT-070 provides the test method to verify SUB-MAIN-041 meets its measurable acceptance criteria. | 2026-03-20T11:50:45.065Z
SUB-MAIN-040 | VER-MAIN-069 | verifies | Verification procedure for SUB-MAIN-040 | REQ-SESURGICALROBOT-069 provides the test method to verify SUB-MAIN-040 meets its measurable acceptance criteria. | 2026-03-20T11:50:44.431Z
SUB-MAIN-065 | VER-MAIN-068 | verifies | Verification procedure for SUB-MAIN-065 | REQ-SESURGICALROBOT-068 provides the test method to verify SUB-MAIN-065 meets its measurable acceptance criteria. | 2026-03-20T11:50:43.817Z
SUB-MAIN-053 | VER-MAIN-067 | verifies | Verification procedure for SUB-MAIN-053 | REQ-SESURGICALROBOT-067 provides the test method to verify SUB-MAIN-053 meets its measurable acceptance criteria. | 2026-03-20T11:50:43.155Z
SUB-MAIN-052 | VER-MAIN-066 | verifies | Verification procedure for SUB-MAIN-052 | REQ-SESURGICALROBOT-066 provides the test method to verify SUB-MAIN-052 meets its measurable acceptance criteria. | 2026-03-20T11:50:42.524Z
SUB-MAIN-050 | VER-MAIN-065 | verifies | Verification procedure for SUB-MAIN-050 | REQ-SESURGICALROBOT-065 provides the test method to verify SUB-MAIN-050 meets its measurable acceptance criteria. | 2026-03-20T11:50:41.931Z
SUB-MAIN-047 | VER-MAIN-064 | verifies | Verification procedure for SUB-MAIN-047 | REQ-SESURGICALROBOT-064 provides the test method to verify SUB-MAIN-047 meets its measurable acceptance criteria. | 2026-03-20T11:50:41.272Z
SUB-MAIN-039 | VER-MAIN-063 | verifies | Verification procedure for SUB-MAIN-039 | REQ-SESURGICALROBOT-063 provides the test method to verify SUB-MAIN-039 meets its measurable acceptance criteria. | 2026-03-20T11:50:40.637Z
SUB-MAIN-029 | VER-MAIN-062 | verifies | Verification procedure for SUB-MAIN-029 | REQ-SESURGICALROBOT-062 provides the test method to verify SUB-MAIN-029 meets its measurable acceptance criteria. | 2026-03-20T11:50:39.991Z
SUB-MAIN-027 | VER-MAIN-061 | verifies | Verification procedure for SUB-MAIN-027 | REQ-SESURGICALROBOT-061 provides the test method to verify SUB-MAIN-027 meets its measurable acceptance criteria. | 2026-03-20T11:50:39.373Z
SUB-MAIN-025 | VER-MAIN-060 | verifies | Verification procedure for SUB-MAIN-025 | REQ-SESURGICALROBOT-060 provides the test method to verify SUB-MAIN-025 meets its measurable acceptance criteria. | 2026-03-20T11:50:38.733Z
SUB-MAIN-024 | VER-MAIN-059 | verifies | Verification procedure for SUB-MAIN-024 | REQ-SESURGICALROBOT-059 provides the test method to verify SUB-MAIN-024 meets its measurable acceptance criteria. | 2026-03-20T11:50:38.050Z
SUB-MAIN-022 | VER-MAIN-058 | verifies | Verification procedure for SUB-MAIN-022 | REQ-SESURGICALROBOT-058 provides the test method to verify SUB-MAIN-022 meets its measurable acceptance criteria. | 2026-03-20T11:50:37.396Z
SUB-MAIN-021 | VER-MAIN-057 | verifies | Verification procedure for SUB-MAIN-021 | REQ-SESURGICALROBOT-057 provides the test method to verify SUB-MAIN-021 meets its measurable acceptance criteria. | 2026-03-20T11:50:36.738Z
SUB-MAIN-020 | VER-MAIN-056 | verifies | Verification procedure for SUB-MAIN-020 | REQ-SESURGICALROBOT-056 provides the test method to verify SUB-MAIN-020 meets its measurable acceptance criteria. | 2026-03-20T11:50:36.074Z
SUB-MAIN-018 | VER-MAIN-055 | verifies | Verification procedure for SUB-MAIN-018 | REQ-SESURGICALROBOT-055 provides the test method to verify SUB-MAIN-018 meets its measurable acceptance criteria. | 2026-03-20T11:50:27.327Z
SUB-MAIN-016 | VER-MAIN-054 | verifies | Verification procedure for SUB-MAIN-016 | REQ-SESURGICALROBOT-054 provides the test method to verify SUB-MAIN-016 meets its measurable acceptance criteria. | 2026-03-20T11:50:26.717Z
SUB-MAIN-015 | VER-MAIN-053 | verifies | Verification procedure for SUB-MAIN-015 | REQ-SESURGICALROBOT-053 provides the test method to verify SUB-MAIN-015 meets its measurable acceptance criteria. | 2026-03-20T11:50:26.073Z
SUB-MAIN-012 | VER-MAIN-052 | verifies | Verification procedure for SUB-MAIN-012 | REQ-SESURGICALROBOT-052 provides the test method to verify SUB-MAIN-012 meets its measurable acceptance criteria. | 2026-03-20T11:50:25.450Z
SUB-MAIN-011 | VER-MAIN-051 | verifies | Verification procedure for SUB-MAIN-011 | REQ-SESURGICALROBOT-051 provides the test method to verify SUB-MAIN-011 meets its measurable acceptance criteria. | 2026-03-20T11:50:24.821Z
SUB-MAIN-010 | VER-MAIN-050 | verifies | Verification procedure for SUB-MAIN-010 | REQ-SESURGICALROBOT-050 provides the test method to verify SUB-MAIN-010 meets its measurable acceptance criteria. | 2026-03-20T11:50:24.151Z
SYS-MAIN-015 | ARC-MAIN-020 | derives | PDR DMA shared memory architecture derives from 1kHz data recording requirement | SYS-MAIN-015 requires 1kHz recording of all kinematic data; the architecture decision to use DMA shared memory for PDR without touching the RTPE thread is necessary to achieve this rate without jitter in the control loop. | 2026-03-20T11:35:41.977Z
SYS-MAIN-002 | ARC-MAIN-020 | derives | Dual-redundant fibre topology derives from single-point failure detection requirement | SYS-MAIN-002 requires detection of any communication loss; the dual-redundant fibre with FPGA-level failover is the architectural response ensuring no single communication failure can halt the 1kHz kinematic command loop. | 2026-03-20T11:35:41.322Z
SUB-MAIN-107 | VER-MAIN-116 | verifies | Verification of NMC link health classification accuracy | VER-MAIN-116 uses BERT injection to verify the three-state classification transitions at the correct BER thresholds within two monitoring cycles. | 2026-03-20T11:15:47.617Z
SUB-MAIN-104 | VER-MAIN-115 | verifies | Verification of Inter-Cart Fibre Link failover time requirement | VER-MAIN-115 verifies the 5ms failover switchover time and command continuity through relay-controlled fault injection across 20 trials. | 2026-03-20T11:15:46.993Z
SUB-MAIN-103 | VER-MAIN-114 | verifies | Verification of Inter-Cart Fibre Link latency requirement | VER-MAIN-114 directly tests the 500us P99 latency criterion in SUB-MAIN-103 using hardware timestamps under peak load. | 2026-03-20T11:15:46.391Z
SYS-MAIN-015 | SUB-MAIN-110 | derives | PDR cryptographic hash derives from procedure data recording and integrity requirement | SYS-MAIN-015 requires all procedural data to be recorded. SUB-MAIN-110 adds the integrity verification dimension (SHA-256) required for regulatory compliance of that recorded data. | 2026-03-20T11:15:36.179Z
SYS-MAIN-001 | SUB-MAIN-108 | derives | Strict-priority traffic shaping derives from end-to-end control loop latency requirement | SYS-MAIN-001 requires <1ms end-to-end latency. SUB-MAIN-108 implements the network QoS policy ensuring safety and kinematic frames are never delayed by lower-priority video traffic, protecting the latency budget. | 2026-03-20T11:15:35.560Z
SYS-MAIN-002 | SUB-MAIN-107 | derives | NMC link health classification derives from failure detection requirement | SYS-MAIN-002 requires detection of communication faults. SUB-MAIN-107 specifies the NMC three-state health classification that enables pre-failure detection of DEGRADED links before complete failure. | 2026-03-20T11:15:34.912Z
SYS-MAIN-002 | SUB-MAIN-106 | derives | RTPE frame loss detection derives from single-point failure detection requirement | SYS-MAIN-002 requires detection of communication loss. SUB-MAIN-106 specifies the RTPE implementation: sequence discontinuity detection within 1ms and COMM_FAULT reporting within 2ms. | 2026-03-20T11:15:34.266Z
IFC-MAIN-048 | VER-MAIN-118 | verifies | Integration test for RTPE to PDR DMA delivery latency | VER-MAIN-118 verifies 100us DMA delivery latency and confirms no RTPE jitter increase, which are the acceptance criteria for IFC-MAIN-048. | 2026-03-20T11:12:55.391Z
IFC-MAIN-047 | VER-MAIN-117 | verifies | Integration test for NMC to SIS COMM_FAULT notification interface | VER-MAIN-117 verifies timing of the LVDS COMM_FAULT signal, which is the concrete pass/fail criterion for IFC-MAIN-047 2ms notification requirement. | 2026-03-20T11:12:54.681Z
SYS-MAIN-015 | SUB-MAIN-109 | derives | PDR throughput requirement derives from 8-hour 1kHz recording mandate | SYS-MAIN-015 mandates 1kHz kinematic plus video recording for 8 hours. SUB-MAIN-109 specifies the 2GB/s NVMe throughput needed to sustain that data rate without buffer overflow, derived from peak data rate calculations. | 2026-03-20T11:12:33.992Z
SYS-MAIN-018 | SUB-MAIN-105 | derives | Frame-level HMAC authentication derives from cryptographic authentication requirement | SYS-MAIN-018 requires cryptographic authentication on safety-critical inter-subsystem interfaces. SUB-MAIN-105 specifies HMAC-SHA256 per kinematic command frame as the concrete implementation for the highest-throughput safety-critical interface in the system. | 2026-03-20T11:12:33.336Z
SYS-MAIN-002 | SUB-MAIN-104 | derives | Fibre link redundancy and failover derives from single-point failure detection requirement | SYS-MAIN-002 requires single-point failure detection and recovery. SUB-MAIN-104 specifies the concrete failover behaviour: 5ms switchover with no command continuity loss, which is the implementation of that requirement for the communication subsystem. | 2026-03-20T11:12:32.733Z
SYS-MAIN-001 | SUB-MAIN-103 | derives | Fibre link latency budget derives from 1ms end-to-end control loop requirement | SYS-MAIN-001 mandates <1ms end-to-end loop latency. SUB-MAIN-103 allocates 500us to the fibre path as the largest single latency contributor in the inter-cart chain. | 2026-03-20T11:12:32.071Z
SYS-MAIN-005 | SUB-MAIN-102 | derives | PMS IEC 60601-1 compliance derives from system power continuity safety requirement | SYS-MAIN-005 requires power backup for controlled arm hold during power loss; achieving that with medical-grade safety mandates IEC 60601-1 F-Type isolation on all patient-coupled supply rails. | 2026-03-20T10:55:42.458Z
SYS-MAIN-007 | SUB-MAIN-047 | derives | Console Computer regulatory qualification derived from motion command system requirement | SYS-MAIN-007 mandates real-time motion command transmission from surgeon console to patient-side cart. The Console Computer is the primary hardware executing this function; EU MDR 2017/745 and FDA 510(k) qualification are required for it to be deployed as a motion-initiating medical device accessory. | 2026-03-20T10:55:10.228Z
SUB-MAIN-042 | VER-MAIN-049 | verifies | UPS mains dropout test verifies power management redundancy requirement | REQ-049 is the mains-dropout test procedure that verifies the UPS transfer time and sustained operation duration specified in REQ-042. | 2026-03-20T10:39:44.476Z
SUB-MAIN-040 | VER-MAIN-048 | verifies | Haptic redundancy fault injection test verifies REQ-040 | REQ-048 is the direct test procedure that verifies the haptic redundancy switchover latency and accuracy floor specified in REQ-040. | 2026-03-20T10:39:36.947Z
SYS-MAIN-012 | SUB-MAIN-046 | derives | Workspace safety enforcer compliance derives from system force limit requirement | SYS-MAIN-012 limits force applied to patient tissue; REQ-046 derives the workspace enforcer's compliance certification and reaction time requirement, which is the mechanism for enforcing the spatial component of that tissue-force limit. | 2026-03-20T10:38:55.100Z
SYS-MAIN-007 | SUB-MAIN-045 | derives | IEC 62304 Class C software requirement derives from system motion transmission accuracy requirement | SYS-MAIN-007 requires accurate real-time motion command transmission from master to slave; REQ-045 derives the software development class requirement ensuring the motion control software's reliability and coverage is commensurate with the patient safety implications of this command path. | 2026-03-20T10:38:44.137Z
SYS-MAIN-002 | SUB-MAIN-044 | derives | Interlock SIL 3 compliance derives from system safety requirement for single-point failure response | SYS-MAIN-002 requires safe single-point failure response; REQ-044 specifies the regulatory certification standard (IEC 61508 SIL 3) the interlock subsystem must achieve to credibly satisfy that safety architecture obligation. | 2026-03-20T10:38:31.677Z
SYS-MAIN-015 | SUB-MAIN-043 | derives | Data recorder redundancy derives from system-level comprehensive recording requirement | SYS-MAIN-015 mandates comprehensive recording of kinematic, video, and system state data throughout all procedures; REQ-043 derives the redundancy mechanism ensuring this obligation is not voided by primary recorder failure. | 2026-03-20T10:38:19.646Z
SYS-MAIN-002 | SUB-MAIN-042 | derives | Power UPS transfer requirement derives from system single-point failure requirement | SYS-MAIN-002 requires single-point failure response; REQ-042 is the power management subsystem's specific redundancy mechanism: UPS transfer within 10 ms to prevent servo watchdog halt on power loss. | 2026-03-20T10:38:07.593Z
SYS-MAIN-002 | SUB-MAIN-041 | derives | Time protocol engine redundancy derives from system-level single-point failure requirement | SYS-MAIN-002 requires safe single-point failure response across the system; REQ-041 derives the time protocol engine's specific mechanism — hot-standby grandmaster switchover within 200 ms — to ensure clock loss does not propagate as an unhandled failure. | 2026-03-20T10:37:56.656Z
SYS-MAIN-002 | SUB-MAIN-040 | derives | Haptic redundancy requirement derives from system-level single-point failure requirement | SYS-MAIN-002 mandates detection and safe response to any single-point failure; REQ-040 derives from this by specifying the haptic subsystem's specific failover mechanism and performance floor when the primary force-rendering processor fails. | 2026-03-20T10:37:35.026Z
IFC-MAIN-002 | VER-MAIN-107 | verifies | IFC-MAIN-002 E-stop interface requirement verified by contactor drop test | VER-MAIN-107 is the direct verification procedure for IFC-MAIN-002 (E-stop chain interface requirement); the test confirms the 10ms contactor drop performance specified in that interface requirement. | 2026-03-20T10:36:53.750Z
IFC-MAIN-040 | VER-MAIN-113 | verifies | Console-to-patient-cart command stream latency and continuity test | Validates the primary teleoperation data path meets 5ms latency and 1kHz frame rate required for surgeon feel; failure mode is motion discontinuity or lag felt by the surgeon | 2026-03-20T10:23:47.350Z
IFC-MAIN-023 | VER-MAIN-112 | verifies | Instrument lockout command delivery and SAFE-HOLD transition test | Validates three lockout trigger scenarios reach the Safe State Manager and enforce SAFE-HOLD; prevents use of sterility-compromised or overused instruments during surgery | 2026-03-20T10:23:26.160Z
IFC-MAIN-022 | VER-MAIN-111 | verifies | TTAC to IDU cable displacement accuracy test | Validates kinematic model accuracy at the tool drive interface; cable displacement errors accumulate into tip position error that affects surgical precision | 2026-03-20T10:23:09.510Z
IFC-MAIN-021 | VER-MAIN-110 | verifies | Cable tension data path and lockout integration test | Validates complete signal path from tension sensor through conditioning to safety subsystem flag logic; overtension lockout must be proven before clinical use to prevent cable fracture injury | 2026-03-20T10:22:51.865Z
IFC-MAIN-007 | VER-MAIN-109 | verifies | Heartbeat interruption SAFE-HOLD test | Tests that loss of MCS heartbeat triggers SAFE-HOLD within 20ms; demonstrates the watchdog monitoring path is operational and the safety system correctly isolates on loss of liveness signal | 2026-03-20T10:22:27.753Z
IFC-MAIN-002 | VER-MAIN-108 | verifies | Hardwired E-stop to contactor drop test | Tests the hardware interlock path from E-stop chain open to servo contactor de-energisation; 10ms limit represents the worst-case allowable delay before an uncontrolled joint can move beyond safe bounds | 2026-03-20T10:20:55.412Z
SUB-MAIN-101 | VER-MAIN-106 | verifies | Test verifies SUB-MAIN-101 authentication requirement | Direct verification of the authentication requirement by exercising both valid and invalid credential paths at the specified rejection latency boundary. | 2026-03-20T10:10:13.787Z
SUB-MAIN-100 | VER-MAIN-105 | verifies | Test verifies SUB-MAIN-100 authentication requirement | Direct verification of the authentication requirement by exercising both valid and invalid credential paths at the specified rejection latency boundary. | 2026-03-20T10:10:12.382Z
SUB-MAIN-099 | VER-MAIN-104 | verifies | Test verifies SUB-MAIN-099 authentication requirement | Direct verification of the authentication requirement by exercising both valid and invalid credential paths at the specified rejection latency boundary. | 2026-03-20T10:10:11.112Z
SUB-MAIN-098 | VER-MAIN-103 | verifies | Test verifies SUB-MAIN-098 authentication requirement | Direct verification of the authentication requirement by exercising both valid and invalid credential paths at the specified rejection latency boundary. | 2026-03-20T10:10:09.840Z
SYS-MAIN-018 | SUB-MAIN-101 | derives | SYS-018 cybersecurity auth cascades to SUB-MAIN-101 | RTPE auth - RTPE handles all inter-cart synchronisation frames at 1kHz, making it a high-impact target for timing attacks on the master-slave control loop | 2026-03-20T10:09:48.083Z
SYS-MAIN-018 | SUB-MAIN-100 | derives | SYS-018 cybersecurity auth cascades to SUB-MAIN-100 | TG waypoint safety - TG is the trajectory planning stage where unsafe workspace boundaries can be enforced before commands reach servo layer | 2026-03-20T10:09:47.467Z
SYS-MAIN-018 | SUB-MAIN-099 | derives | SYS-018 cybersecurity auth cascades to SUB-MAIN-099 | KE auth - KE parses all Cartesian commands before IK, making it the upstream command validation point for arm motion | 2026-03-20T10:09:46.822Z
SYS-MAIN-018 | SUB-MAIN-098 | derives | Cybersecurity auth cascades to TTAC | SYS-MAIN-018 mandates authentication for all safety-critical interfaces; SUB-MAIN-098 implements this for TTAC which directly drives distal instrument DOFs, making it a high-risk command injection target. | 2026-03-20T10:09:33.900Z
SYS-MAIN-018 | SUB-MAIN-101 | derives | Cybersecurity auth requirement cascades to SUB-MAIN-101 | SYS-MAIN-018 mandates authentication across all safety-critical interfaces; SUB-MAIN-101 implements this mandate for its specific subsystem component, specifying the HMAC/CRC scheme and rejection timing appropriate to that component's control cycle. | 2026-03-20T10:09:26.126Z
SYS-MAIN-018 | SUB-MAIN-100 | derives | Cybersecurity auth requirement cascades to SUB-MAIN-100 | SYS-MAIN-018 mandates authentication across all safety-critical interfaces; SUB-MAIN-100 implements this mandate for its specific subsystem component, specifying the HMAC/CRC scheme and rejection timing appropriate to that component's control cycle. | 2026-03-20T10:09:25.472Z
SYS-MAIN-018 | SUB-MAIN-099 | derives | Cybersecurity auth requirement cascades to SUB-MAIN-099 | SYS-MAIN-018 mandates authentication across all safety-critical interfaces; SUB-MAIN-099 implements this mandate for its specific subsystem component, specifying the HMAC/CRC scheme and rejection timing appropriate to that component's control cycle. | 2026-03-20T10:09:24.859Z
SYS-MAIN-018 | SUB-MAIN-098 | derives | Cybersecurity auth requirement cascades to SUB-MAIN-098 | SYS-MAIN-018 mandates authentication across all safety-critical interfaces; SUB-MAIN-098 implements this mandate for its specific subsystem component, specifying the HMAC/CRC scheme and rejection timing appropriate to that component's control cycle. | 2026-03-20T10:09:24.217Z
SUB-MAIN-083 | VER-MAIN-039 | verifies | Degraded-mode fault injection test verifies WSE bounding-box fallback | SUB-MAIN-083 specifies that the WSE must fall back to conservative bounding-box limits when the anatomy mesh is unavailable. The test injects the specific fault condition and verifies the fallback behaviour, which cannot be exercised under normal integration testing. | 2026-03-20T09:28:09.915Z
SUB-MAIN-082 | VER-MAIN-038 | verifies | Proximity enforcement test verifies WSE anatomy boundary protection | SUB-MAIN-082 requires the WSE to enforce anatomy boundaries without penetration. The test directly exercises the constraint activation path under representative surgical motion trajectories and verifies the non-penetration guarantee. | 2026-03-20T09:28:09.313Z
SUB-MAIN-071 | VER-MAIN-037 | verifies | TDM cycle-to-cycle jitter test verifies RTPE scheduler timing budget | SUB-MAIN-071 specifies the +/-50us jitter budget for the TDM frame scheduler. The measurement procedure directly characterises the jitter distribution and determines compliance with the specified limit. The test exists because this requirement exists. | 2026-03-20T09:28:08.698Z
IFC-MAIN-040 | VER-MAIN-036 | verifies | Protocol integrity test verifies CC-to-ICFL command frame format compliance | IFC-MAIN-040 specifies the frame format, rate, and error detection for the kinematic command interface. The protocol integrity test directly verifies each specified parameter (frame size, rate, sequence number, CRC) under sustained load. The test is derived from this requirement and would not exist without it. | 2026-03-20T09:27:26.382Z
SUB-MAIN-094 | VER-MAIN-035 | verifies | IEC 62304 Class C qualification inspection procedure verifies SUB-MAIN-094 software lifecycle compliance | SUB-MAIN-094 requires the Motion Control System software to be qualified to IEC 62304 Class C. The verification procedure (audit of lifecycle artefacts) directly addresses each clause referenced in the requirement. This is a one-to-one requirement-to-verification relationship. | 2026-03-20T09:27:07.661Z
SUB-MAIN-097 | VER-MAIN-102 | verifies | HMAC penetration test verifies cybersecurity authentication for MCS | VER-MAIN-102 tests the authentication boundary with valid, corrupt, and replay frames — the three attack scenarios that HMAC-SHA256 must defeat per SUB-MAIN-097. | 2026-03-20T09:12:33.854Z
SUB-MAIN-096 | VER-MAIN-101 | verifies | Haptic FSM link-loss test verifies force-blind degraded mode | VER-MAIN-101 directly tests the 50ms entry timing, 0.3N braking force, and surgeon kinematic control retention specified in SUB-MAIN-096. | 2026-03-20T09:12:32.584Z
SUB-MAIN-095 | VER-MAIN-100 | verifies | Watchdog failover fault injection test verifies Console Computer redundancy | VER-MAIN-100 directly tests the 500ms failover timing, position preservation, and notification requirements in SUB-MAIN-095 through repeatable fault injection. | 2026-03-20T09:12:11.788Z
SUB-MAIN-093 | VER-MAIN-099 | verifies | Independent SIL 3 safety assessment verifies compliance | IEC 62061 requires independent competent body review for SIL 3 certification; this is the direct verification method for the PFH figure in SUB-MAIN-093. | 2026-03-20T09:12:10.471Z
SUB-MAIN-037 | VER-MAIN-098 | verifies | Lifecycle enforcement boundary test verifies use-count rejection | VER-MAIN-098 tests all boundary conditions of the instrument use-count enforcement logic in SUB-MAIN-037, including the exact rejection threshold and the audit logging requirement. | 2026-03-20T09:11:49.128Z
SUB-MAIN-036 | VER-MAIN-097 | verifies | Articulation controller cable displacement test verifies tip accuracy | VER-MAIN-097 validates that cable displacement commands computed by the Tool Tip Articulation Controller maintain clinical accuracy under worn-cable conditions — the specific scenario SUB-MAIN-036 is designed to handle. | 2026-03-20T09:11:47.723Z
SUB-MAIN-033 | VER-MAIN-096 | verifies | Instrument drive actuation test verifies DOF performance | VER-MAIN-096 exercises all four instrument DOF under commanded motion and measures positional error and latency at the tool tip — the direct observable outputs of SUB-MAIN-033. | 2026-03-20T09:11:32.508Z
SYS-MAIN-007 | SUB-MAIN-097 | derives | Motion command integrity requirement drives cybersecurity authentication | SYS-MAIN-007 requires that surgeon hand motion commands are faithfully transmitted; authentication of command frames ensures that transmitted commands are genuinely from the surgeon console and have not been tampered with in transit. | 2026-03-20T09:11:18.940Z
SYS-MAIN-004 | SUB-MAIN-096 | derives | Tissue force measurement mandate drives haptic degraded-mode specification | SYS-MAIN-004 requires force measurement and display; SUB-MAIN-096 defines what the subsystem must do when that measurement chain fails — the degraded behaviour is a direct consequence of the force measurement mandate. | 2026-03-20T09:11:02.979Z
SYS-MAIN-016 | SUB-MAIN-095 | derives | System degraded-mode requirement drives console failover design | SYS-MAIN-016 establishes the principle that the system maintains some capability on component loss; SUB-MAIN-095 applies this principle to the Console Computer specifically, ensuring surgical session continuity on a software fault. | 2026-03-20T09:10:47.752Z
SYS-MAIN-002 | SUB-MAIN-094 | derives | Single-point failure detection drives IEC 62304 Class C classification for motion control software | SYS-MAIN-002 requires detection of any single-point failure; motion control software executing this function on a Class IIb device must be Class C under IEC 62304 to provide adequate lifecycle assurance. | 2026-03-20T09:10:31.214Z
SYS-MAIN-002 | SUB-MAIN-093 | derives | System single-point failure detection derives SIS SIL 3 compliance | SYS-MAIN-002 mandates detection and safe response to any single-point failure; meeting that requirement for a Class IIb device requires the Safety and Interlock Subsystem to be qualified at SIL 3 per IEC 62061. | 2026-03-20T09:10:17.552Z
SYS-MAIN-002 | ARC-MAIN-019 | derives | Verification coverage architecture derives from system safety detection requirement | ARC-MAIN-019 documents the prioritisation rationale for verification entries — the SIL-based ordering flows directly from SYS-MAIN-002's mandate that every single-point failure be detected and handled safely. | 2026-03-20T09:09:42.110Z
SUB-MAIN-092 | VER-MAIN-095 | verifies | Load regulation and battery endurance test verifies AUX PSU spec | VER-MAIN-095 tests both the plus or minus 2% voltage regulation and 20-minute independent operation specified in SUB-MAIN-092. | 2026-03-20T08:42:30.213Z
SUB-MAIN-090 | VER-MAIN-094 | verifies | SoC accuracy and low-battery alert test verifies UPS telemetry requirement | VER-MAIN-094 tests the plus or minus 2% accuracy, 1 Hz reporting rate, and 25% alert threshold in SUB-MAIN-090. | 2026-03-20T08:42:29.564Z
IFC-MAIN-001 | uht-bot:se-surgical-robot:VER-MAIN-073 | verifies | Joint force monitoring test verifies SIS-MCS torque interface | VER-MAIN-073 tests the calibrated torque load injection and monitoring response specified in IFC-MAIN-001, confirming the SIS can detect and respond to force limit exceedances at the joint level. | 2026-03-20T08:42:07.860Z
SYS-MAIN-005 | SUB-MAIN-092 | derives | Auxiliary Power Supply spec derives from safety circuit continuity requirement | SYS-MAIN-005 requires controlled arm position after complete power loss; SUB-MAIN-092 defines the dedicated 24V rail that maintains the Safe State Manager and Emergency Stop Chain beyond UPS depletion, which is the last line of power-backed safety function. | 2026-03-20T08:42:07.266Z
SYS-MAIN-005 | SUB-MAIN-090 | derives | UPS SOC reporting derives from system power continuity requirement | SYS-MAIN-005 requires 10-minute controlled operation after mains loss; SUB-MAIN-090 is the telemetry mechanism that allows the Power Sequencing Controller to predict and manage battery depletion — without accurate SOC, the system cannot provide the 8-minute warning margin needed for orderly shutdown. | 2026-03-20T08:42:06.665Z
SUB-MAIN-091 | VER-MAIN-093 | verifies | Mains collapse test verifies PSC transfer timing | VER-MAIN-093 measures end-to-end transfer time at PDU rail, directly verifying the 30 ms constraint in SUB-MAIN-091. | 2026-03-20T08:40:35.630Z
SUB-MAIN-089 | VER-MAIN-092 | verifies | Ground fault injection test verifies PDU leakage detection | VER-MAIN-092 tests the 600 µA threshold, 100 ms response, and branch isolation isolation behaviour specified in SUB-MAIN-089. | 2026-03-20T08:40:35.047Z
IFC-MAIN-046 | VER-MAIN-091 | verifies | PSC-AUX mode transition test verifies discrete control interface | VER-MAIN-091 tests the 50 ms mode acknowledgement and CAN-bus fallback behaviour specified in IFC-MAIN-046. | 2026-03-20T08:40:24.026Z
IFC-MAIN-045 | VER-MAIN-090 | verifies | DC link load test verifies UPS-PDU impedance and current rating | VER-MAIN-090 tests the 50 mOhm impedance and 200 A pulse rating directly specified in IFC-MAIN-045. | 2026-03-20T08:40:23.431Z
IFC-MAIN-044 | VER-MAIN-089 | verifies | CAN bus integration test verifies PDU-PSC interface | VER-MAIN-089 directly tests the CAN FD frame timing and galvanic isolation specified in IFC-MAIN-044; pass/fail criteria map one-to-one to the 5 ms latency and 500 VDC isolation requirements. | 2026-03-20T08:40:22.820Z
SYS-MAIN-005 | SUB-MAIN-091 | derives | PSC mains-loss transfer derives directly from system UPS continuity requirement | SYS-MAIN-005 mandates controlled arm function for 10 minutes after mains loss; SUB-MAIN-091 is the specific mechanism that achieves seamless transition to UPS — if transfer exceeds 30 ms the servo controllers fault and system continuity fails. | 2026-03-20T08:40:22.214Z
SYS-MAIN-005 | SUB-MAIN-088 | derives | PDU branch circuit requirements derive from system power integrity requirement | SYS-MAIN-005 requires the system maintain controlled operation after mains loss; SUB-MAIN-088 defines the PDU branch circuit protection that prevents single-branch faults from disrupting the overall power distribution, which is a direct architectural contributor to system-level power integrity. | 2026-03-20T08:40:21.384Z
SYS-MAIN-001 | VER-MAIN-084 | verifies | End-to-end teleoperation chain test against SYS-MAIN-001 latency | SYS-MAIN-001 requires <1ms end-effector response to surgeon input; VER-MAIN-084 measures the full chain latency under concurrent load, which is the only way to validate this system-level requirement. | 2026-03-20T08:13:00.835Z
IFC-MAIN-002 | VER-MAIN-072 | verifies | E-stop de-energisation test (prior session) | Supplements VER-MAIN-075; both verify IFC-MAIN-002 confirming the hardwired E-stop interlock meets IEC 60601-1 de-energisation requirements. | 2026-03-20T08:12:52.080Z
IFC-MAIN-001 | VER-MAIN-071 | verifies | Additional joint force monitor latency test (prior session) | Supplements VER-MAIN-074 with a variant torque injection protocol; both verify IFC-MAIN-001 from different test angles providing redundant evidence for SIL 2. | 2026-03-20T08:12:51.472Z
IFC-MAIN-020 | VER-MAIN-088 | verifies | Instrument recognition completeness and rejection test for IFC-MAIN-020 | Randomised instrument sequence including unknown and expired instruments validates both positive recognition and safety rejection paths; zero false acceptance is a patient safety pass criterion. | 2026-03-20T08:10:44.906Z
IFC-MAIN-013 | VER-MAIN-087 | verifies | ESG activation and de-activation timing test for IFC-MAIN-013 | Resistive load test validates power-on and power-off timing; return electrode check sequence verifies patient safety interlock is active before each activation. | 2026-03-20T08:10:44.283Z
IFC-MAIN-012 | VER-MAIN-086 | verifies | UPS switchover and sustain test for IFC-MAIN-012 | Physical mains interruption is necessary to validate true switchover timing including inrush transient that power supply emulators cannot reproduce. | 2026-03-20T08:10:16.323Z
IFC-MAIN-011 | VER-MAIN-085 | verifies | Haptic force rendering latency and accuracy test for IFC-MAIN-011 | Calibrated force applicator at tip validates the round-trip haptic rendering latency and accuracy under known reference loads. | 2026-03-20T08:10:15.716Z
IFC-MAIN-010 | VER-MAIN-083 | verifies | Stereo endoscope dual-channel sync and SNR test for IFC-MAIN-010 | Inter-channel sync measurement validates stereopsis quality; SNR floor confirms illumination-independent image fidelity. | 2026-03-20T08:09:46.576Z
IFC-MAIN-009 | VER-MAIN-082 | verifies | Instrument torque feedback accuracy and rate test for IFC-MAIN-009 | Calibrated fixture comparison is the only method to validate torque accuracy independently of the instrument drive; 500Hz continuity must be proven across dynamic manoeuvres. | 2026-03-20T08:09:31.433Z
IFC-MAIN-008 | VER-MAIN-081 | verifies | Stereo video latency and colour fidelity test for IFC-MAIN-008 | End-to-end photon-to-pixel latency and ΔE colour measurement validate both the temporal and chromatic requirements of the surgical imaging interface. | 2026-03-20T08:09:16.185Z
IFC-MAIN-007 | VER-MAIN-080 | verifies | Safety heartbeat interruption test for IFC-MAIN-007 | Fault injection at three interruption intervals covers worst-case liveness miss detection and validates auto-recovery without manual intervention. | 2026-03-20T08:09:00.367Z
IFC-MAIN-006 | VER-MAIN-079 | verifies | 250Hz joint command stream jitter test for IFC-MAIN-006 | Jitter measurement over 30 minutes at operating temperature validates that timing drift does not degrade servo control quality. | 2026-03-20T08:08:47.430Z
IFC-MAIN-005 | VER-MAIN-078 | verifies | Surgeon console to motion scaling command latency test for IFC-MAIN-005 | Latency measurement at operational and elevated rates validates both nominal transparency and overload behaviour of the command interface. | 2026-03-20T08:08:24.959Z
IFC-MAIN-004 | VER-MAIN-077 | verifies | Safe State Manager broadcast latency test for IFC-MAIN-004 | Tests all three transition paths to confirm broadcast latency stays within budget that drives subsystem safe-state adoption. | 2026-03-20T08:07:56.918Z
IFC-MAIN-003 | VER-MAIN-076 | verifies | Fibre link CRC error detection test for IFC-MAIN-003 | BER injection is the only practical method to validate detection latency and false alarm rate at required statistical confidence; real faults cannot be induced repeatably. | 2026-03-20T08:07:46.095Z
IFC-MAIN-002 | VER-MAIN-075 | verifies | E-stop hardware de-energisation test for IFC-MAIN-002 | Physical test is the only way to validate hardwired de-energisation latency; the 150ms budget must be demonstrated under representative servo load to confirm the interlock chain meets IEC 60601-1 requirements. | 2026-03-20T08:07:32.652Z
IFC-MAIN-001 | VER-MAIN-074 | verifies | Joint Force Monitor latency test for IFC-MAIN-001 | Demonstrates the 2ms detection latency specified in IFC-MAIN-001 is achieved in hardware under representative load conditions; pass/fail criteria are directly derived from the requirement threshold. | 2026-03-20T08:07:18.608Z
SYS-MAIN-001 | ARC-MAIN-018 | derives | Linear pipeline topology derives from 1ms end-to-end latency requirement | ARC-MAIN-018 documents the decision to use a single-node linear pipeline topology, directly motivated by the 1ms end-to-end latency budget in SYS-MAIN-001. The rejected distributed alternative adds 0.2-0.4ms per pipeline stage — making the 1ms budget unachievable — and ARC-MAIN-018 records this trade-off explicitly. | 2026-03-20T07:09:17.454Z
SYS-MAIN-002 | ARC-MAIN-015 | derives | FPGA/algorithm separation architecture supports single-point-failure safe-state | ARC-MAIN-015 documents the architectural decision to separate FPGA infrastructure from software algorithms, where the FPGA watchdog ensures safe-state within 250ms even if software hangs — directly implementing SYS-MAIN-002's single-point failure safe-state requirement. | 2026-03-20T07:08:47.508Z
SUB-MAIN-031 | VER-MAIN-034 | derives | RTP Engine PTP authentication test procedure | REQ-SESURGICALROBOT-034 verifies the three SHALL clauses of the RTP Engine timing authentication requirement: MAC rejection, safe-hold timing, and surgeon alert. | 2026-03-20T07:08:46.855Z
SUB-MAIN-030 | VER-MAIN-033 | derives | Trajectory Generator envelope verification procedure | REQ-SESURGICALROBOT-033 verifies both the boot-time RSA signature check and the per-waypoint envelope validation for the Trajectory Generator cybersecurity requirement. | 2026-03-20T07:08:46.171Z
SUB-MAIN-029 | VER-MAIN-032 | derives | Kinematic engine authentication verification | REQ-SESURGICALROBOT-032 is the direct verification procedure for the Kinematics Engine authentication requirement, confirming both rejection and logging behaviour at 1kHz operational rate. | 2026-03-20T07:08:45.557Z
SYS-MAIN-002 | SUB-MAIN-031 | derives | RTP Engine timing authentication derives from communication-loss safe-state requirement | SYS-MAIN-002 includes communication loss as a failure requiring safe state within 250ms. A time-injection attack on the Real-Time Protocol Engine induces a synthetic communication loss — all servo controllers receive a corrupted time base and desynchronise. Authenticating PTP sync messages prevents this failure mode, and the 10ms safe-hold threshold is derived directly from the 250ms safe-state budget. | 2026-03-20T07:07:40.252Z
SYS-MAIN-002 | SUB-MAIN-030 | derives | Trajectory Generator envelope verification derives from single-point-failure safe-state requirement | SYS-MAIN-002 requires the system to achieve safe state within 250ms of any single-point failure. A maliciously injected trajectory waypoint is a software-exception failure mode that must be detected and rejected before it propagates to joint servo execution. Workspace-envelope signature verification closes this attack vector at the Trajectory Generator, directly supporting the 250ms safe-state guarantee. | 2026-03-20T07:07:39.626Z
SYS-MAIN-002 | SUB-MAIN-029 | derives | Kinematics Engine command authentication derives from software-exception safe-state requirement | SYS-MAIN-002 identifies software exceptions as a class of single-point failure requiring safe state within 250ms. Authentication of kinematics command inputs prevents a specific class of software exception — malicious or corrupted motion commands — from ever entering the solver pipeline, addressing the root cause rather than just detecting the failure. | 2026-03-20T07:07:38.985Z
SYS-MAIN-009 | VER-MAIN-070 | verifies | End-to-end pipeline test also validates tremor attenuation at system level | SYS-MAIN-009 requires 40dB tremor attenuation at the instrument tip; VER-MAIN-070 measures this end-to-end through the real Tremor Rejection Filter under full pipeline load. | 2026-03-20T06:13:21.715Z
SYS-MAIN-001 | VER-MAIN-070 | verifies | End-to-end MC pipeline test verifies system-level latency and precision requirements | SYS-MAIN-001 requires 1mm RMS tip accuracy and 1ms latency; VER-MAIN-070 is the integration test that validates these at the full pipeline level with all three arms under load. | 2026-03-20T06:13:21.083Z
SYS-MAIN-009 | SUB-MAIN-081 | derives | Tremor monitoring and logging derives from system tremor attenuation requirement | SYS-MAIN-009 requires 40dB tremor attenuation; SUB-MAIN-081 implements the monitoring and logging that provides observability into tremor levels, supporting post-operative analysis and future adaptive filter calibration against the system figure. | 2026-03-20T06:13:20.453Z
SYS-MAIN-007 | SUB-MAIN-087 | derives | MC&S command authentication requirement derived from motion command integrity need | SYS-MAIN-007 requires motion commands to be delivered reliably within 1ms; SUB-MAIN-087 adds the integrity guarantee that only authenticated commands enter the pipeline, preventing injection of malicious scaling or workspace commands. | 2026-03-20T06:12:49.837Z
IFC-MAIN-028 | VER-MAIN-069 | verifies | Integration test for Workspace Safety Enforcer to Kinematics Engine proximity constraint channel | Validates proximity enforcement latency and velocity limiting under anatomy-boundary approach, confirming patient safety guarantee. | 2026-03-20T06:09:58.607Z
IFC-MAIN-027 | VER-MAIN-068 | verifies | Integration test for Kinematics Engine to Joint Servo Controller joint setpoint delivery | Validates joint setpoint timing, accuracy, and fault rejection at the final joint-level interface. | 2026-03-20T06:09:40.534Z
IFC-MAIN-026 | VER-MAIN-067 | verifies | Integration test for Trajectory Generator to Kinematics Engine pose delivery | Validates S-curve interpolation quality and abort-on-violation at this safety-critical internal interface. | 2026-03-20T06:09:22.741Z
IFC-MAIN-025 | VER-MAIN-066 | verifies | Integration test for Motion Scaling to Trajectory Generator interface | Validates scaling accuracy, safety clamping, and data integrity at the MC subsystem internal boundary. | 2026-03-20T06:09:07.606Z
IFC-MAIN-024 | VER-MAIN-065 | verifies | Integration test for Tremor Rejection Filter output interface | Validates that IFC-MAIN-024 spectral attenuation and latency requirements are met at the component boundary under realistic tremor-frequency injection. | 2026-03-20T06:08:35.811Z
SYS-MAIN-005 | SUB-MAIN-086 | derives | Real-Time Compute Node thermal management derives from degraded operation continuity requirement | SYS-MAIN-005 requires 60 seconds of controlled operation after a single fault; SUB-MAIN-086 preserves real-time scheduling under thermal stress and defines the transition condition that ensures safe hold is entered before hardware damage can cause unpredictable behavior. | 2026-03-20T06:07:39.316Z
SYS-MAIN-010 | SUB-MAIN-085 | derives | JSC fault isolation and brake engagement derives from system safe-stop requirement | SYS-MAIN-010 requires all joints to reach safe hold within 150ms of any stop command; SUB-MAIN-085 implements the servo-level detection and brake engagement that satisfies this at the joint level (2ms brake engage vs. 150ms system budget). | 2026-03-20T06:07:38.663Z
SYS-MAIN-007 | SUB-MAIN-084 | derives | Kinematics Engine singularity handling derives from end-to-end latency and motion fidelity requirement | SYS-MAIN-007 requires motion commands delivered within 1ms end-to-end; SUB-MAIN-084 specifies DLS singularity avoidance that maintains this timing guarantee and directional accuracy through near-singular configurations. | 2026-03-20T06:07:38.001Z
SYS-MAIN-005 | SUB-MAIN-083 | derives | WSE degraded mode derives from system degraded operation requirement | SYS-MAIN-005 requires 60 seconds of controlled operation after any single fault; SUB-MAIN-083 provides the fallback workspace restriction that maintains controlled instrument state when anatomy mesh is unavailable. | 2026-03-20T06:07:24.509Z
SYS-MAIN-001 | SUB-MAIN-082 | derives | Workspace safety enforcement derives from system precision and safety requirements | SYS-MAIN-001 requires 1mm RMS tip accuracy; SUB-MAIN-082 implements the proximity enforcement that prevents gross workspace violations by constraining approach velocity near anatomy boundaries. | 2026-03-20T06:07:23.867Z
SYS-MAIN-009 | SUB-MAIN-080 | derives | Tremor filter cutoff and attenuation spec derived from system tremor requirement | SYS-MAIN-009 requires 40dB attenuation above 6Hz at the instrument tip; SUB-MAIN-080 allocates this requirement to the Tremor Rejection Filter as the specific mechanism, specifying the filter order and topology needed to achieve the system figure. | 2026-03-20T06:07:23.248Z
SYS-MAIN-006 | SUB-MAIN-079 | derives | System-level patient safety/sterile compliance drives galvanic isolation requirement | SYS-MAIN-006 mandates IEC 60601-1 compliance for patient-contact components. SUB-MAIN-079 is the specific IEC 60601-1 Clause 8.8 reinforced insulation requirement as applied to the inter-cart link — the fibre link crosses the patient/non-patient boundary, making galvanic isolation not just good practice but a regulatory requirement mandated by SYS-MAIN-006. | 2026-03-20T05:14:26.021Z
SYS-MAIN-007 | SUB-MAIN-078 | derives | System-level transmission latency mandate requires transparent fibre failover | SYS-MAIN-007 requires continuous <1ms command delivery. SUB-MAIN-078 ensures the fibre link switching mechanism is transparent to the kinematic pipeline — a failover that dropped frames would interrupt command delivery and violate SYS-MAIN-007 as surely as a sustained link failure. | 2026-03-20T05:14:25.402Z
SYS-MAIN-015 | SUB-MAIN-076 | derives | System-level data recording mandate drives PDR dual-mirror NVMe integrity requirement | SYS-MAIN-015 requires complete kinematic data recording for the entire procedure. SUB-MAIN-076 implements the storage-layer redundancy necessary for SYS-MAIN-015 to remain satisfied even under a single drive failure — without dual-mirror NVMe, a single NVMe failure could cause total data loss for the remainder of a long procedure, directly violating SYS-MAIN-015. | 2026-03-20T05:14:24.771Z
SYS-MAIN-002 | SUB-MAIN-073 | derives | System-level single-fault detection drives NMC clock synchronisation accuracy | SYS-MAIN-002 requires detection of any single-point failure. A timing skew beyond 500ns between servo nodes manifests as coordinated inter-joint position error — a failure mode that SYS-MAIN-002's fault detection chain must catch. SUB-MAIN-073 defines the threshold below which inter-node errors remain within acceptable position tolerance, preventing false-positive fault detections from clock drift. | 2026-03-20T05:14:11.274Z
SYS-MAIN-007 | IFC-MAIN-040 | derives | System-level command transmission latency drives CC-to-ICFL frame format and timing | IFC-MAIN-040 defines the frame format and 1kHz rate at the Console Computer output — the first link in the command chain that SYS-MAIN-007 constrains to 1ms end-to-end. The explicit sequence number and CRC exist to satisfy SYS-MAIN-007's implicit requirement for reliable, ordered delivery within that budget. | 2026-03-20T05:14:10.672Z
IFC-MAIN-025 | VER-MAIN-064 | verifies | Ring-buffer latency and drop test for Motion Scaling to Trajectory Generator interface | VER-MAIN-064 fills the gap in verification coverage for IFC-MAIN-025, which previously had no verification entry. The test validates the shared-memory path latency and data fidelity that IFC-MAIN-025 specifies. | 2026-03-20T05:10:50.376Z
IFC-MAIN-043 | VER-MAIN-063 | verifies | PCIe DMA soak test verifying Real-Time Compute Node to PDR data integrity | VER-MAIN-063 provides the sequence-number continuity check that is the only way to prove zero frame loss through the PCIe DMA path under sustained NVMe load, which is the critical pass criterion of IFC-MAIN-043. | 2026-03-20T05:10:49.779Z
IFC-MAIN-042 | VER-MAIN-062 | verifies | Bench EtherCAT integration test for NMC to Joint Servo Controller interface | VER-MAIN-062 verifies the PDO structure, timing, and per-node fault isolation in IFC-MAIN-042 at the subsystem bench level before patient-side cart integration. | 2026-03-20T05:10:49.186Z
IFC-MAIN-041 | VER-MAIN-061 | verifies | Integration test for fibre-to-RTPE frame delivery timing and CRC validity | VER-MAIN-061 directly tests the timing and validity flag assertions in IFC-MAIN-041, providing the only measurable evidence of SERDES jitter compliance at the RTPE boundary. | 2026-03-20T05:10:37.016Z
SYS-MAIN-007 | SUB-MAIN-077 | derives | System-level command transmission latency cascades to fibre link propagation budget | SYS-MAIN-007 requires surgeon hand motion commands to reach the instrument tip within 1ms. The Inter-Cart Fibre Link optical propagation latency is one of five sequential latency contributors; the 100-microsecond ceiling in SUB-MAIN-077 represents the allocated fibre link portion of the total 1ms budget, leaving 200μs per remaining stage. | 2026-03-20T05:10:36.143Z
SYS-MAIN-015 | SUB-MAIN-075 | derives | System-level 1kHz recording mandate drives PDR sample rate and write latency | SYS-MAIN-015 specifies 1kHz kinematic recording. SUB-MAIN-075 is the direct realisation of that requirement in the Procedure Data Recorder, adding the 10ms write latency ceiling that translates the system-level recording mandate into a verifiable hardware performance constraint. | 2026-03-20T05:10:25.055Z
SYS-MAIN-016 | SUB-MAIN-074 | derives | System degraded-mode arm isolation cascades to NMC per-node fault containment | SYS-MAIN-016 requires that loss of one arm's servo communication does not affect remaining arms. SUB-MAIN-074 implements this at the bus level — the NMC must isolate the faulted EtherCAT node within a single bus cycle to prevent bus contention from disrupting healthy nodes. The requirement exists because SYS-MAIN-016 cannot be met without active fault isolation at the network layer. | 2026-03-20T05:10:24.430Z
SYS-MAIN-010 | SUB-MAIN-072 | derives | System-level emergency stop response drives RTPE comms-loss controlled halt | SYS-MAIN-010 mandates controlled arm halt on any safety fault including communication loss. SUB-MAIN-072 implements the RTPE's portion of this requirement — detecting frame absence and issuing zero-velocity commands — and exists directly because SYS-MAIN-010 applies to communication-path failures, not only hardwired stop signals. | 2026-03-20T05:10:12.555Z
SYS-MAIN-001 | SUB-MAIN-071 | derives | System end-to-end latency budget cascades to RTPE frame jitter ceiling | SYS-MAIN-001 requires <1ms end-to-end latency. The RTPE frame scheduler is the first link in the surgeon-to-instrument chain and its jitter directly reduces the available budget for downstream algorithmic stages; the 1-microsecond ceiling in SUB-MAIN-071 exists specifically because SYS-MAIN-001 leaves only a small jitter margin. | 2026-03-20T05:10:11.902Z
SUB-MAIN-070 | VER-MAIN-060 | verifies | Console failsafe injection test | VER-MAIN-060 verifies SUB-MAIN-070 by direct exception injection, measuring safe-state timing and data retention under controlled fault conditions. | 2026-03-20T04:15:53.776Z
SUB-MAIN-069 | VER-MAIN-059 | verifies | IEC 62304 compliance audit | VER-MAIN-059 verifies SUB-MAIN-069 via documentary audit of the IEC 62304 software lifecycle record and MDR Annex I checklist. | 2026-03-20T04:15:53.125Z
SYS-MAIN-002 | SUB-MAIN-070 | derives | Console computer safe-hold on exception derives from system single-point failure detection requirement | SYS-MAIN-002 requires the system to detect any single-point failure and transition to safe state. A console computer software exception is a single-point failure in the primary control node; SUB-MAIN-070 specifies the maximum safe-state entry time and data retention for this failure mode. | 2026-03-20T04:15:52.206Z
SYS-MAIN-002 | SUB-MAIN-069 | derives | Console computer IEC 62304 Class C compliance derives from system safety integrity requirements | SYS-MAIN-002 mandates detection and safe-state transition on any single-point failure. IEC 62304 Class C development is the recognised mechanism to achieve the software integrity level needed for a system where console software failure could directly cause patient harm without an opportunity for detection. | 2026-03-20T04:15:51.541Z
SUB-MAIN-068 | VER-MAIN-058 | verifies | Handle disengagement detection latency and propagation test | VER-MAIN-058 verifies SUB-MAIN-068 by measuring sensor-to-inhibit-signal latency and confirming downstream instrument motion cessation to close the complete inhibit chain. | 2026-03-20T04:11:53.077Z
SUB-MAIN-067 | VER-MAIN-057 | verifies | Haptic force rendering step response test | VER-MAIN-057 verifies SUB-MAIN-067 through step force input characterisation, directly measuring rendering latency and steady-state accuracy at the master arm fingertips. | 2026-03-20T04:11:52.472Z
SUB-MAIN-066 | VER-MAIN-056 | verifies | SIP pose rate and quantisation spectral test | VER-MAIN-056 verifies SUB-MAIN-066 using spectral analysis to confirm the 1kHz pose rate and quantisation floor measurement from recorded trajectory data. | 2026-03-20T04:11:51.860Z
IFC-MAIN-039 | VER-MAIN-055 | verifies | EtherCAT bidirectional latency and continuity test | VER-MAIN-055 verifies IFC-MAIN-039 by measuring actual bidirectional latency under sustained load, confirming the 1ms criterion at the 99.9th percentile over 30 minutes. | 2026-03-20T04:11:51.242Z
SUB-MAIN-061 | VER-MAIN-054 | verifies | Voice module degraded mode fault injection test | VER-MAIN-054 verifies SUB-MAIN-061 through software fault injection, confirming graceful degradation without motion interruption when the voice module is disabled. | 2026-03-20T04:11:50.673Z
SUB-MAIN-060 | VER-MAIN-053 | verifies | Console startup self-test timing verification | VER-MAIN-053 verifies SUB-MAIN-060 by cold-start testing across 10 power cycles, confirming the 90-second self-test completion criterion under realistic thermal and power conditions. | 2026-03-20T04:11:50.053Z
SUB-MAIN-058 | VER-MAIN-052 | verifies | Authentication demonstration test | VER-MAIN-052 verifies SUB-MAIN-058 by demonstration, confirming that motion is always inhibited without valid authentication and that the audit trail is preserved. | 2026-03-20T04:11:49.469Z
SUB-MAIN-057 | VER-MAIN-051 | verifies | Voice dispatch latency covered by combined WER/latency test | SUB-MAIN-057 latency criterion is verified in the same test session as SUB-MAIN-056, as dispatch latency is measured per utterance alongside WER during the 500-item vocabulary test. | 2026-03-20T04:11:33.194Z
SUB-MAIN-056 | VER-MAIN-051 | verifies | Voice WER and dispatch latency test | VER-MAIN-051 covers both WER and dispatch latency requirements in a single combined test, as both properties are measured simultaneously from the same utterance capture. | 2026-03-20T04:11:32.560Z
SUB-MAIN-055 | VER-MAIN-050 | verifies | Foot pedal CAN latency mechanical test | VER-MAIN-050 uses mechanical jig actuation to verify SUB-MAIN-055 latency directly at the CAN frame level, eliminating human timing uncertainty from the measurement. | 2026-03-20T04:11:31.903Z
IFC-MAIN-037 | VER-MAIN-049 | verifies | USB session management interface integration test | IFC-MAIN-037 defines session management message integrity over USB 3.0. VER-MAIN-049 verifies this through sustained injection testing, validating both latency and reliability characteristics at the interface boundary. | 2026-03-20T04:11:31.277Z
SYS-MAIN-002 | SUB-MAIN-059 | derives | Arm positioning lockout during OPERATIONAL state derives from safe-state requirements | SYS-MAIN-002 requires the system to transition to safe state on any single-point failure. Arm movement during OPERATIONAL mode changes the kinematic relationship between master and slave, which is a patient safety hazard equivalent to a control failure. SUB-MAIN-059 implements the physical guard that prevents this class of failure. | 2026-03-20T04:11:18.251Z
SYS-MAIN-010 | SUB-MAIN-055 | derives | Foot pedal CAN latency derives from E-stop command propagation requirement | SYS-MAIN-010 requires the system to enter safe state within 1 second of any emergency stop command including surgeon foot pedal. The 50ms CAN transmission budget for all pedal events is derived by dividing the total 1-second E-stop latency budget across its constituent links; the pedal interface is the first link in the chain. | 2026-03-20T04:11:17.570Z
SYS-MAIN-001 | IFC-MAIN-039 | derives | SIP-CC EtherCAT interface derives from system <1ms motion scaling requirement | SYS-MAIN-001 specifies <1ms end-to-end motion scaling. The bidirectional haptic bus between Surgeon Interface Panel and Console Computer must consume no more than 1ms in each direction to meet this system-level budget; IFC-MAIN-039 sets this allocation explicitly. | 2026-03-20T04:11:01.472Z
SYS-MAIN-002 | SUB-MAIN-068 | derives | Handle disengagement detection derives from single-point failure detection requirement | SYS-MAIN-002 requires detection and safe-state transition on any single-point failure. Loss of surgeon hand presence is a safety-critical event that the system must detect and act on; SUB-MAIN-068 specifies the detection latency budget for this specific failure mode at the console. | 2026-03-20T04:11:00.866Z
SYS-MAIN-004 | SUB-MAIN-067 | derives | SIP haptic force rendering derives from system force sensing and feedback requirement | SYS-MAIN-004 requires instrument-tissue force measurement with haptic feedback to the surgeon. SUB-MAIN-067 specifies the rendering half of the closed haptic loop at the Surgeon Interface Panel - the end-effector through which the system requirement's feedback is delivered. | 2026-03-20T04:11:00.280Z
SYS-MAIN-001 | SUB-MAIN-066 | derives | SIP 1kHz pose rate derives from system <1ms motion scaling requirement | SYS-MAIN-001 requires end-to-end motion scaling latency below 1ms. The Surgeon Interface Panel must sample at 1kHz or higher to contribute no more than 1ms sampling latency to the pipeline; this is the direct kinematic chain from surgeon hand to tooltip that the system requirement governs. | 2026-03-20T04:10:48.510Z
SUB-MAIN-064 | VER-MAIN-028 | verifies | Master Handle STANDBY backdrive force verification | REQ-SESURGICALROBOT-028 is the test procedure that verifies the zero-force STANDBY requirement in SUB-MAIN-064, covering all handle positions and mode-transition transients. | 2026-03-20T03:09:58.501Z
SUB-MAIN-063 | uht-bot:se-surgical-robot:REQ-SESURGICALROBOT-027 | verifies | HIL stability test for haptic feedback loop | REQ-SESURGICALROBOT-027 defines the HIL stability sweep test that verifies the no-oscillation requirement in SUB-MAIN-063. | 2026-03-20T03:09:51.724Z
SUB-MAIN-062 | uht-bot:se-surgical-robot:REQ-SESURGICALROBOT-026 | verifies | SIL 2 verification for Haptic Controller | REQ-SESURGICALROBOT-026 defines the independent safety assessment procedure that verifies the SIL 2 development claims in SUB-MAIN-062. | 2026-03-20T03:09:41.473Z
SUB-MAIN-024 | VER-MAIN-025 | verifies | Network isolation VER entry | REQ-SESURGICALROBOT-025 defines the penetration test procedure that verifies the network isolation claimed in REQ-SESURGICALROBOT-024. | 2026-03-20T03:09:20.056Z
SUB-MAIN-024 | SYS-MAIN-002 | derives | Console Computer network isolation supports fault prevention required by SYS-MAIN-002 | Network isolation (REQ-SESURGICALROBOT-024) is required to prevent a cybersecurity-induced fault from reaching the surgical control bus — without isolation, a network compromise could trigger exactly the kind of control loss that SYS-MAIN-002 mandates detection and safe-state response for. | 2026-03-20T03:09:10.104Z
SUB-MAIN-065 | SYS-MAIN-002 | derives | SUB-MAIN-065 IEC 60601-1 compliance is the regulatory path for Haptic subsystem safety | IEC 60601-1 compliance (SUB-MAIN-065) is the regulatory mechanism through which the safety integrity required by SYS-MAIN-002 is certified for the Haptic Feedback Subsystem; the standard mandates the failure mode analysis, leakage current limits, and safety classification that underpin safe energy delivery. | 2026-03-20T03:06:45.052Z
ARC-MAIN-014 | SYS-MAIN-004 | derives | ARC-MAIN-014 is the enabling architecture for haptic force sensing in SYS-MAIN-004 | The galvanic isolation architecture in ARC-MAIN-014 is the safety boundary that enables force sensing to be transmitted from the sterile field to the surgeon console without violating IEC 60601-1 leakage current limits, which is the enabling architectural decision for SYS-MAIN-004. | 2026-03-20T03:06:16.129Z
ARC-MAIN-012 | SYS-MAIN-010 | derives | ARC-MAIN-012 supports emergency stop architecture as required by SYS-MAIN-010 | ARC-MAIN-012 records the architectural decision to hardware-wire safety-critical foot pedal inputs independently of the Console Computer, directly supporting the emergency stop response requirement in SYS-MAIN-010. | 2026-03-20T03:06:15.478Z
ARC-MAIN-009 | SYS-MAIN-017 | derives | ARC-MAIN-009 records design rationale supporting electrosurgical energy delivery | This architecture decision records the dual-modality RF/ultrasonic selection that directly implements the electrosurgical energy delivery requirement in SYS-MAIN-017. | 2026-03-20T03:06:04.835Z
SYS-MAIN-001 | VER-MAIN-046 | verifies | End-to-end console-to-tip test verifies surgeon input-to-instrument motion chain | VER-MAIN-046 tests the full console-to-instrument-tip chain which is the primary verification vehicle for the motion scaling and command transmission system requirements. | 2026-03-20T02:08:43.678Z
SYS-MAIN-016 | SUB-MAIN-061 | derives | Voice module failure degraded mode derives from system degraded operation requirement | SYS-MAIN-016 mandates full function retention when one arm loses servo communication; the Surgeon Input Console must equally maintain full function when a non-critical input module (voice) fails. | 2026-03-20T02:08:43.039Z
SYS-MAIN-002 | SUB-MAIN-060 | derives | Console startup self-test derives from system fault detection requirement | SYS-MAIN-002 requires single-point failure detection; startup self-test is the mechanism by which the Console Computer verifies its own integrity before entering operational state. | 2026-03-20T02:08:42.364Z
SYS-MAIN-007 | SUB-MAIN-057 | derives | Voice command dispatch latency derives from surgeon input transmission requirement | SYS-MAIN-007 requires real-time command transmission; voice command dispatch latency directly affects end-to-end input responsiveness. | 2026-03-20T02:08:41.742Z
SYS-MAIN-007 | SUB-MAIN-056 | derives | Voice command recognition derives from surgeon input transmission requirement | SYS-MAIN-007 requires all surgeon input modalities to be transmitted to the robot; voice commands are an input channel that must meet recognition performance to be useful in the surgical workflow. | 2026-03-20T02:08:41.021Z
SYS-MAIN-007 | IFC-MAIN-037 | derives | Console-to-RTPE session interface derives from system motion command transmission requirement | SYS-MAIN-007 requires surgeon hand motion commands to be transmitted to instruments; the Console Computer to Real-Time Protocol Engine interface is the primary inter-subsystem path for this command data. | 2026-03-20T02:08:40.339Z
SUB-MAIN-063 | VER-MAIN-048 | verifies | Stiffness sweep test verifies haptic stability across full operating envelope | Stability is a boundary phenomenon — only full envelope sweep can confirm the requirement is met. | 2026-03-20T02:07:09.133Z
SUB-MAIN-062 | VER-MAIN-047 | verifies | SIL2 certification assessment verifies haptic controller integrity level | Third-party SIL2 assessment is the only accepted verification method for IEC 62061 safety integrity claims. | 2026-03-20T02:07:08.486Z
SYS-MAIN-001 | SUB-MAIN-064 | derives | Backdrive compliance derives from motion scaling requirement — transparent kinaesthetic operation required for 1:1 to 10:1 scaling modes | SYS-MAIN-001 specifies surgeon hand motion transmission; force-transparent back-drive is a precondition for accurate motion capture across all scaling ratios including 1:1. | 2026-03-20T02:06:35.347Z
SYS-MAIN-004 | SUB-MAIN-063 | derives | Stability requirement derives from system force measurement requirement — unstable haptic loop corrupts the force measurement and feedback chain | SUB-MAIN-063 stability constraint is necessary because SYS-MAIN-004 requires reliable force measurement during contact; haptic loop instability introduces force signal noise that corrupts measurement integrity. | 2026-03-20T02:06:34.669Z
SYS-MAIN-004 | SUB-MAIN-062 | derives | SIL2 integrity requirement for haptic controller derives from system force measurement safety requirement | SYS-MAIN-004 mandates 0.1N force measurement at patient contact — the safety integrity level on the render path is a direct consequence of this safety-critical measurement function. | 2026-03-20T02:06:34.014Z
SUB-MAIN-059 | VER-MAIN-045 | verifies | State-transition lock-out test for Arm Positioning System | Verifies the timing bound (500ms) and command rejection behaviour specified in SUB-MAIN-059. | 2026-03-20T01:10:22.328Z
IFC-MAIN-038 | VER-MAIN-044 | verifies | USB traffic analysis to verify no raw audio leaves Voice Command Module | Verifies the privacy constraint in IFC-MAIN-038: that raw audio is not transmitted outside the module. | 2026-03-20T01:10:10.070Z
IFC-MAIN-036 | VER-MAIN-043 | verifies | Integration test for Foot Pedal clutch interface to Motion Control | Verifies latency under load — the only measurable acceptance criterion of IFC-MAIN-036 that cannot be verified at component level. | 2026-03-20T01:10:08.839Z
IFC-MAIN-035 | VER-MAIN-042 | verifies | Integration test for Foot Pedal Array to Energy Delivery Controller interface | Verifies both latency compliance (10ms message delivery) and electrical isolation requirement, which are the two measurable acceptance criteria of IFC-MAIN-035. | 2026-03-20T01:09:49.946Z
SYS-MAIN-015 | SUB-MAIN-058 | derives | System recording requirement derives console authentication logging | SYS-MAIN-015 requires recording of all system events at 1kHz. SUB-MAIN-058 specifically derives the authentication and case-start logging requirement, which is a system event that must be captured as part of the complete procedure record. | 2026-03-20T01:09:34.231Z
SYS-MAIN-010 | SUB-MAIN-055 | derives | E-stop system requirement derives foot pedal latency requirement | SYS-MAIN-010 requires emergency stop within 200ms from any input including surgeon foot pedal. SUB-MAIN-055 allocates 50ms to the pedal event transmission as the first budget element in the end-to-end E-stop latency chain. | 2026-03-20T01:09:33.554Z
SYS-MAIN-013 | SUB-MAIN-023 | derives | Power supply specification derived from 8-hour operational requirement | SYS-MAIN-013 requires 8 consecutive hours of full operation; each subsystem power allocation specifies the supply voltage and current budget that the Power Management Subsystem must deliver to sustain the system for this duration. | 2026-03-20T00:39:21.549Z
SYS-MAIN-013 | SUB-MAIN-022 | derives | Power supply specification derived from 8-hour operational requirement | SYS-MAIN-013 requires 8 consecutive hours of full operation; each subsystem power allocation specifies the supply voltage and current budget that the Power Management Subsystem must deliver to sustain the system for this duration. | 2026-03-20T00:39:20.875Z
SYS-MAIN-013 | SUB-MAIN-021 | derives | Power supply specification derived from 8-hour operational requirement | SYS-MAIN-013 requires 8 consecutive hours of full operation; each subsystem power allocation specifies the supply voltage and current budget that the Power Management Subsystem must deliver to sustain the system for this duration. | 2026-03-20T00:39:20.212Z
SYS-MAIN-013 | SUB-MAIN-020 | derives | Power supply specification derived from 8-hour operational requirement | SYS-MAIN-013 requires 8 consecutive hours of full operation; each subsystem power allocation specifies the supply voltage and current budget that the Power Management Subsystem must deliver to sustain the system for this duration. | 2026-03-20T00:39:19.588Z
SYS-MAIN-013 | SUB-MAIN-019 | derives | Power supply specification derived from 8-hour operational requirement | SYS-MAIN-013 requires 8 consecutive hours of full operation; each subsystem power allocation specifies the supply voltage and current budget that the Power Management Subsystem must deliver to sustain the system for this duration. | 2026-03-20T00:39:18.918Z
SYS-MAIN-013 | SUB-MAIN-018 | derives | Power supply specification derived from 8-hour operational requirement | SYS-MAIN-013 requires 8 consecutive hours of full operation; each subsystem power allocation specifies the supply voltage and current budget that the Power Management Subsystem must deliver to sustain the system for this duration. | 2026-03-20T00:39:18.296Z
SYS-MAIN-013 | SUB-MAIN-017 | derives | Power supply specification derived from 8-hour operational requirement | SYS-MAIN-013 requires 8 consecutive hours of full operation; each subsystem power allocation specifies the supply voltage and current budget that the Power Management Subsystem must deliver to sustain the system for this duration. | 2026-03-20T00:39:17.610Z
SYS-MAIN-017 | VER-MAIN-039 | verifies | End-to-end system integration test for energy delivery | VER-MAIN-039 directly tests the activation latency, deactivation latency, and REM safety requirements defined in SYS-MAIN-017. | 2026-03-20T00:38:33.940Z
SYS-MAIN-017 | SUB-MAIN-053 | derives | Activation timeout safety derived from controlled energy delivery requirement | SYS-MAIN-017 requires controlled delivery; the 15s/5s timeouts define the upper bound of 'controlled' activation duration — any longer risks unintended tissue damage. | 2026-03-20T00:38:33.356Z
SYS-MAIN-017 | SUB-MAIN-052 | derives | Tissue effect monitoring derived from controlled energy delivery requirement | SYS-MAIN-017 requires controlled energy delivery; TEM automatic cutoff on seal endpoint is a key control mechanism preventing over-application — directly derived from the 'controlled' qualifier. | 2026-03-20T00:38:32.700Z
SUB-MAIN-054 | VER-MAIN-041 | verifies | Type CF leakage current test per IEC 60601-1 | Electrical safety testing is the only valid verification method for leakage current; analysis alone cannot substitute. | 2026-03-20T00:37:48.146Z
SUB-MAIN-051 | VER-MAIN-040 | verifies | Return electrode monitor threshold and inhibit response test | This directly tests the 135Ω threshold and 500ms response that define the SUB-MAIN-051 safety requirement. | 2026-03-20T00:37:46.869Z
IFC-MAIN-034 | VER-MAIN-038 | verifies | E-STOP response latency test for EDC-Safety interface | Repeated E-STOP testing validates that the 20ms response is reliable under varying energy delivery states, not just a single nominal case. | 2026-03-20T00:37:10.958Z
IFC-MAIN-033 | VER-MAIN-037 | verifies | Impedance feedback and endpoint detection test for TEM-ESG interface | This test verifies both the sampling rate and the shutoff latency on the same bench setup, providing confidence in the end-to-end impedance feedback chain. | 2026-03-20T00:37:09.684Z
IFC-MAIN-032 | VER-MAIN-036 | verifies | Hardware interlock line response test for REM-ESG interface | Safety interlock verification must be performed at the physical layer to confirm it cannot be bypassed by software failure. | 2026-03-20T00:37:08.442Z
IFC-MAIN-031 | VER-MAIN-035 | verifies | RS-485 telemetry rate test for EDC-UEM interface | The 50Hz minimum is the basis for the blade temperature inhibit response; this test verifies the link performs reliably under operational load. | 2026-03-20T00:36:32.788Z
IFC-MAIN-030 | VER-MAIN-034 | verifies | CAN bus timing and error test for EDC-ESG interface | This integration test proves the 5ms acknowledgement and error-free operation requirements of IFC-MAIN-030, which cannot be validated by inspection alone. | 2026-03-20T00:36:31.556Z
SYS-MAIN-017 | SUB-MAIN-050 | derives | Mutual exclusion safety derived from energy delivery system requirement | SYS-MAIN-017 does not explicitly state mutual exclusion, but 'controlled energy delivery' in a dual-modality system implies exactly one modality active at a time — SUB-MAIN-050 makes this explicit. | 2026-03-20T00:36:12.327Z
SYS-MAIN-017 | SUB-MAIN-049 | derives | Ultrasonic energy specification derived from energy delivery system requirement | SYS-MAIN-017 covers both modalities; SUB-MAIN-049 specifies the ultrasonic branch performance parameters. | 2026-03-20T00:36:11.663Z
SYS-MAIN-017 | SUB-MAIN-048 | derives | Activation/deactivation latency derived from energy delivery system requirement | SYS-MAIN-017's ≤100ms activation and ≤50ms deactivation bounds cascade directly to the generator hardware responsible for physical energy delivery. | 2026-03-20T00:36:11.070Z
SYS-MAIN-017 | SUB-MAIN-047 | derives | RF power specification derived from energy delivery system requirement | SYS-MAIN-017 mandates controlled RF energy delivery; SUB-MAIN-047 specifies the power envelope and accuracy that makes 'controlled' measurable at the generator level. | 2026-03-20T00:36:10.465Z
ARC-MAIN-008 | SYS-MAIN-002 | derives | Isolated safety power supply architecture derives from single-point failure safety requirement | ARC-MAIN-008 specifies a dedicated isolated 24V supply for safety circuits; derives from SYS-MAIN-002 requiring detection and safe-state entry for any single-point failure — hardware isolation prevents a main bus fault from defeating safety supervision. | 2026-03-20T00:19:21.985Z
ARC-MAIN-007 | SYS-MAIN-001 | derives | Dedicated trajectory generator derives from control loop latency and motion safety requirement | ARC-MAIN-007 specifies a dedicated trajectory generation stage with S-curve profiling; derives from SYS-MAIN-001 requiring <1ms end-to-end control loop with selectable scaling — S-curve limiting ensures tissue-safe acceleration within the latency budget. | 2026-03-20T00:19:21.401Z
ARC-MAIN-006 | SYS-MAIN-014 | derives | Cable-driven remote motor architecture derives from sterilisation compatibility requirement | ARC-MAIN-006 specifies cable-driven actuation with remote motors in the IDU; derives from SYS-MAIN-014 requiring patient-side surfaces to be compatible with sterilisation — keeping motors proximal avoids autoclave-incompatible components in the sterile field. | 2026-03-20T00:19:20.802Z
ARC-MAIN-005 | SYS-MAIN-007 | derives | FPGA real-time communications architecture derives from motion command latency requirement | ARC-MAIN-005 specifies FPGA-based real-time protocol for communications; derives from SYS-MAIN-007 requiring motion command transmission with <1ms end-to-end latency. | 2026-03-20T00:19:20.189Z
ARC-MAIN-004 | SYS-MAIN-004 | derives | Galvanic isolation at force conditioner derives from instrument-tissue force measurement requirement | ARC-MAIN-004 specifies galvanic isolation at the Force Signal Conditioner; derives from SYS-MAIN-004 requiring instrument-tissue force measurement — electrical isolation ensures patient safety and signal integrity for force sensing. | 2026-03-20T00:19:19.544Z
ARC-MAIN-003 | SYS-MAIN-003 | derives | FPGA image pipeline architecture derives from stereo HD video latency requirement | ARC-MAIN-003 specifies FPGA-based deterministic image processing with separated optical and processing stages; derives from SYS-MAIN-003 requiring stereoscopic HD video with <33ms display latency. | 2026-03-20T00:19:00.214Z
ARC-MAIN-002 | SYS-MAIN-009 | derives | Cartesian-space pipeline with tremor filter stage derives from tremor attenuation requirement | ARC-MAIN-002 specifies the Cartesian-space pipeline with independent tremor filter stage; derives from SYS-MAIN-009 requiring >40dB tremor attenuation above 6Hz. | 2026-03-20T00:18:59.579Z
ARC-MAIN-001 | SYS-MAIN-002 | derives | Safety processor architecture derives from single-point failure detection requirement | ARC-MAIN-001 documents the dedicated safety processor decision; derives from SYS-MAIN-002 requiring single-point failure detection. | 2026-03-20T00:18:58.952Z
SYS-MAIN-003 | SUB-MAIN-046 | derives | Video integrity safety constraint derived from stereo HD video system requirement | SYS-MAIN-003 requires synchronised stereo HD video to the surgeon; the integrity constraint in SUB-MAIN-046 is the safety enforcement mechanism that ensures what reaches the surgeon meets the HD synchronised video standard, rather than a degraded or corrupted frame. | 2026-03-20T00:09:53.744Z
SYS-MAIN-013 | SUB-MAIN-044 | derives | Power sequencing derived from 8-hour operational duration requirement | SYS-MAIN-013 requires 8 consecutive hours of operation; SUB-MAIN-044 specifies the power sequencing order that ensures controlled state transitions during this period without creating unsafe intermediate states. | 2026-03-20T00:09:52.439Z
SUB-MAIN-043 | VER-MAIN-033 | verifies | Full-load UPS duration test at 80% charge for SUB-MAIN-043 | Tests the 30-minute operational continuity requirement at pre-conditioned battery state (80%), representing realistic worst-case. Duration measured to first failure or safe shutdown — directly verifies the 30-minute requirement floor. | 2026-03-20T00:08:39.445Z
IFC-MAIN-029 | VER-MAIN-032 | verifies | Integration test for IFC-MAIN-029 under mains loss | Tests the voltage continuity specification at the exact interface boundary under the worst-case transient. Pass/fail criterion is directly derived from the 10ms interruption limit and 22-26V range in IFC-MAIN-029. | 2026-03-20T00:08:38.825Z
SYS-MAIN-002 | IFC-MAIN-029 | derives | Power-to-E-stop interface derived from single-point failure requirement | SYS-MAIN-002 requires no single power fault to prevent safe-state. IFC-MAIN-029 specifies the voltage continuity on the E-stop contactor coil supply — the specific interface that must not fail during a mains power fault. | 2026-03-20T00:08:24.840Z
SYS-MAIN-002 | SUB-MAIN-045 | derives | Aux power isolation derived from single-point failure detection requirement | SYS-MAIN-002 requires the system to detect and respond to any single-point power fault. SUB-MAIN-045 implements the auxiliary power isolation that ensures safety supervision persists through main bus failures — a specific instantiation of the single-point failure survivability requirement. | 2026-03-20T00:08:06.397Z
SYS-MAIN-005 | SUB-MAIN-043 | derives | UPS 30-min backup derived from 30-min operational capability requirement | SYS-MAIN-005 requires 30 minutes of controlled arm capability on backup power. SUB-MAIN-043 implements this specifically for the Power Management Subsystem by specifying the UPS battery minimum capacity. | 2026-03-20T00:08:05.625Z
SYS-MAIN-002 | SUB-MAIN-042 | derives | Motion Scaling Module velocity limit implements per-stage fault detection for input faults | SYS-MAIN-002 requires single-point failure detection; SUB-MAIN-042 allocates velocity-limit detection to the Motion Scaling Module, catching sensor faults or uncontrolled inputs at the earliest possible pipeline stage | 2026-03-19T23:37:00.865Z
SYS-MAIN-012 | SUB-MAIN-041 | derives | Trajectory acceleration limits derive from system-level tissue force safety constraint | SYS-MAIN-012 limits tip force to 5N; SUB-MAIN-041 enforces acceleration limits at the trajectory level to prevent motion commands that would generate unsafe tissue contact forces | 2026-03-19T23:37:00.250Z
SYS-MAIN-002 | SUB-MAIN-041 | derives | Trajectory Generator clamp implements per-stage fault detection for motion violations | SYS-MAIN-002 requires single-point failure detection and safe-state transition; SUB-MAIN-041 allocates this to the Trajectory Generator for workspace and acceleration violations, contributing to the overall fault detection chain | 2026-03-19T23:36:59.624Z
SYS-MAIN-001 | VER-MAIN-031 | verifies | End-to-end pipeline test validates system-level motion control performance | SYS-MAIN-001 specifies tip repeatability and scaling; this test exercises the complete pipeline that implements those capabilities, validating the integrated behaviour that component tests cannot | 2026-03-19T23:33:02.442Z
SYS-MAIN-007 | SUB-MAIN-039 | derives | Trajectory Generator 1kHz rate contributes to end-to-end latency budget | SYS-MAIN-007 requires <100ms total pipeline latency; the Trajectory Generator's 1kHz cycle and S-curve profiling add at most 1ms to the pipeline, allocated from the system latency budget. The acceleration limits in SUB-MAIN-039 are driven by the tissue safety constraint in SYS-MAIN-012 | 2026-03-19T23:32:52.802Z
SYS-MAIN-001 | SUB-MAIN-040 | derives | Scaling accuracy contributes to tip repeatability budget | SYS-MAIN-001 requires ±0.1mm tip repeatability; SUB-MAIN-040 allocates ±0.5% gain accuracy to the Motion Scaling Module, which at 10:1 ratio consumes 50μm of the 100μm system budget — a direct numerical derivation | 2026-03-19T23:32:52.185Z
SYS-MAIN-008 | SUB-MAIN-040 | derives | Motion Scaling Module implements the selectable ratio capability | SYS-MAIN-008 specifies the 3:1/5:1/10:1 scaling ratios; SUB-MAIN-040 allocates this function to the Motion Scaling Module with specific accuracy and timing constraints derived from the system-level tip repeatability budget | 2026-03-19T23:32:51.576Z
IFC-MAIN-025 | VER-MAIN-030 | verifies | Scaling ratio transition test for Motion Scaling to Trajectory Generator interface | Validates metadata propagation during ratio changes, the primary risk scenario for IFC-MAIN-025 | 2026-03-19T23:32:31.742Z
IFC-MAIN-028 | VER-MAIN-029 | verifies | Dynamic boundary update integrity test for Workspace Safety Enforcer to Kinematics Engine | Validates atomic update integrity under concurrent read/write, which is the safety-critical aspect of IFC-MAIN-028 | 2026-03-19T23:32:25.131Z
IFC-MAIN-027 | VER-MAIN-028 | verifies | EtherCAT frame delivery timing test for Kinematics Engine to Joint Servo Controller | Validates the deterministic delivery requirement of IFC-MAIN-027 under sustained multi-joint load, the primary integration risk | 2026-03-19T23:32:18.125Z
IFC-MAIN-026 | VER-MAIN-027 | verifies | Ring buffer stress test for Trajectory Generator to Kinematics Engine interface | Validates lock-free buffer depth and data integrity under scheduling jitter, which is the primary risk for IFC-MAIN-026 | 2026-03-19T23:32:11.317Z
IFC-MAIN-024 | VER-MAIN-026 | verifies | Integration test for Tremor Filter to Motion Scaling shared-memory interface | Validates the critical shared-memory timing and data integrity requirements of IFC-MAIN-024 under real-time operating conditions | 2026-03-19T23:32:04.549Z
SUB-MAIN-032 | VER-MAIN-022 | verifies | Instrument recognition timing requirement → multi-instrument coupling verification | VER-022 tests instrument identity chip reading from 5 different instrument types to confirm SUB-032 200ms recognition requirement on the physical hardware. | 2026-03-19T23:14:27.905Z
SYS-MAIN-002 | SUB-MAIN-037 | derives | System single-fault detection → instrument lifecycle overuse prevention requirement | SYS-002 requires detection and safe response to single-point faults; instrument overuse is a mechanical degradation hazard. The lifecycle controller prevents coupling of overused instruments before they fail in service. | 2026-03-19T23:14:11.912Z
SYS-MAIN-001 | SUB-MAIN-036 | derives | System motion precision → TTAC cable displacement computation requirement | SYS-001 requires ±0.1mm tip precision; TTAC cable displacement computation is the mechanism that translates Cartesian setpoints to per-cable commands at the required precision. | 2026-03-19T23:14:11.300Z
STK-MAIN-005 | SYS-MAIN-013 | derives | Ergonomic console operability need → system 8-hour availability requirement | STK-005 ergonomic operability for 4-hour procedures is a sub-constraint of the 8-hour availability in SYS-013; ergonomic design ensures the surgeon can operate the console for the full procedure duration without physical fatigue. | 2026-03-19T23:14:10.623Z
SYS-MAIN-002 | IFC-MAIN-009 | derives | System safety safe-state requirement → SIS to Motion Control stop command interface | SYS-002 requires the SIS to command all arms to safe state within 250ms; the SIS-to-Motion-Control command interface is the mechanism that must propagate the brake command within this budget. | 2026-03-19T23:14:09.965Z
IFC-MAIN-023 | VER-MAIN-016 | verifies | IFC-MAIN-023 interface verification | Quantified functional test confirming IFC-MAIN-023 interface data accuracy, latency, and safety chain completeness. | 2026-03-19T23:13:38.038Z
IFC-MAIN-022 | VER-MAIN-015 | verifies | IFC-MAIN-022 interface verification | Quantified functional test confirming IFC-MAIN-022 interface data accuracy, latency, and safety chain completeness. | 2026-03-19T23:13:37.414Z
IFC-MAIN-021 | VER-MAIN-014 | verifies | IFC-MAIN-021 interface verification | Quantified functional test confirming IFC-MAIN-021 interface data accuracy, latency, and safety chain completeness. | 2026-03-19T23:13:36.779Z
IFC-MAIN-020 | VER-MAIN-013 | verifies | IFC-MAIN-020 interface verification | Quantified functional test confirming IFC-MAIN-020 interface data accuracy, latency, and safety chain completeness. | 2026-03-19T23:13:36.122Z
SYS-MAIN-002 | IFC-MAIN-023 | derives | System fault detection and safe state → lifecycle lockout to safe state manager interface | SYS-002 requires safe state on single-point fault; overuse instrument lockout communicated from Lifecycle Controller to Safe State Manager ensures the SIS can interlock an instrument that has exceeded its use count. | 2026-03-19T23:12:47.887Z
SYS-MAIN-001 | IFC-MAIN-022 | derives | System motion precision → TTAC to IDU cable displacement command interface | SYS-001 precision requires accurate cable displacement commands delivered from TTAC to IDU; IFC-022 specifies the command resolution and latency of this critical motion interface. | 2026-03-19T23:12:47.228Z
SYS-MAIN-002 | IFC-MAIN-021 | derives | System safety fault detection → cable tension to SIS monitoring interface | SYS-002 requires single-point fault detection and safe state within 250ms; cable tension telemetry to the SIS is the signal path that enables detection of instrument mechanical faults. | 2026-03-19T23:12:46.616Z
SYS-MAIN-001 | IFC-MAIN-020 | derives | System precision requirement → instrument recognition to TTAC data interface | SYS-001 requires ±0.1mm tip precision; instrument-specific calibration offsets communicated from the Recognition Module to the TTAC are essential for accurate cable displacement calculation. | 2026-03-19T23:12:46.006Z
STK-MAIN-013 | SYS-MAIN-006 | derives | Rapid sterile instrument exchange → sterility compliance requirement | Rapid instrument exchange without breaking sterile field (STK-013) requires single-use sterile instruments compatible with hospital disinfectant protocols. SYS-006 captures the sterility compliance requirement. | 2026-03-19T23:12:25.397Z
STK-MAIN-003 | SYS-MAIN-015 | derives | OR infrastructure integration need → data recording system requirement | OR infrastructure integration (STK-003) requires data recording to hospital PACS/EMR; SYS-015 is the system requirement for encrypted procedure data recording and retention. | 2026-03-19T23:12:24.746Z
SUB-MAIN-017 | VER-MAIN-012 | verifies | SUB-MAIN-017 optical/display verification | Quantified physical measurement confirming SUB-MAIN-017 requirement on integrated optical/display assembly. | 2026-03-19T23:10:22.407Z
SUB-MAIN-014 | VER-MAIN-011 | verifies | SUB-MAIN-014 optical/display verification | Quantified physical measurement confirming SUB-MAIN-014 requirement on integrated optical/display assembly. | 2026-03-19T23:10:21.769Z
SUB-MAIN-013 | VER-MAIN-010 | verifies | SUB-MAIN-013 optical/display verification | Quantified physical measurement confirming SUB-MAIN-013 requirement on integrated optical/display assembly. | 2026-03-19T23:10:21.073Z
SUB-MAIN-009 | VER-MAIN-009 | verifies | Servo tracking accuracy SUB→VER | Quantified test procedure confirming SUB-MAIN-009 performance requirement on target hardware. | 2026-03-19T23:09:41.440Z
SUB-MAIN-008 | VER-MAIN-008 | verifies | Kinematics WCET SUB→VER | Quantified test procedure confirming SUB-MAIN-008 performance requirement on target hardware. | 2026-03-19T23:09:40.815Z
SUB-MAIN-007 | VER-MAIN-007 | verifies | Workspace velocity rejection SUB→VER | Quantified test procedure confirming SUB-MAIN-007 performance requirement on target hardware. | 2026-03-19T23:09:39.987Z
SUB-MAIN-005 | VER-MAIN-006 | verifies | Hardware E-stop series loop → physical break-point tests | SUB-MAIN-005 requires hardware E-stop loop independence from software; REQ-006 verifies by physical loop interruption at each segment. | 2026-03-19T23:09:01.147Z
SUB-MAIN-004 | VER-MAIN-005 | verifies | Watchdog processor isolation → power interrupt verification | SUB-MAIN-004 requires physical processor isolation for the watchdog; REQ-005 verifies this by demonstrating watchdog persistence through main power fault. | 2026-03-19T23:09:00.582Z
SUB-MAIN-002 | VER-MAIN-004 | verifies | SIS comm latency detection threshold → fault injection test | SUB-MAIN-002 specifies the communication latency threshold that triggers SIS response. REQ-004 injects known latency values to verify detection specificity and sensitivity. | 2026-03-19T23:08:29.479Z
SUB-MAIN-003 | VER-MAIN-003 | verifies | SIS safe-state timing → multi-path E-stop timing test | SUB-MAIN-003 specifies 150ms maximum time to complete safe state. REQ-003 tests all three trigger paths to confirm this timing is met across the full E-stop circuit diversity. | 2026-03-19T23:08:28.808Z
SUB-MAIN-001 | VER-MAIN-001 | verifies | SIS joint torque threshold detection → bench test procedure | SUB-MAIN-001 defines the SIS joint torque detection threshold — a SIL 3 safety function. This VER entry provides the quantified bench test confirming detection timing meets the 50ms/20ms specification. | 2026-03-19T23:07:29.487Z
SUB-MAIN-038 | VER-MAIN-025 | verifies | Degraded mode requirement verified by fault injection and isolation test | VER-MAIN-025 tests the full degraded-mode response chain specified in SUB-MAIN-038: anomaly injection, arm shutdown timing, remaining arm isolation, and operator notification. | 2026-03-19T23:01:40.317Z
SUB-MAIN-035 | VER-MAIN-024 | verifies | Sterile adapter requirement verified by endurance and dye penetration test | VER-MAIN-024 directly tests the sterile barrier integrity and torque transmission requirements in SUB-MAIN-035 through endurance cycling and dye penetration testing. | 2026-03-19T23:01:39.722Z
SUB-MAIN-034 | VER-MAIN-023 | verifies | Cable tensioning requirement verified by frequency sweep and anomaly injection test | VER-MAIN-023 directly tests the cable tension maintenance and anomaly detection requirements specified in SUB-MAIN-034. | 2026-03-19T23:01:39.094Z
SYS-MAIN-016 | SUB-MAIN-038 | derives | System degraded mode derives instrument-level cable anomaly response | SYS-MAIN-016 specifies system-level degraded operation when one arm loses servo communication. SUB-MAIN-038 derives the instrument-specific degraded mode for cable tension anomalies: the 50ms shutdown and arm isolation are the instrument subsystems implementation of the system graceful degradation requirement. | 2026-03-19T23:01:38.459Z
SYS-MAIN-006 | SUB-MAIN-035 | derives | System sterilisation requirement derives sterile adapter integrity spec | SYS-MAIN-006 requires all sterile field components to be sterilisable. SUB-MAIN-035 is the direct instrument-level derivation, specifying the Sterile Adapters barrier integrity under operating loads as the mechanism for achieving sterile compliance at the instrument-arm boundary. | 2026-03-19T23:01:23.620Z
SYS-MAIN-012 | SUB-MAIN-034 | derives | System force limit derives cable tension monitoring requirement | SYS-MAIN-012 limits force at the instrument tip. The Cable Tensioning Systems tension monitoring is the primary mechanism for detecting force anomalies that could exceed the system force limit. Tension deviation directly maps to uncontrolled tip force. | 2026-03-19T23:01:20.171Z
SYS-MAIN-001 | SUB-MAIN-033 | derives | System motion scaling derives instrument tip accuracy allocation | SYS-MAIN-001 specifies master-to-slave motion scaling. SUB-MAIN-033 allocates the instrument subsystems share of the overall position accuracy budget: 0.1mm tip accuracy is the direct instrument-level derivation of the system motion fidelity requirement. | 2026-03-19T23:01:16.726Z
SYS-MAIN-014 | SUB-MAIN-032 | derives | System instrument exchange requirement derives recognition speed sub-budget | SYS-MAIN-014 requires patient-side arms to enable instrument exchange. SUB-MAIN-032 derives the recognition speed sub-budget: without fast instrument identification, the system cannot reconfigure kinematic models in time to meet the rapid exchange requirement. | 2026-03-19T23:01:13.060Z
SYS-MAIN-003 | IFC-MAIN-008 | derives | Vision-to-console interface derives from stereo video display requirement | SYS-MAIN-003 requires delivery of stereo HD video to the surgeon; IFC-MAIN-008 specifies the transmission interface that carries the two independent video streams from Vision subsystem to the Surgeon Console display. | 2026-03-19T22:10:52.254Z
SYS-MAIN-007 | IFC-MAIN-006 | derives | Motion Control to Patient-Side Cart interface derives from command transmission requirement | SYS-MAIN-007 requires Cartesian command transmission; IFC-MAIN-006 specifies the physical joint-space interface that carries the resulting joint setpoints to the Patient-Side Cart actuators. | 2026-03-19T22:10:51.640Z
SYS-MAIN-003 | SUB-MAIN-015 | derives | Illumination regulation derives from video quality requirement | SYS-MAIN-003 requires HD stereo video with consistent image quality; SUB-MAIN-015 controls the illumination component of that quality by preventing over/under-exposure which would degrade surgical visibility. | 2026-03-19T22:10:51.035Z
SYS-MAIN-001 | SUB-MAIN-009 | derives | Joint servo tracking error derives from motion control precision requirement | SYS-MAIN-001 specifies motion scaling precision; the joint servo tracking error in SUB-MAIN-009 is the hardware allocation of the motion control precision budget. | 2026-03-19T22:10:50.403Z
SYS-MAIN-001 | SUB-MAIN-008 | derives | Kinematics engine computation derives from motion control latency requirement | SYS-MAIN-001 requires 1ms end-to-end motion control; SUB-MAIN-008 specifies the kinematics computation timing for the engine responsible for converting Cartesian commands to joint setpoints within that budget. | 2026-03-19T22:10:49.777Z
SYS-MAIN-007 | SUB-MAIN-030 | derives | Frame-level error recovery derives from command transmission integrity requirement | SYS-MAIN-007 requires reliable kinematic command transmission; SUB-MAIN-030 specifies how the Communications subsystem maintains that reliability through frame-level error containment rather than channel resets. | 2026-03-19T22:10:49.154Z
SYS-MAIN-006 | SUB-MAIN-025 | derives | Galvanic isolation derives from sterile field safety requirement | SYS-MAIN-006 requires sterile field components to be fluid-immersion compatible; SUB-MAIN-025 addresses patient leakage current isolation which is the electrical safety aspect of the same IEC 60601-1 compliance requirement. | 2026-03-19T22:10:48.500Z
SUB-MAIN-028 | VER-MAIN-021 | verifies | Failover test also verifies comms failover timing requirement | Switchover timing and zero-loss frame continuity are both tested in VER-MAIN-021. | 2026-03-19T22:09:11.454Z
IFC-MAIN-019 | VER-MAIN-021 | verifies | Fibre failover test for Network Management Controller / SIS interface | VER-MAIN-021 verifies failover timing and zero-frame-loss for both IFC-MAIN-019 (NMC-SIS status interface) and SUB-MAIN-028 (switchover requirement). | 2026-03-19T22:09:10.816Z
IFC-MAIN-018 | VER-MAIN-020 | verifies | Peak traffic latency test for Real-Time Protocol Engine / Fibre interface | VER-MAIN-020 verifies both bandwidth capacity and per-frame kinematic latency specified in IFC-MAIN-018 under maximum concurrent load. | 2026-03-19T22:09:10.142Z
SYS-MAIN-015 | SUB-MAIN-029 | derives | Procedure data recording derives from system-level data logging requirement | SYS-MAIN-015 requires 1kHz kinematic and event recording for the full procedure; SUB-MAIN-029 specifies the Procedure Data Recorder implementation of that function including 8-hour capacity and WORM protection. | 2026-03-19T22:08:43.358Z
SYS-MAIN-002 | SUB-MAIN-028 | derives | Comms failover timing derives from single-point failure handling requirement | SYS-MAIN-002 requires graceful single-point failure recovery; SUB-MAIN-028 specifies the timing and continuity constraints for fibre link failover specifically. | 2026-03-19T22:08:42.718Z
SYS-MAIN-001 | SUB-MAIN-027 | derives | Comms latency budget derives from end-to-end 1ms control loop | SYS-MAIN-001 specifies 1ms total control latency. SUB-MAIN-027 assigns the comms layer its share of that budget. | 2026-03-19T22:08:42.096Z
IFC-MAIN-017 | VER-MAIN-019 | verifies | Hardware torque limit test verifies CAN FD HC-MHA interface | VER-MAIN-019 verifies both the hardware torque limit and 1kHz command rate specified in IFC-MAIN-017. | 2026-03-19T22:06:59.841Z
IFC-MAIN-016 | VER-MAIN-018 | verifies | SPI latency test verifies digital FSC-HC interface | VER-MAIN-018 measures propagation latency and sample rate of the SPI bus specified in IFC-MAIN-016. | 2026-03-19T22:06:59.231Z
IFC-MAIN-015 | VER-MAIN-017 | verifies | CMRR test verifies analog FSM-FSC interface | VER-MAIN-017 directly tests the 80dB CMRR and cable specification of IFC-MAIN-015 under realistic OR noise frequencies. | 2026-03-19T22:06:58.619Z
SYS-MAIN-004 | SUB-MAIN-022 | derives | Force measurement resolution allocation | SYS-MAIN-004 system-level force sensing requirement derives to subsystem sensor resolution spec. | 2026-03-19T22:05:59.575Z
SYS-MAIN-002 | SUB-MAIN-026 | derives | Haptic fault containment derived from single-point failure detection requirement | SYS-MAIN-002 requires detection and graceful handling of single-point failures; SUB-MAIN-026 specifies the haptic subsystem's response to a force sensor failure, isolating the fault to the affected arm only. | 2026-03-19T22:05:53.346Z
SYS-MAIN-012 | SUB-MAIN-024 | derives | Master handle force limit derived from system instrument tip force constraint | SYS-MAIN-012 limits instrument tip force to protect patient tissue; SUB-MAIN-024 extends this principle to the master handle, limiting feedback force to protect surgeon from uncontrolled actuator response. | 2026-03-19T22:05:52.701Z
SYS-MAIN-001 | SUB-MAIN-023 | derives | Haptic latency budget derived from end-to-end control loop timing | SYS-MAIN-001 specifies 1ms end-to-end control latency; SUB-MAIN-023 allocates a 2ms haptic feedback latency which operates as a parallel channel outside the primary motion control loop, needing its own timing budget. | 2026-03-19T22:05:52.078Z
SYS-MAIN-004 | SUB-MAIN-022 | derives | Force measurement resolution allocation to Haptic Feedback Subsystem | SYS-MAIN-004 specifies system-level force sensing resolution; SUB-MAIN-022 allocates that resolution to the sensor hardware in the Haptic Feedback Subsystem, the subsystem that implements the measurement function. | 2026-03-19T22:05:51.394Z
SYS-MAIN-003 | VER-MAIN-016 | verifies | End-to-end vision chain test verifies system stereo video requirement | The end-to-end test validates the complete optical path latency, resolution, and depth accuracy that implement SYS-MAIN-003 | 2026-03-19T21:32:59.968Z
SYS-MAIN-011 | SUB-MAIN-021 | derives | Multimodal surgical visualization drives CCU enhancement modes | SYS-MAIN-011 surgical field visualisation is served by multiple imaging modalities; CCU enhancement modes implement the switching between standard, NBI, and fluorescence views | 2026-03-19T21:32:59.348Z
SYS-MAIN-011 | SUB-MAIN-020 | derives | Multimodal surgical visualization drives NIR fluorescence capability | SYS-MAIN-011 requires visualisation of the surgical field; ICG fluorescence extends this to subsurface vascular imaging needed for perfusion assessment during anastomosis | 2026-03-19T21:32:58.723Z
SYS-MAIN-003 | SUB-MAIN-016 | derives | Stereo video requirement drives IPP latency budget | System stereo video requirement implicitly demands real-time display; the 2ms IPP latency budget is part of the total visual latency chain that delivers real-time surgical feedback | 2026-03-19T21:32:58.145Z
SUB-MAIN-019 | VER-MAIN-015 | verifies | Degraded-mode test for vision system single-channel failure | Physical disconnection test validates graceful degradation path from stereo to mono operation | 2026-03-19T21:31:49.529Z
IFC-MAIN-014 | VER-MAIN-014 | verifies | Endurance test for IPP-recorder interface | 8-hour recording endurance test validates continuous recording with timestamp synchronisation accuracy | 2026-03-19T21:31:48.293Z
IFC-MAIN-013 | VER-MAIN-013 | verifies | Integration test for IPP-display interface latency and resolution | Photodiode latency and display analyser resolution/colour depth measurement validates the display interface end-to-end | 2026-03-19T21:31:47.040Z
IFC-MAIN-012 | VER-MAIN-012 | verifies | Integration test for CCU-illumination feedback loop latency | Oscilloscope measurement of step response validates the closed-loop illumination control latency | 2026-03-19T21:31:22.531Z
IFC-MAIN-011 | VER-MAIN-011 | verifies | Integration test for CCU-IPP 3G-SDI interface synchronisation | Precision skew measurement validates stereoscopic synchronisation at the CCU-IPP boundary | 2026-03-19T21:31:21.352Z
IFC-MAIN-010 | VER-MAIN-010 | verifies | Integration test for endoscope-CCU HD-SDI interface | PRBS-31 BER test validates physical layer integrity; independent channel disconnection validates graceful degradation | 2026-03-19T21:31:20.129Z
SYS-MAIN-015 | IFC-MAIN-014 | derives | System recording requirement drives recorder interface specification | SYS-MAIN-015 recording requirement drives the need for composited video output with embedded metadata for post-operative audit | 2026-03-19T21:30:55.466Z
SYS-MAIN-003 | IFC-MAIN-013 | derives | Stereo video requirement drives display interface specification | The system stereo HD video requirement determines the display interface resolution, frame rate, and colour depth needed to present the surgical field | 2026-03-19T21:30:54.863Z
SYS-MAIN-003 | IFC-MAIN-010 | derives | Stereo video requirement drives endoscope-CCU interface specification | The system stereo video requirement determines the data rate and channel independence requirements for the endoscope-to-CCU HD-SDI interface | 2026-03-19T21:30:54.243Z
SYS-MAIN-015 | SUB-MAIN-018 | derives | System recording requirement drives video recorder duration and sync | SYS-MAIN-015 requires 1kHz kinematic and video recording; the recorder's 8-hour continuous recording and frame-accurate timestamp synchronisation directly implement this system requirement | 2026-03-19T21:30:45.500Z
SYS-MAIN-002 | SUB-MAIN-019 | derives | Failure detection cascades to vision graceful degradation | The system-level single-point failure detection requirement drives the need for the vision system to detect channel loss and degrade gracefully to 2D rather than losing all visual feedback | 2026-03-19T21:30:44.877Z
SYS-MAIN-003 | SUB-MAIN-017 | derives | Stereoscopic video quality requires low display crosstalk | The system requirement for stereoscopic 3D video directly requires the display to achieve low crosstalk to deliver usable depth perception | 2026-03-19T21:30:44.299Z
SYS-MAIN-003 | SUB-MAIN-014 | derives | Stereoscopic video quality requires CCU inter-channel synchronisation | Stereoscopic 3D display requires tightly synchronised left/right channels; CCU sync requirement directly derives from the system stereo video capability | 2026-03-19T21:30:43.690Z
SYS-MAIN-003 | SUB-MAIN-013 | derives | System stereo video requirement drives endoscope optical resolution | The system-level 1080p stereo video requirement requires the endoscope to resolve at 20 lp/mm to provide sufficient optical quality for the digital imaging chain | 2026-03-19T21:30:43.227Z
SYS-MAIN-002 | SUB-MAIN-005 | derives | SYS fault response → SIS hardwired E-stop chain | Hardwired E-stop is required because software E-stop cannot provide the independence needed for SYS-MAIN-002's single-point failure coverage (software fault is one of the enumerated failure modes). | 2026-03-19T21:04:36.481Z
SYS-MAIN-002 | SUB-MAIN-004 | derives | SYS fault response → SIS watchdog physical isolation architecture | Physical processor isolation for Watchdog is an architectural necessity to satisfy the SIL 3 Hardware Fault Tolerance requirement embedded in SYS-MAIN-002's single-point failure coverage claim. | 2026-03-19T21:04:35.838Z
SYS-MAIN-007 | VER-MAIN-009 | verifies | End-to-end system latency verification test | SYS-MAIN-007 is a system-level latency requirement that can only be verified by measuring the complete signal chain from console input to instrument tip response. | 2026-03-19T21:04:06.790Z
SUB-MAIN-006 | VER-MAIN-008 | verifies | Real-time pipeline deadline verification on target hardware | SUB-MAIN-006 is a timing requirement that must be proven on production hardware under worst-case load; simulation cannot replicate OS scheduling jitter. | 2026-03-19T21:04:05.578Z
IFC-MAIN-007 | VER-MAIN-007 | verifies | Fault injection test for MC-Safety heartbeat timeout | IFC-MAIN-007 is a SIL 3 safety function; fault injection with measured brake response time is the only adequate verification method — analysis alone is insufficient. | 2026-03-19T21:03:36.542Z
IFC-MAIN-005 | VER-MAIN-006 | verifies | Integration test for console-MC command interface | Tests 100% packet delivery and 3ms latency constraint of IFC-MAIN-005 which are the two failure modes that would degrade surgical control fidelity. | 2026-03-19T21:03:35.335Z
SYS-MAIN-002 | SUB-MAIN-012 | derives | SYS fault response → SIS automatic initiation (no operator dependency) | Auto-initiation without operator confirmation is a necessary architectural property to achieve SYS-MAIN-002's 250ms fault response — operator reaction time (~500ms) would blow the budget. | 2026-03-19T21:02:59.144Z
SYS-MAIN-012 | SUB-MAIN-010 | derives | System force limit → workspace enforcer | The workspace safety enforcer is the primary implementation mechanism for the system-level 5N/8N force limits; it checks force before commanding the servos. | 2026-03-19T21:02:23.237Z
SYS-MAIN-010 | SUB-MAIN-011 | derives | E-stop spec → RT watchdog spec | The 5ms watchdog assertion deadline is derived from the 50ms emergency stop budget: 5ms watchdog + 45ms brake engagement = 50ms total, implementing the system-level emergency stop requirement. | 2026-03-19T21:02:22.613Z
SYS-MAIN-009 | SUB-MAIN-007 | derives | System tremor spec → MC tremor filter | The subsystem tremor filter requirement restates the system tremor requirement at component level, adding pipeline ordering constraint (before scaling) that is architectural rather than functional. | 2026-03-19T21:02:21.957Z
SYS-MAIN-007 | SUB-MAIN-006 | derives | System latency → MC computation budget | The 10ms Motion Control computation budget is the allocated share of the 100ms end-to-end latency; this subsystem req exists specifically because of the system latency constraint. | 2026-03-19T21:02:21.328Z
SYS-MAIN-002 | VER-MAIN-005 | verifies | End-to-end safe state verification | System-level fault injection test is the only means to verify the 250ms safe-state budget as a whole, accounting for all subsystem response times in combination. | 2026-03-19T21:01:22.676Z
IFC-MAIN-004 | VER-MAIN-004 | verifies | Protocol analyser test for safe state broadcast latency | Bus protocol verification under load proves guaranteed delivery claim; cannot be established by inspection of code alone. | 2026-03-19T21:01:06.798Z
IFC-MAIN-003 | VER-MAIN-003 | verifies | Fault injection test for communication monitor response | Latency injection and packet loss simulation are required to verify monitor response at both nominal and threshold conditions without waiting for real link faults. | 2026-03-19T21:01:05.564Z
IFC-MAIN-002 | VER-MAIN-002 | verifies | Oscilloscope verification of E-stop contactor timing | Hardware timing verification cannot be replaced by inspection; physical measurement across multiple trials at temperature extremes confirms deterministic behaviour. | 2026-03-19T21:01:04.329Z
IFC-MAIN-001 | VER-MAIN-001 | verifies | Integration test for joint torque bus latency and rate | Logic analyser measurement at integration provides objective evidence of ≥1000Hz and <200µs latency that cannot be verified by inspection alone. | 2026-03-19T21:01:03.124Z
SYS-MAIN-002 | SUB-MAIN-003 | derives | SYS fault response time budget → SIS end-to-end safe state timing | SUB-MAIN-003 allocates the full SYS-MAIN-002 250ms system budget to SIS as the orchestrating safety subsystem. | 2026-03-19T20:59:56.235Z
SYS-MAIN-002 | SUB-MAIN-002 | derives | SYS fault response → SIS communication loss detection | SUB-MAIN-002 implements SYS-MAIN-002's safe-state requirement specifically for the communication loss failure mode. | 2026-03-19T20:59:55.623Z
SYS-MAIN-002 | SUB-MAIN-001 | derives | SYS fault response → SIS joint force monitoring | SUB-MAIN-001 is the primary mechanism implementing SYS-MAIN-002's 250ms safe-state requirement for the joint force fault mode. | 2026-03-19T20:59:55.016Z
STK-MAIN-011 | SYS-MAIN-014 | derives | Sterility need → sterilisation spec | The EN ISO 17665 sterilisation requirement directly implements the sterile field stakeholder requirement; material constraints cascade from this trace. | 2026-03-19T20:59:00.831Z
STK-MAIN-010 | SYS-MAIN-016 | derives | Safety need → degraded mode | Degraded operation without abort implements safety need: abrupt mid-procedure shutdown is itself a patient hazard requiring graceful degradation instead. | 2026-03-19T20:58:52.652Z
STK-MAIN-015 | SYS-MAIN-013 | derives | Availability need → 8h operation | Continuous 8-hour operation directly implements the surgical list availability stakeholder need. | 2026-03-19T20:58:52.070Z
STK-MAIN-014 | SYS-MAIN-015 | derives | Data recording need → recording spec | System recording requirement directly implements the clinical audit stakeholder need; retention period and resolution derived from regulatory requirements. | 2026-03-19T20:58:51.486Z
STK-MAIN-012 | SYS-MAIN-011 | derives | Visualisation need → video spec | The 1080p/60Hz stereo video specification directly implements the situational awareness stakeholder need; parameters derived from clinical evidence on depth perception. | 2026-03-19T20:58:50.872Z
STK-MAIN-010 | SYS-MAIN-012 | derives | Safety need → force limit | Force limiting prevents inadvertent tissue damage even under normal operation; it directly addresses the safety stakeholder requirement at the instrument-tissue interface. | 2026-03-19T20:58:50.270Z
STK-MAIN-010 | SYS-MAIN-010 | derives | Safety need → emergency stop | The emergency stop function is the primary safety mechanism preventing unintended motion; it directly implements the do-no-harm stakeholder requirement. | 2026-03-19T20:58:49.650Z
STK-MAIN-009 | SYS-MAIN-009 | derives | Precision need → tremor filtration | Tremor filtration is the second mechanism enabling precision: without filtration, scaled motion amplifies tremor, negating the precision benefit. | 2026-03-19T20:58:49.003Z
STK-MAIN-009 | SYS-MAIN-008 | derives | Precision need → motion scaling | Motion scaling is the primary technical mechanism delivering surgical precision beyond freehand; scales derive directly from the precision stakeholder need. | 2026-03-19T20:58:48.412Z
STK-MAIN-009 | SYS-MAIN-007 | derives | Precision need → latency budget | Surgeon precision requires real-time motion response; 100ms latency budget directly implements the precision stakeholder need by bounding perceptible lag. | 2026-03-19T20:58:47.792Z
STK-MAIN-004 | SYS-MAIN-006 | derives | STK sterile field → SYS sterile compliance | ISO 11135 sterility and disinfectant compatibility directly implement the sterile-field stakeholder requirement. | 2026-03-19T20:56:49.115Z
STK-MAIN-002 | SYS-MAIN-005 | derives | STK safety → SYS power continuity | 60s UPS bridge is the quantified implementation of the no-uncontrolled-force stakeholder safety need during mains failure. | 2026-03-19T20:56:48.488Z
STK-MAIN-001 | SYS-MAIN-004 | derives | STK precision → SYS force sensing resolution | Force sensing at ≤0.1N resolution is needed to achieve the tissue discrimination required for sub-millimetre precision surgery. | 2026-03-19T20:56:47.878Z
STK-MAIN-001 | SYS-MAIN-003 | derives | STK precision → SYS imaging quality | Sub-millimetre precision depends on imaging quality; 1080p/60Hz/100ms latency bounds are minimum for ±0.1mm tip operation. | 2026-03-19T20:56:47.228Z
STK-MAIN-002 | SYS-MAIN-002 | derives | STK patient safety → SYS fault response requirement | SYS-MAIN-002 allocates a quantified 250ms fault response budget derived from the patient harm hazard analysis backing STK-MAIN-002. | 2026-03-19T20:56:46.612Z
STK-MAIN-001 | SYS-MAIN-001 | derives | STK precision need → SYS motion scaling and latency spec | SYS-MAIN-001 quantifies the precision performance (±0.1mm, <1ms) directly required to satisfy the surgeon sub-millimetre accuracy need. | 2026-03-19T20:56:45.844Z

## Linksets
Source Document | Target Document | Default Link Type | Link Count
architecture-decisions | subsystem-requirements |  | 0
architecture-decisions | system-requirements |  | 16
system-requirements | verification-plan |  | 11
interface-requirements | verification-plan |  | 75
subsystem-requirements | verification-plan |  | 113
system-requirements | interface-requirements |  | 14
system-requirements | subsystem-requirements |  | 166
stakeholder-requirements | system-requirements |  | 20

## Sections
Document | Section | Short Code | Order | Requirements
architecture-decisions | Main |  | 0 | 15
interface-requirements | Main |  | 0 | 48
stakeholder-requirements | Main |  | 0 | 12
subsystem-requirements | Main |  | 0 | 162
system-requirements | Main |  | 0 | 19
verification-plan | Main |  | 0 | 193